## Data
Example data only |  |  |  | Version: Pb_Auto_V1ad, updated 27 Jan 2015, validated by TJS 27 Jan 2015. | [13 empty] | The formulae below need to be copied below if the spreadsheet needs to be "livened up". See PowerCalc A11 | [48 empty] | source select table
2 | HEG | 10 | 230 | 0 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | #N/A | 6 | 1 |  | ACal All done. 00.00°C 00.00%. | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [41 empty] | HEG | CH | HV | Ioff
cell number | [39 empty] | current_row_number | 166 | [25 empty] | 0 | 1 | 10 | 100
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80 | 81 | 82 | 83 | 84 | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105
| FLUHIGH | Current | Voltage | Phase | DivRange | Shunt |  | Freqy | [18 empty] | Exact Freq | Wdial | Wsign | Vdial | Vsign | [7 empty] | Trial Balance Calculations | [27 empty] | W count | W sign | V count | V sign
|  | This is NOT a data entry block. (Used by trial balance to find dial settings.) | [24 empty] | Trial Balance Data |  |  |  |  | FFT Det Volts | FFT Det Phase | RefVolts | FFT RefVolts | FFT Ref phase |  | Det gain | ExactFreq | Divider ratio K | Divider ratio K | Ref Volts V1(cor'd) | Shunt impedance R3 | Shunt impedance R3 | Watts dial alpha | Watts buffer gain | R2/R1 | R2/R1 | Vars dial beta | Vars buffer gain | g_1 conductance | Y1 admittance | R2 impedance | R2 impedance | Detector amp volts | Detector phase | Det Amp Gain G | Det Amp phase Gj | Det amp conductance g_d | Det amp admittance Yd | Watts | Vars | Line Current | Line Potential | Phase | Power Factor | [6 empty] | watts or vars | MeanWatts | MeanVars | Meanvolts | MeanAmps | MeanDeg | Watts_stdev | Var_stdev | Volts_stdev | Amps_stdev | Phase_stdev | Imp/w-sec Error (%) | Combined std unc | degrees of freedom | Coverage factor | Expanded Uncertainty | Power/var factor | Mean Ratio | Ratio Stdev | Impulses/w-hr or var-hr |  |  | PLT Reading (Ohms) | Temperature (°C) | C1 (T corrected)
166 | #N/A |  |  |  | 230 | ESI(1) | 100 | 50 | [18 empty] | 50.0005 | 0.7041 | WP- | 0.52539 | VP- | 2.2491 | 11.4058 |  | 1.1324 | 107.9567 |  | Max | 50.0005 | 202.9816 | 0.021862 | 1.1324 | 0.99999 | 1.0367e-05 | 0.7041 | -1.0003 | 1 | -1.8848e-07 | 0.52539 | -1.0003 | 1.8311e-08 | 0.00010145 | 10000.0599 | -0.03016 | 2.2491 | -1.6851 | 1135.2724 | -0.045847 | 1e-07 | 2e-10 | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? |  |  | watts | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? |  |  | 109.115 | 23.5249 | 0.32293
Test Setup Columns: Enter test condition data here. Test sequence will stop if there is no number in "Row Number" | [17 empty] | Results | [7 empty] | 13:57 03/15/04
Row Number | Set source | Set Current | Set Voltage | Set Phase. For HEG lag is positive angle (IEC), and for CH or FLUKE lag is negative (Non-IEC) | PB divider range | Shunt | Current Transformer Ratio | Frequency | Counter GateTime | Number of Readings | Impulses/Wh or varh Ch1 | watts/vars | Type B unc Ch 1 | Ch 1 type B deg of freedom | RD31 Phase connected | DVM DCV Range Test or Ch2 gate time. | Remarks | Voltage | Current | Non-IEC Phase Difference | Power /var factor | Error (%) | Relative Expanded Uncertainty | Coverage Factor | Impulses/w or var hr | Time & Date | Exact Freq | Wdial | Wsign | Vdial | Vsign | FFT Det Volts | FFT Det Phase | RefVolts | FFT RefVolts | FFT Ref phase | Counter | Counter | FFT Det Volts | FFT Det Phase | RefVolts | FFT RefVolts | FFT Ref phase | Counter | Counter | FFT Det Volts | FFT Det Phase | RefVolts | FFT RefVolts | FFT Ref phase | Counter | Counter | FFT Det Volts | FFT Det Phase | RefVolts | FFT RefVolts | FFT Ref phase | Counter | Counter | FFT Det Volts | FFT Det Phase | RefVolts | FFT RefVolts | FFT Ref phase | Counter | Counter | [7 empty] | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 |  |  | Do not copy this column when livening.
| Enter: HEG or FLUKE FLUHIGH CH HV Ioff |  |  |  | 60 120 230 | ESI(.1) (1 ohm P) T(122) C(.2) ESI(1) (10 ohm P) ESI(10) (10 ohm S-P) NoShunt |  |  | Channel 1 | 1 to 5 |  | watts vars | enter type B uncertainty of DUT | enter type B dof of DUT | 1,2,3 or 123 | Other info if not the above | [15 empty] | volts | degrees | volts | volts | degrees | Ch 1 Freqy Hz | Ch 2 Freqy Hz | volts | degrees | volts | volts | degrees | Ch 1 Freqy Hz | Ch 2 Freqy Hz | volts | degrees | volts | volts | degrees | Ch 1 Freqy Hz | Ch 2 Freqy Hz | volts | degrees | volts | volts | degrees | Ch 1 Freqy Hz | Ch 2 Freqy Hz | volts | degrees | volts | volts | degrees | Ch 1 Freqy Hz | Ch 2 Freqy Hz |  |  |  |  | Shunt Volts | [23 empty] | Pt Resistance (Ohms) | Temperature (°C) | Corrected C1 (uF)
Row 12 is start row. This row left blank on purpose. Example of the data inserted is in Row 2
12 | FLUHIGH | 2 | 230 | 0 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0006 | 6 | 123 |  | ACal All done & RD31 reset 19.76°C 65.69%. | 229.9965 | 2.0001 | -0.0017044 | 1 | 0.0020602 | 0.0035449 | 2.002 | 100002.0602 | 2015-01-29 14:26 | 50.0002 | 0.35254 | WP- | 0 | VP+ | 0.21171 | 83.6624 | 1.1331 | 1.1332 | 107.9576 | 12778.4795 | 460.0298 | 0.2138 | 84.2708 | 1.1331 | 1.133 | 107.9456 | 12778.4589 | 460.0307 | 0.21233 | 83.4215 | 1.1331 | 1.133 | 107.9541 | 12778.447 | 460.0304 | 0.21147 | 84.1761 | 1.1331 | 1.1332 | 107.9435 | 12778.4327 | 460.0298 | 0.21563 | 84.3992 | 1.1331 | 1.1327 | 107.9749 | 12778.4041 | 460.0298 | [7 empty] | 460.0145 | 0.013684 | 229.9965 | 2.0001 | -0.0017044 | 0.0019392 | 0.0016898 | 9.0498e-05 | 7.8157e-06 | 0.00021047 | 0.0020602 | 0.0017707 | 57.0064 | 2.002 | 0.0035449 | 1 | 114.9929 | 0.00092994 | 100002.0602 | 0 | 0 | 109.147 | 23.6074 | 0.32293
13 | FLUHIGH | 2 | 230 | 0 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9968 | 2.0001 | -0.0025417 | 1 | 0.0032208 | 0.0047755 | 2.0756 | 100003.2208 | 2015-01-29 14:30 | 50.0006 | 0.35352 | WP- | 0 | VP+ | 0.34923 | -104.8376 | 1.1331 | 1.1333 | 107.9573 | 12778.6854 | 460.0299 | 0.35007 | -104.5433 | 1.1331 | 1.1332 | 107.9418 | 12778.5736 | 460.03 | 0.35062 | -104.3905 | 1.1331 | 1.1327 | 107.9472 | 12778.463 | 460.0292 | 0.34779 | -104.1258 | 1.1331 | 1.133 | 107.943 | 12778.5832 | 460.0312 | 0.34806 | -104.5034 | 1.1331 | 1.1332 | 107.9307 | 12778.5821 | 460.0273 | [7 empty] | 460.014 | 0.020407 | 229.9968 | 2.0001 | -0.0025417 | 0.0011011 | 0.0016419 | 8.547e-05 | 5.0027e-06 | 0.0002045 | 0.0032208 | 0.0023008 | 21.1843 | 2.0756 | 0.0047755 | 1 | 114.9933 | 0.00068315 | 100003.2208 | 0 | 0 | 109.153 | 23.6228 | 0.32293
14 | FLUHIGH | 2 | 230 | 0 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9967 | 2.0001 | -0.0018758 | 1 | 0.0029408 | 0.0048602 | 2.0678 | 100002.9408 | 2015-01-29 14:34 | 50.0002 | 0.35254 | WP- | 0 | VP+ | 0.21011 | 83.776 | 1.1331 | 1.1334 | 107.9542 | 12778.3136 | 460.0261 | 0.21245 | 83.7716 | 1.1331 | 1.1332 | 107.9575 | 12778.4391 | 460.0305 | 0.213 | 84.3381 | 1.1331 | 1.1326 | 107.9543 | 12778.4475 | 460.0313 | 0.21091 | 83.4945 | 1.1331 | 1.1334 | 107.9632 | 12778.4908 | 460.0279 | 0.21154 | 83.893 | 1.1331 | 1.1329 | 107.9265 | 12778.7866 | 460.0319 | [7 empty] | 460.0123 | 0.01506 | 229.9967 | 2.0001 | -0.0018758 | 0.0010299 | 0.0012444 | 8.0652e-05 | 4.582e-06 | 0.000155 | 0.0029408 | 0.0023504 | 22.6551 | 2.0678 | 0.0048602 | 1 | 114.9936 | 0.0006155 | 100002.9408 | 0 | 0 | 109.162 | 23.646 | 0.32292
15 | FLUHIGH | 2 | 230 | -60 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9966 | 2 | -60.0003 | 0.5 | -0.0038166 | 0.0086789 | 2.0452 | 99996.1834 | 2015-01-29 14:38 | 50.0002 | 0.17676 | WP- | 0.30176 | VP+ | 0.28297 | -165.3511 | 1.1331 | 1.1331 | 107.9527 | 6388.732 | 229.9904 | 0.28426 | -165.486 | 1.1331 | 1.1332 | 107.9371 | 6388.5001 | 229.9907 | 0.28299 | -165.2974 | 1.1331 | 1.133 | 107.9453 | 6388.5996 | 229.9897 | 0.283 | -165.167 | 1.1331 | 1.1329 | 107.924 | 6388.6674 | 229.9897 | 0.27877 | -164.9589 | 1.1331 | 1.1334 | 107.9341 | 6388.6789 | 229.9894 | [7 empty] | 229.9997 | 398.3752 | 229.9966 | 2 | -60.0003 | 0.00056932 | 0.0026214 | 9.5657e-05 | 9.7766e-06 | 0.00014262 | -0.0038166 | 0.0042436 | 28.4735 | 2.0452 | 0.0086789 | 0.5 | 114.9959 | 0.0011592 | 99996.1834 | 0 | 0 | 109.173 | 23.6744 | 0.32292
16 | FLUHIGH | 2 | 230 | -60 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9967 | 2 | -60.0006 | 0.49999 | -0.0024019 | 0.0086921 | 2.0447 | 99997.5981 | 2015-01-29 14:43 | 50.0005 | 0.17676 | WP- | 0.30176 | VP+ | 0.28385 | -164.7058 | 1.1331 | 1.1332 | 107.9724 | 6388.6999 | 229.9921 | 0.28388 | -165.1433 | 1.1331 | 1.1335 | 107.9243 | 6388.6422 | 229.991 | 0.28088 | -165.0272 | 1.1331 | 1.133 | 107.957 | 6388.7024 | 229.9907 | 0.28164 | -164.7448 | 1.1331 | 1.1328 | 107.94 | 6388.605 | 229.9902 | 0.28124 | -165.4931 | 1.1331 | 1.1329 | 107.9528 | 6388.7164 | 229.99 | [7 empty] | 229.9978 | 398.3782 | 229.9967 | 2 | -60.0006 | 0.0017768 | 0.0013841 | 6.9402e-05 | 6.811e-06 | 0.00019377 | -0.0024019 | 0.0042511 | 28.6386 | 2.0447 | 0.0086921 | 0.49999 | 114.9956 | 0.00084661 | 99997.5981 | 0 | 0 | 109.18 | 23.6924 | 0.32292
17 | FLUHIGH | 2 | 230 | -60 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9967 | 2 | -60.0004 | 0.49999 | -0.0032777 | 0.0086925 | 2.0447 | 99996.7223 | 2015-01-29 14:47 | 50.0003 | 0.17676 | WP- | 0.30176 | VP+ | 0.2809 | -165.3927 | 1.1331 | 1.1328 | 107.9395 | 6388.5905 | 229.9882 | 0.28291 | -165.1911 | 1.1331 | 1.1331 | 107.971 | 6388.6674 | 229.9909 | 0.28378 | -164.9136 | 1.1331 | 1.1328 | 107.9528 | 6388.6385 | 229.9907 | 0.28206 | -165.0593 | 1.1331 | 1.1327 | 107.968 | 6388.637 | 229.9862 | 0.2841 | -164.8251 | 1.1331 | 1.1333 | 107.9399 | 6388.5563 | 229.9894 | [7 empty] | 229.9978 | 398.374 | 229.9967 | 2 | -60.0004 | 0.0015789 | 0.00085142 | 3.8223e-05 | 6.1485e-06 | 0.00013233 | -0.0032777 | 0.0042513 | 28.6439 | 2.0447 | 0.0086925 | 0.49999 | 114.9965 | 0.00077502 | 99996.7223 | 0 | 0 | 109.189 | 23.7156 | 0.32292
18 | FLUHIGH | 2 | 230 | 36.87 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9969 | 2 | 36.8669 | 0.80003 | 0.0045598 | 0.0057273 | 2.0523 | 100004.5598 | 2015-01-29 14:51 | 50 | 0.28223 | WP- | 0.20898 | VP- | 0.14344 | -3.6624 | 1.1331 | 1.1327 | 107.9267 | 10223.1668 | 368.0416 | 0.1421 | -4.0438 | 1.1331 | 1.1331 | 107.9066 | 10222.9412 | 368.0371 | 0.14607 | -4.542 | 1.1331 | 1.1334 | 107.9454 | 10223.0619 | 368.0376 | 0.14663 | -3.6263 | 1.1331 | 1.1329 | 107.9517 | 10223.2674 | 368.04 | 0.14779 | -3.0573 | 1.1331 | 1.1333 | 107.9734 | 10223.2842 | 368.0367 | [7 empty] | 368.0164 | 275.9825 | 229.9969 | 2 | 36.8669 | 0.0013944 | 0.00242 | 5.7144e-05 | 5.9635e-06 | 0.00030099 | 0.0045598 | 0.0027906 | 26.3271 | 2.0523 | 0.0057273 | 0.80003 | 114.9963 | 0.0007557 | 100004.5598 | 0 | 0 | 109.198 | 23.7388 | 0.32292
19 | FLUHIGH | 2 | 230 | 36.87 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9966 | 2 | 36.8664 | 0.80004 | 0.0036358 | 0.0057159 | 2.0531 | 100003.6358 | 2015-01-29 14:55 | 50.0006 | 0.28223 | WP- | 0.20898 | VP- | 0.14852 | -2.5781 | 1.1331 | 1.1335 | 107.9936 | 10223.1454 | 368.0432 | 0.14543 | -3.963 | 1.1331 | 1.1329 | 107.975 | 10223.0039 | 368.0405 | 0.14945 | -3.1754 | 1.1331 | 1.1333 | 107.9462 | 10223.2672 | 368.0364 | 0.14558 | -3.341 | 1.1331 | 1.1336 | 107.9433 | 10223.0989 | 368.0411 | 0.1484 | -3.5778 | 1.1331 | 1.133 | 107.9306 | 10223.0556 | 368.0375 | [7 empty] | 368.0187 | 275.9792 | 229.9966 | 2 | 36.8664 | 0.0014372 | 0.0018659 | 8.7555e-05 | 4.4176e-06 | 0.00025939 | 0.0036358 | 0.002784 | 26.1103 | 2.0531 | 0.0057159 | 0.80004 | 114.996 | 0.0005349 | 100003.6358 | 0 | 0 | 109.208 | 23.7646 | 0.32292
20 | FLUHIGH | 2 | 230 | 36.87 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9967 | 2 | 36.8664 | 0.80004 | 0.0035533 | 0.0058122 | 2.0477 | 100003.5533 | 2015-01-29 14:59 | 50.0006 | 0.28223 | WP- | 0.20898 | VP- | 0.14554 | -3.1104 | 1.1331 | 1.1335 | 107.9394 | 10223.0954 | 368.0374 | 0.1471 | -2.7169 | 1.1331 | 1.1332 | 107.9598 | 10223.2285 | 368.0411 | 0.14822 | -3.24 | 1.1331 | 1.1332 | 107.9578 | 10223.3731 | 368.0398 | 0.14732 | -3.1338 | 1.1331 | 1.1334 | 107.9521 | 10222.9643 | 368.0396 | 0.14731 | -3.2633 | 1.1331 | 1.1332 | 107.9427 | 10223.0738 | 368.0382 | [7 empty] | 368.0202 | 275.9801 | 229.9967 | 2 | 36.8664 | 0.00058493 | 0.00091586 | 0.00010897 | 3.0093e-06 | 0.00010508 | 0.0035533 | 0.0028384 | 27.6836 | 2.0477 | 0.0058122 | 0.80004 | 114.9956 | 0.00042715 | 100003.5533 | 0 | 0 | 109.216 | 23.7852 | 0.32292
21 | FLUHIGH | 5 | 230 | 0 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9969 | 5.0002 | -0.0012672 | 1 | 0.0029146 | 0.0048031 | 2.0737 | 100002.9146 | 2015-01-29 15:04 | 50.0011 | 0.88281 | WP- | 0 | VP+ | 0.31931 | -110.1093 | 1.1331 | 1.1329 | 107.9295 | 31946.3068 | 1150.0745 | 0.31876 | -109.764 | 1.1331 | 1.1336 | 107.9634 | 31945.9847 | 1150.0712 | 0.32041 | -110.4943 | 1.1331 | 1.1324 | 107.937 | 31945.9682 | 1150.0681 | 0.32031 | -110.227 | 1.1331 | 1.1332 | 107.9606 | 31946.4752 | 1150.0684 | 0.32024 | -110.6311 | 1.1331 | 1.1329 | 107.949 | 31946.8934 | 1150.0688 | [7 empty] | 1150.0342 | 0.025436 | 229.9969 | 5.0002 | -0.0012672 | 0.00064261 | 0.0019747 | 3.2279e-05 | 2.4083e-06 | 9.8382e-05 | 0.0029146 | 0.0023162 | 21.509 | 2.0737 | 0.0048031 | 1 | 45.9974 | 4.4944e-05 | 100002.9146 | 0 | 0 | 109.218 | 23.7903 | 0.32292
22 | FLUHIGH | 5 | 230 | 0 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9967 | 5.0002 | -0.001244 | 1 | 0.0029853 | 0.0047237 | 2.0819 | 100002.9853 | 2015-01-29 15:08 | 50 | 0.88281 | WP- | 0 | VP+ | 0.32124 | -110.2646 | 1.1331 | 1.1334 | 107.9252 | 31946.517 | 1150.0679 | 0.31933 | -110.2341 | 1.1331 | 1.1332 | 107.9745 | 31946.2262 | 1150.0719 | 0.31755 | -110.3005 | 1.1331 | 1.1331 | 107.9555 | 31946.128 | 1150.0751 | 0.32262 | -110.3356 | 1.1331 | 1.1331 | 107.948 | 31946.1382 | 1150.0642 | 0.32013 | -110.2234 | 1.1331 | 1.1329 | 107.9378 | 31946.3732 | 1150.0687 | [7 empty] | 1150.0316 | 0.02497 | 229.9967 | 5.0002 | -0.001244 | 0.0014882 | 0.00042841 | 5.4405e-05 | 7.3797e-06 | 2.1343e-05 | 0.0029853 | 0.0022689 | 20.117 | 2.0819 | 0.0047237 | 1 | 45.9974 | 0.00017191 | 100002.9853 | 0 | 0 | 109.214 | 23.78 | 0.32292
23 | FLUHIGH | 5 | 230 | 0 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9969 | 5.0002 | -0.0012191 | 1 | 0.0025811 | 0.0047098 | 2.0838 | 100002.5811 | 2015-01-29 15:12 | 50.0008 | 0.88281 | WP- | 0 | VP+ | 0.32042 | -110.1733 | 1.1331 | 1.1328 | 107.9563 | 31946.2288 | 1150.0717 | 0.31845 | -110.8853 | 1.1331 | 1.1325 | 107.9571 | 31946.1782 | 1150.0668 | 0.31934 | -109.8434 | 1.1331 | 1.1334 | 107.9548 | 31946.3457 | 1150.0732 | 0.31928 | -110.4763 | 1.1331 | 1.1338 | 107.9364 | 31946.145 | 1150.072 | 0.32249 | -110.2043 | 1.1331 | 1.133 | 107.943 | 31946.1221 | 1150.0676 | [7 empty] | 1150.0337 | 0.024469 | 229.9969 | 5.0002 | -0.0012191 | 0.0019139 | 0.0020829 | 5.2995e-05 | 7.878e-06 | 0.00010378 | 0.0025811 | 0.0022603 | 19.8314 | 2.0838 | 0.0047098 | 1 | 45.9974 | 0.00015603 | 100002.5811 | 0 | 0 | 109.212 | 23.7749 | 0.32292
24 | FLUHIGH | 5 | 230 | -60 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9967 | 5 | -59.9986 | 0.50002 | -0.0045711 | 0.008531 | 2.0515 | 99995.4289 | 2015-01-29 15:16 | 49.9999 | 0.44141 | WP- | 0.75293 | VP+ | 0.22386 | -18.5739 | 1.1331 | 1.1334 | 107.9457 | 15972.27 | 574.9882 | 0.22237 | -19.1531 | 1.1331 | 1.1335 | 107.9313 | 15971.8324 | 574.985 | 0.22079 | -18.3379 | 1.1331 | 1.1333 | 107.9586 | 15972.2713 | 574.9791 | 0.22318 | -19.1859 | 1.1331 | 1.1331 | 107.9365 | 15971.9936 | 574.9858 | 0.2219 | -18.3998 | 1.1331 | 1.1331 | 107.9585 | 15971.872 | 574.9816 | [7 empty] | 575.02 | 995.9075 | 229.9967 | 5 | -59.9986 | 0.0018052 | 0.0011803 | 7.1631e-05 | 4.8048e-06 | 8.5314e-05 | -0.0045711 | 0.0041584 | 26.5514 | 2.0515 | 0.008531 | 0.50002 | 45.999 | 8.8509e-05 | 99995.4289 | 0 | 0 | 109.21 | 23.7697 | 0.32292
25 | FLUHIGH | 5 | 230 | -60 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9968 | 5.0001 | -59.9988 | 0.50002 | -0.0042056 | 0.0086098 | 2.0483 | 99995.7944 | 2015-01-29 15:20 | 50.0002 | 0.44141 | WP- | 0.75293 | VP+ | 0.22114 | -19.237 | 1.1331 | 1.1332 | 107.9518 | 15972.0679 | 574.9835 | 0.22318 | -18.344 | 1.1331 | 1.1333 | 107.9593 | 15971.8891 | 574.9827 | 0.22172 | -18.9442 | 1.1331 | 1.1332 | 107.9545 | 15971.8296 | 574.9882 | 0.22386 | -18.4771 | 1.1331 | 1.1333 | 107.9565 | 15972.2579 | 574.981 | 0.22407 | -18.6041 | 1.1331 | 1.1333 | 107.9277 | 15972.5378 | 574.9857 | [7 empty] | 575.0204 | 995.915 | 229.9968 | 5.0001 | -59.9988 | 0.0013974 | 0.0015463 | 4.8583e-05 | 8.3549e-06 | 2.8848e-05 | -0.0042056 | 0.0042033 | 27.485 | 2.0483 | 0.0086098 | 0.50002 | 45.9988 | 0.00016719 | 99995.7944 | 0 | 0 | 109.211 | 23.7723 | 0.32292
26 | FLUHIGH | 5 | 230 | -60 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9968 | 5 | -59.9988 | 0.50002 | -0.0039859 | 0.0085568 | 2.0504 | 99996.0141 | 2015-01-29 15:24 | 50.0001 | 0.44141 | WP- | 0.75293 | VP+ | 0.22345 | -18.258 | 1.1331 | 1.1333 | 107.961 | 15972.0984 | 574.9803 | 0.22454 | -19.396 | 1.1331 | 1.1328 | 107.9419 | 15971.7824 | 574.9836 | 0.22399 | -19.9231 | 1.1331 | 1.1328 | 107.9495 | 15972.2341 | 574.9838 | 0.22132 | -18.3353 | 1.1331 | 1.133 | 107.934 | 15972.3763 | 574.9858 | 0.22301 | -18.7464 | 1.1331 | 1.1337 | 107.9347 | 15971.9756 | 574.9829 | [7 empty] | 575.0183 | 995.9134 | 229.9968 | 5 | -59.9988 | 0.0029253 | 0.00078916 | 8.4769e-05 | 4.8096e-06 | 0.00014075 | -0.0039859 | 0.0041733 | 26.8776 | 2.0504 | 0.0085568 | 0.50002 | 45.9989 | 0.0001134 | 99996.0141 | 0 | 0 | 109.211 | 23.7723 | 0.32292
27 | FLUHIGH | 5 | 230 | 36.87 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9968 | 5 | 36.8671 | 0.80003 | 0.006758 | 0.0057236 | 2.0533 | 100006.758 | 2015-01-29 15:29 | 49.9999 | 0.70605 | WP- | 0.52148 | VP- | 0.3068 | -141.7191 | 1.1331 | 1.133 | 107.9455 | 25557.7401 | 920.086 | 0.30762 | -142.694 | 1.1331 | 1.1332 | 107.9491 | 25557.9955 | 920.0854 | 0.30721 | -141.9703 | 1.1331 | 1.1331 | 107.9546 | 25557.7014 | 920.0893 | 0.31079 | -142.1201 | 1.1331 | 1.134 | 107.9624 | 25557.7943 | 920.0873 | 0.30618 | -142.1221 | 1.1331 | 1.1335 | 107.9366 | 25558.6621 | 920.0861 | [7 empty] | 920.0251 | 689.9487 | 229.9968 | 5 | 36.8671 | 0.0022625 | 0.0019231 | 7.8197e-05 | 8.9238e-06 | 9.4855e-05 | 0.006758 | 0.0027875 | 26.0652 | 2.0533 | 0.0057236 | 0.80003 | 45.9991 | 0.00016825 | 100006.758 | 0 | 0 | 109.21 | 23.7697 | 0.32292
28 | FLUHIGH | 5 | 230 | 36.87 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9968 | 5 | 36.8673 | 0.80003 | 0.0063187 | 0.0056737 | 2.0562 | 100006.3187 | 2015-01-29 15:33 | 50.0002 | 0.70605 | WP- | 0.52148 | VP- | 0.30717 | -142.5118 | 1.1331 | 1.1333 | 107.9531 | 25558.3411 | 920.0869 | 0.30655 | -142.4675 | 1.1331 | 1.1335 | 107.9327 | 25557.8009 | 920.0923 | 0.30928 | -142.0718 | 1.1331 | 1.133 | 107.9692 | 25557.8479 | 920.0859 | 0.30721 | -142.3679 | 1.1331 | 1.1329 | 107.9302 | 25557.766 | 920.0909 | 0.30863 | -142.1793 | 1.1331 | 1.1325 | 107.9415 | 25557.7431 | 920.0964 | [7 empty] | 920.0263 | 689.9533 | 229.9968 | 5 | 36.8673 | 0.0016434 | 0.00039263 | 0.00010176 | 5.2338e-06 | 4.6597e-05 | 0.0063187 | 0.0027593 | 25.285 | 2.0562 | 0.0056737 | 0.80003 | 45.999 | 0.00010716 | 100006.3187 | 0 | 0 | 109.212 | 23.7749 | 0.32292
29 | FLUHIGH | 5 | 230 | 36.87 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9967 | 5 | 36.8673 | 0.80003 | 0.0064108 | 0.0056157 | 2.0607 | 100006.4108 | 2015-01-29 15:37 | 50.0003 | 0.70605 | WP- | 0.52148 | VP- | 0.3051 | -142.425 | 1.1331 | 1.1331 | 107.9419 | 25557.9686 | 920.0937 | 0.30802 | -142.0126 | 1.1331 | 1.1331 | 107.9573 | 25557.7593 | 920.088 | 0.30947 | -142.3414 | 1.1331 | 1.1331 | 107.9408 | 25557.747 | 920.0906 | 0.30786 | -142.5426 | 1.1331 | 1.133 | 107.9642 | 25558.1187 | 920.0935 | 0.30883 | -142.414 | 1.1331 | 1.1334 | 107.9601 | 25558.0076 | 920.0924 | [7 empty] | 920.0261 | 689.9542 | 229.9967 | 5 | 36.8673 | 0.0013944 | 0.0013901 | 4.2238e-05 | 4.8469e-06 | 8.1833e-05 | 0.0064108 | 0.0027251 | 24.1873 | 2.0607 | 0.0056157 | 0.80003 | 45.999 | 0.00010568 | 100006.4108 | 0 | 0 | 109.218 | 23.7903 | 0.32292
30 | FLUHIGH | 2 | 230 | -90 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0006 | 6 | 123 |  |  | 229.9967 | 2 | -90.0004 | 1 | 0.00093678 | 0.01116 | 2.1226 | 100000.9368 | 2015-01-29 15:41 | 50.0004 | 0 | WP+ | 0.34766 | VP+ | 0.19066 | 6.725 | 1.1331 | 1.1332 | 107.9866 | 12777.4798 | 460.029 | 0.19127 | 6.4208 | 1.1331 | 1.1327 | 107.9774 | 12777.4681 | 460.0314 | 0.19047 | 7.1379 | 1.1331 | 1.1327 | 107.9516 | 12777.4844 | 460.0294 | 0.19057 | 6.6883 | 1.1331 | 1.1329 | 107.9456 | 12777.4964 | 460.0245 | 0.19037 | 6.7255 | 1.1331 | 1.1332 | 107.9486 | 12777.4594 | 460.0254 | [7 empty] | -0.0028848 | 459.9849 | 229.9967 | 2 | -90.0004 | 0.00077292 | 0.00060556 | 8.233e-05 | 2.4757e-06 | 9.6275e-05 | 0.00093678 | 0.005258 | 15.2901 | 2.1226 | 0.01116 | 1 | 115.0005 | 0.00032405 | 100000.9368 | 0 | 0 | 109.231 | 23.8238 | 0.32292
31 | FLUHIGH | 2 | 230 | -90 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0040824 | 6 | 123 |  |  | 229.9969 | 2 | -90 | 1 | 0.00045892 | 0.013853 | 2.08 | 100000.4589 | 2015-01-29 15:46 | 50 | 0 | WP- | 0.34766 | VP+ | 0.18175 | 8.0686 | 1.1331 | 1.1329 | 107.9588 | 12777.4931 | 460.0264 | 0.183 | 7.3856 | 1.1331 | 1.1329 | 107.9368 | 12777.462 | 460.0282 | 0.18459 | 7.4443 | 1.1331 | 1.1332 | 107.9536 | 12777.4079 | 460.0298 | 0.18784 | 7.7658 | 1.1331 | 1.1331 | 107.9243 | 12777.3783 | 460.0288 | 0.18329 | 8.6461 | 1.1331 | 1.1333 | 107.9652 | 12777.4031 | 460.0285 | [7 empty] | 0.0002681 | 459.9853 | 229.9969 | 2 | -90 | 0.0015775 | 0.0024368 | 3.2279e-05 | 1.0544e-05 | 0.00019649 | 0.00045892 | 0.0066601 | 20.4312 | 2.08 | 0.013853 | 1 | 115.0006 | 0.00135 | 100000.4589 | 0 | 0 | 109.241 | 23.8496 | 0.32292
32 | FLUHIGH | 2 | 230 | -90 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0040824 | 6 | 123 |  |  | 229.9969 | 1.9999 | -89.9998 | 1 | 0.0014318 | 0.013811 | 2.0819 | 100001.4318 | 2015-01-29 15:50 | 49.9999 | 0 | WP+ | 0.34766 | VP+ | 0.19147 | 7.3485 | 1.1331 | 1.1332 | 107.9258 | 12777.3512 | 460.028 | 0.18943 | 7.9663 | 1.1331 | 1.1324 | 107.9589 | 12777.3365 | 460.0279 | 0.19046 | 6.8403 | 1.1331 | 1.1332 | 107.9278 | 12777.3262 | 460.0285 | 0.19135 | 7.1802 | 1.1331 | 1.1333 | 107.9338 | 12777.3033 | 460.0255 | 0.19096 | 7.4199 | 1.1331 | 1.1333 | 107.932 | 12777.2368 | 460.0293 | [7 empty] | 0.0017268 | 459.9766 | 229.9969 | 1.9999 | -89.9998 | 0.0011477 | 0.0009831 | 5.755e-05 | 4.5077e-06 | 0.00014296 | 0.0014318 | 0.0066337 | 20.1173 | 2.0819 | 0.013811 | 1 | 115.0027 | 0.00061492 | 100001.4318 | 0 | 0 | 109.248 | 23.8676 | 0.32292
33 | FLUHIGH | 2 | 230 | -30 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0046547 | 6 | 123 |  |  | 229.9967 | 2.0001 | -30.0003 | 0.50001 | -0.0025019 | 0.014834 | 2.0823 | 99997.4981 | 2015-01-29 15:54 | 50 | 0.30566 | WP- | 0.17383 | VP+ | 0.072323 | -5.7336 | 1.1331 | 1.133 | 107.9547 | 6388.9574 | 230.0426 | 0.072833 | -7.9109 | 1.1331 | 1.1329 | 107.9723 | 6388.8933 | 230.0389 | 0.072816 | -4.5402 | 1.1331 | 1.1327 | 107.9783 | 6388.9813 | 230.0413 | 0.073554 | -7.3781 | 1.1331 | 1.1328 | 107.9481 | 6388.866 | 230.0415 | 0.071451 | -8.4849 | 1.1331 | 1.1332 | 107.9782 | 6388.9041 | 230.0382 | [7 empty] | 398.3782 | 230.0069 | 229.9967 | 2.0001 | -30.0003 | 0.0020394 | 0.00078644 | 9.1489e-05 | 8.0086e-06 | 0.00015137 | -0.0025019 | 0.0071238 | 20.0624 | 2.0823 | 0.014834 | 0.50001 | 114.9945 | 0.0010535 | 99997.4981 | 0 | 0 | 109.255 | 23.8857 | 0.32291
34 | FLUHIGH | 2 | 230 | -30 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0046547 | 6 | 123 |  |  | 229.9968 | 2.0001 | -30.0001 | 0.5 | -0.0037709 | 0.014875 | 2.0806 | 99996.2291 | 2015-01-29 15:58 | 50.0004 | 0.30566 | WP- | 0.17383 | VP+ | 0.072845 | -5.9179 | 1.1331 | 1.1331 | 107.9601 | 6388.8608 | 230.0383 | 0.068112 | -6.562 | 1.1331 | 1.1329 | 107.9679 | 6388.789 | 230.0417 | 0.074127 | -5.0964 | 1.1331 | 1.1328 | 107.9508 | 6388.8582 | 230.0432 | 0.071638 | -6.1601 | 1.1331 | 1.1333 | 107.9699 | 6388.8391 | 230.0402 | 0.073168 | -5.9004 | 1.1331 | 1.1335 | 107.9474 | 6388.8066 | 230.0357 | [7 empty] | 398.3808 | 230.0066 | 229.9968 | 2.0001 | -30.0001 | 0.00086323 | 0.002257 | 8.1084e-05 | 8.0157e-06 | 0.00019915 | -0.0037709 | 0.0071495 | 20.3375 | 2.0806 | 0.014875 | 0.5 | 114.994 | 0.0010558 | 99996.2291 | 0 | 0 | 109.261 | 23.9011 | 0.32291
35 | FLUHIGH | 2 | 230 | -30 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0046547 | 6 | 123 |  |  | 229.9968 | 2.0001 | -30 | 0.5 | -0.0027283 | 0.014872 | 2.0807 | 99997.2717 | 2015-01-29 16:02 | 50 | 0.30566 | WP- | 0.17383 | VP+ | 0.070291 | -5.8136 | 1.1331 | 1.1331 | 107.9536 | 6388.9282 | 230.0404 | 0.072839 | -5.8606 | 1.1331 | 1.1327 | 107.941 | 6388.8371 | 230.0408 | 0.073141 | -4.1399 | 1.1331 | 1.1326 | 107.9412 | 6388.9362 | 230.0441 | 0.074096 | -5.7401 | 1.1331 | 1.1328 | 107.9503 | 6388.8714 | 230.0384 | 0.073057 | -5.5051 | 1.1331 | 1.1329 | 107.9557 | 6388.8878 | 230.0397 | [7 empty] | 398.3826 | 230.0064 | 229.9968 | 2.0001 | -30 | 0.00096415 | 0.0014863 | 4.3394e-05 | 5.6161e-06 | 0.0001463 | -0.0027283 | 0.0071476 | 20.3181 | 2.0807 | 0.014872 | 0.5 | 114.9937 | 0.00070554 | 99997.2717 | 0 | 0 | 109.264 | 23.9089 | 0.32291
36 | FLUHIGH | 5 | 230 | -90 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0040824 | 6 | 123 |  |  | 229.9967 | 4.9999 | -89.9997 | 1 | 1.4967e-05 | 0.013797 | 2.0829 | 100000.015 | 2015-01-29 16:07 | 50.0003 | 0 | WP- | 0.87012 | VP+ | 0.19491 | 118.3763 | 1.1331 | 1.1331 | 107.9556 | 31943.3847 | 1150.0704 | 0.20046 | 118.1655 | 1.1331 | 1.133 | 107.9495 | 31943.426 | 1150.0728 | 0.1996 | 118.1434 | 1.1331 | 1.1332 | 107.952 | 31943.3407 | 1150.0677 | 0.19497 | 117.8735 | 1.1331 | 1.1332 | 107.9538 | 31943.4827 | 1150.071 | 0.19691 | 117.1728 | 1.1331 | 1.1329 | 107.9488 | 31943.3433 | 1150.0629 | [7 empty] | 0.0060449 | 1149.9621 | 229.9967 | 4.9999 | -89.9997 | 0.0021435 | 0.0023084 | 6.3581e-05 | 9.7483e-06 | 0.0001068 | 1.4967e-05 | 0.0066237 | 19.9611 | 2.0829 | 0.013797 | 1 | 46.0002 | 0.00019854 | 100000.015 | 0 | 0 | 109.262 | 23.9037 | 0.32291
37 | FLUHIGH | 5 | 230 | -90 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0040824 | 6 | 123 |  |  | 229.9967 | 4.9999 | -89.9997 | 1 | -0.00097383 | 0.013795 | 2.083 | 99999.0262 | 2015-01-29 16:11 | 50.0003 | 0 | WP- | 0.87012 | VP+ | 0.19762 | 116.8637 | 1.1331 | 1.1337 | 107.9281 | 31943.2483 | 1150.072 | 0.19644 | 118.0598 | 1.1331 | 1.1329 | 107.9485 | 31943.2905 | 1150.0649 | 0.19986 | 118.0072 | 1.1331 | 1.1331 | 107.9505 | 31943.1761 | 1150.0665 | 0.19704 | 117.9759 | 1.1331 | 1.1327 | 107.9637 | 31943.1771 | 1150.0651 | 0.19476 | 117.9506 | 1.1331 | 1.1336 | 107.9347 | 31943.1785 | 1150.0701 | [7 empty] | 0.0064886 | 1149.9669 | 229.9967 | 4.9999 | -89.9997 | 0.0019323 | 0.0021048 | 6.7078e-05 | 8.1859e-06 | 9.6275e-05 | -0.00097383 | 0.0066227 | 19.9497 | 2.083 | 0.013795 | 1 | 46 | 0.00015189 | 99999.0262 | 0 | 0 | 109.258 | 23.8934 | 0.32291
38 | FLUHIGH | 5 | 230 | -90 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0040824 | 6 | 123 |  |  | 229.9969 | 4.9999 | -89.9998 | 1 | 0.00035219 | 0.013797 | 2.0829 | 100000.3522 | 2015-01-29 16:15 | 49.9998 | 0 | WP- | 0.87012 | VP+ | 0.19938 | 118.504 | 1.1331 | 1.1332 | 107.9579 | 31943.3681 | 1150.067 | 0.19573 | 118.2578 | 1.1331 | 1.1332 | 107.973 | 31943.4314 | 1150.0745 | 0.19546 | 118.3703 | 1.1331 | 1.1334 | 107.9324 | 31943.3016 | 1150.0688 | 0.19939 | 117.0667 | 1.1331 | 1.1329 | 107.922 | 31943.2862 | 1150.0718 | 0.19346 | 118.2071 | 1.1331 | 1.1331 | 107.9761 | 31943.3359 | 1150.0687 | [7 empty] | 0.0041793 | 1149.9564 | 229.9969 | 4.9999 | -89.9998 | 0.0028035 | 0.0019132 | 8.4476e-05 | 7.8026e-06 | 0.00013969 | 0.00035219 | 0.0066242 | 19.9673 | 2.0829 | 0.013797 | 1 | 46.0005 | 0.00015844 | 100000.3522 | 0 | 0 | 109.257 | 23.8908 | 0.32291
39 | FLUHIGH | 5 | 230 | -30 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0046547 | 6 | 123 |  |  | 229.9969 | 5.0002 | -29.9999 | 0.5 | -0.00084606 | 0.014777 | 2.0849 | 99999.1539 | 2015-01-29 16:19 | 50.0003 | 0.76465 | WP- | 0.43457 | VP+ | 0.32077 | -66.2902 | 1.1331 | 1.1332 | 107.9294 | 15972.6836 | 575.0906 | 0.32045 | -66.1495 | 1.1331 | 1.1327 | 107.9372 | 15972.7164 | 575.0839 | 0.31926 | -65.8429 | 1.1331 | 1.1338 | 107.9443 | 15972.5761 | 575.0889 | 0.3229 | -65.9517 | 1.1331 | 1.1331 | 107.9459 | 15972.5079 | 575.0906 | 0.32074 | -65.8062 | 1.1331 | 1.1332 | 107.934 | 15972.4475 | 575.0867 | [7 empty] | 995.9633 | 575.018 | 229.9969 | 5.0002 | -29.9999 | 0.0018336 | 0.0015654 | 9.4247e-05 | 7.0617e-06 | 6.3062e-05 | -0.00084606 | 0.0070874 | 19.6549 | 2.0849 | 0.014777 | 0.5 | 45.9972 | 0.00010989 | 99999.1539 | 0 | 0 | 109.256 | 23.8883 | 0.32291
40 | FLUHIGH | 5 | 230 | -30 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0046547 | 6 | 123 |  |  | 229.9968 | 5.0002 | -29.9999 | 0.5 | -0.00064106 | 0.014758 | 2.0858 | 99999.3589 | 2015-01-29 16:23 | 50.0003 | 0.76465 | WP- | 0.43457 | VP+ | 0.32275 | -66.5368 | 1.1331 | 1.133 | 107.9219 | 15972.5614 | 575.0894 | 0.31913 | -66.5754 | 1.1331 | 1.1331 | 107.9444 | 15972.5307 | 575.0884 | 0.321 | -66.3828 | 1.1331 | 1.1331 | 107.9649 | 15972.5731 | 575.0883 | 0.32196 | -66.1578 | 1.1331 | 1.1337 | 107.9621 | 15972.5201 | 575.0905 | 0.32067 | -65.94 | 1.1331 | 1.1327 | 107.9524 | 15972.4258 | 575.0905 | [7 empty] | 995.9595 | 575.0145 | 229.9968 | 5.0002 | -29.9999 | 0.0018837 | 0.0015976 | 8.4545e-05 | 7.2572e-06 | 7.2989e-05 | -0.00064106 | 0.0070756 | 19.527 | 2.0858 | 0.014758 | 0.5 | 45.9974 | 0.00012568 | 99999.3589 | 0 | 0 | 109.257 | 23.8908 | 0.32291
41 | FLUHIGH | 5 | 230 | -30 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0046547 | 6 | 123 |  |  | 229.9968 | 5.0002 | -29.9997 | 0.49999 | -0.00011599 | 0.01476 | 2.0857 | 99999.884 | 2015-01-29 16:28 | 50 | 0.76465 | WP- | 0.43457 | VP+ | 0.32319 | -66.0212 | 1.1331 | 1.1328 | 107.9627 | 15972.5626 | 575.092 | 0.31924 | -66.3595 | 1.1331 | 1.1329 | 107.9384 | 15972.6208 | 575.092 | 0.32057 | -66.5554 | 1.1331 | 1.1327 | 107.9497 | 15972.5611 | 575.0928 | 0.31837 | -65.9096 | 1.1331 | 1.1331 | 107.9557 | 15972.4623 | 575.0839 | 0.3194 | -65.8342 | 1.1331 | 1.1328 | 107.9159 | 15972.4284 | 575.0932 | [7 empty] | 995.9629 | 575.0116 | 229.9968 | 5.0002 | -29.9997 | 0.0022742 | 0.0015266 | 6.7251e-05 | 8.5032e-06 | 8.9151e-05 | -0.00011599 | 0.0070768 | 19.5405 | 2.0857 | 0.01476 | 0.49999 | 45.9973 | 0.0001691 | 99999.884 | 0 | 0 | 109.256 | 23.8883 | 0.32291
42 | FLUHIGH | 2 | 230 | -126.87 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0042105 | 6 | 123 |  |  | 229.9969 | 2 | -126.8733 | 0.79996 | -0.0014387 | 0.014047 | 2.081 | 99998.5614 | 2015-01-29 16:32 | 50.0002 | 0.21191 | WP+ | 0.27832 | VP+ | 0.11119 | 58.2659 | 1.1331 | 1.1328 | 107.936 | 10221.27 | 368.0111 | 0.11498 | 57.5308 | 1.1331 | 1.1334 | 107.9463 | 10221.2528 | 368.0021 | 0.11266 | 59.0411 | 1.1331 | 1.133 | 107.9552 | 10221.2483 | 368.005 | 0.11329 | 57.6162 | 1.1331 | 1.1328 | 107.9452 | 10221.265 | 368.0023 | 0.11495 | 57.5942 | 1.1331 | 1.133 | 107.9585 | 10221.2258 | 368.0036 | [7 empty] | -276.0117 | 367.9704 | 229.9969 | 2 | -126.8733 | 0.0012591 | 0.0016682 | 1.1819e-05 | 4.3736e-06 | 0.00022824 | -0.0014387 | 0.0067503 | 20.2616 | 2.081 | 0.014047 | 0.79996 | 115.0011 | 0.00056304 | 99998.5614 | 0 | 0 | 109.26 | 23.8986 | 0.32291
43 | FLUHIGH | 2 | 230 | -126.87 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0042105 | 6 | 123 |  |  | 229.9969 | 2 | -126.8732 | 0.79997 | -0.0019611 | 0.014059 | 2.0805 | 99998.0388 | 2015-01-29 16:36 | 50.0001 | 0.21191 | WP+ | 0.27832 | VP+ | 0.11276 | 58.1548 | 1.1331 | 1.1331 | 107.9601 | 10221.2052 | 368.0025 | 0.11272 | 58.4557 | 1.1331 | 1.1334 | 107.9202 | 10221.2016 | 368.0052 | 0.11482 | 58.0041 | 1.1331 | 1.133 | 107.9414 | 10221.1775 | 368.0045 | 0.11458 | 56.3514 | 1.1331 | 1.1328 | 107.9396 | 10221.1867 | 368.0074 | 0.11435 | 58.0505 | 1.1331 | 1.1331 | 107.9487 | 10221.2027 | 368.0063 | [7 empty] | -276.011 | 367.9702 | 229.9969 | 2 | -126.8732 | 0.00086998 | 0.0020044 | 8.7206e-05 | 7.0089e-06 | 0.0001631 | -0.0019611 | 0.0067574 | 20.3454 | 2.0805 | 0.014059 | 0.79997 | 115.0012 | 0.00082881 | 99998.0388 | 0 | 0 | 109.265 | 23.9114 | 0.32291
44 | FLUHIGH | 2 | 230 | -126.87 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0042105 | 6 | 123 |  |  | 229.9969 | 2 | -126.8734 | 0.79996 | -0.001348 | 0.014025 | 2.0821 | 99998.652 | 2015-01-29 16:40 | 50.0003 | 0.21191 | WP+ | 0.27832 | VP+ | 0.11169 | 57.6914 | 1.1331 | 1.133 | 107.9309 | 10221.3294 | 368.0065 | 0.11474 | 57.9815 | 1.1331 | 1.1331 | 107.9284 | 10221.2558 | 368.0038 | 0.11137 | 57.3959 | 1.1331 | 1.1333 | 107.9508 | 10221.2876 | 368.0035 | 0.11505 | 57.7844 | 1.1331 | 1.1329 | 107.9598 | 10221.2528 | 368.0035 | 0.11135 | 58.1518 | 1.1331 | 1.1334 | 107.9422 | 10221.2151 | 368.0061 | [7 empty] | -276.013 | 367.9706 | 229.9969 | 2 | -126.8734 | 0.0019058 | 0.00069967 | 3.2815e-05 | 3.8575e-06 | 0.00022528 | -0.001348 | 0.0067362 | 20.0963 | 2.0821 | 0.014025 | 0.79996 | 115.0008 | 0.00047839 | 99998.652 | 0 | 0 | 109.266 | 23.914 | 0.32291
45 | FLUHIGH | 5 | 230 | -126.87 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0042105 | 6 | 123 |  |  | 229.9968 | 4.9999 | -126.8724 | 0.79997 | -7.4262e-05 | 0.014004 | 2.0833 | 99999.9257 | 2015-01-29 16:44 | 50.0004 | 0.5293 | WP+ | 0.69629 | VP+ | 0.18731 | 171.9707 | 1.1331 | 1.1333 | 107.9274 | 25553.7873 | 920.017 | 0.18633 | 173.0967 | 1.1331 | 1.133 | 107.9404 | 25554.0606 | 920.0182 | 0.18609 | 172.8052 | 1.1331 | 1.1334 | 107.9657 | 25553.906 | 920.0207 | 0.18546 | 171.8218 | 1.1331 | 1.1328 | 107.9479 | 25553.8356 | 920.0167 | 0.18605 | 173.5951 | 1.1331 | 1.1322 | 107.9672 | 25553.925 | 920.0205 | [7 empty] | -690.0196 | 919.9412 | 229.9968 | 4.9999 | -126.8724 | 0.0024413 | 0.00071071 | 7.1996e-05 | 7.4217e-06 | 9.223e-05 | -7.4262e-05 | 0.0067219 | 19.8974 | 2.0833 | 0.014004 | 0.79997 | 46.0001 | 0.00015617 | 99999.9257 | 0 | 0 | 109.266 | 23.914 | 0.32291
46 | FLUHIGH | 5 | 230 | -126.87 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0042105 | 6 | 123 |  |  | 229.9966 | 4.9999 | -126.8724 | 0.79997 | -0.00020332 | 0.014003 | 2.0834 | 99999.7967 | 2015-01-29 16:49 | 50.0003 | 0.5293 | WP+ | 0.69629 | VP+ | 0.18403 | 171.6326 | 1.1331 | 1.1327 | 107.9368 | 25553.7932 | 920.0158 | 0.18642 | 172.6003 | 1.1331 | 1.1329 | 107.9645 | 25553.6685 | 920.0196 | 0.18509 | 172.4772 | 1.1331 | 1.1328 | 107.9643 | 25553.8422 | 920.0181 | 0.18502 | 172.2145 | 1.1331 | 1.1329 | 107.9413 | 25553.8569 | 920.0195 | 0.18886 | 172.6816 | 1.1331 | 1.1323 | 107.9619 | 25553.7883 | 920.0159 | [7 empty] | -690.0161 | 919.9383 | 229.9966 | 4.9999 | -126.8724 | 0.0015048 | 0.0016355 | 5.4458e-05 | 3.1457e-06 | 0.00010675 | -0.00020332 | 0.0067216 | 19.8937 | 2.0834 | 0.014003 | 0.79997 | 46.0002 | 8.2323e-05 | 99999.7967 | 0 | 0 | 109.267 | 23.9166 | 0.32291
47 | FLUHIGH | 5 | 230 | -126.87 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0042105 | 6 | 123 |  |  | 229.9969 | 4.9999 | -126.8724 | 0.79997 | -0.00068697 | 0.014002 | 2.0834 | 99999.313 | 2015-01-29 16:53 | 50.0002 | 0.5293 | WP+ | 0.69629 | VP+ | 0.18405 | 172.6163 | 1.1331 | 1.133 | 107.9333 | 25553.7437 | 920.0133 | 0.18491 | 172.6256 | 1.1331 | 1.1327 | 107.9581 | 25553.6856 | 920.0138 | 0.18744 | 172.504 | 1.1331 | 1.1334 | 107.9227 | 25553.7511 | 920.0161 | 0.18499 | 172.5882 | 1.1331 | 1.1327 | 107.9563 | 25553.751 | 920.0167 | 0.18604 | 171.7081 | 1.1331 | 1.1333 | 107.9421 | 25553.7578 | 920.0168 | [7 empty] | -690.0183 | 919.9409 | 229.9969 | 4.9999 | -126.8724 | 0.00122 | 0.0014076 | 3.4586e-05 | 6.5272e-06 | 5.3698e-05 | -0.00068697 | 0.006721 | 19.8866 | 2.0834 | 0.014002 | 0.79997 | 46.0002 | 0.00013443 | 99999.313 | 0 | 0 | 109.266 | 23.914 | 0.32291
48 | FLUHIGH | 0.5 | 230 | 0 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9964 | 0.50002 | -0.0016197 | 1 | 0.002022 | 0.0050581 | 2.0614 | 100002.022 | 2015-01-29 17:14 | 50 | 0.44141 | WP- | 0 | VP+ | 0.15683 | -112.6822 | 1.1331 | 1.133 | 107.9737 | 3194.5662 | 115.0062 | 0.15429 | -112.5438 | 1.1331 | 1.1331 | 107.9635 | 3194.4945 | 115.0065 | 0.15765 | -110.8351 | 1.1331 | 1.1334 | 107.9551 | 3194.5847 | 115.0062 | 0.15973 | -111.3165 | 1.1331 | 1.1328 | 107.9506 | 3194.6148 | 115.0054 | 0.15481 | -111.1624 | 1.1331 | 1.1327 | 107.9414 | 3194.5535 | 115.0059 | [7 empty] | 115.0019 | 0.0032509 | 229.9964 | 0.50002 | -0.0016197 | 0.00039031 | 0.00044298 | 0.00015167 | 1.8742e-06 | 0.0002207 | 0.002022 | 0.0024537 | 24.0335 | 2.0614 | 0.0050581 | 1 | 459.978 | 0.0042966 | 100002.022 | 0 | 0 | 109.286 | 23.9656 | 0.32291
49 | FLUHIGH | 0.5 | 230 | 0 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9966 | 0.50001 | -0.001992 | 1 | 0.0018318 | 0.004974 | 2.0631 | 100001.8318 | 2015-01-29 17:18 | 50 | 0.44141 | WP- | 0 | VP+ | 0.15728 | -112.0614 | 1.1331 | 1.1329 | 107.9659 | 3194.5306 | 115.0057 | 0.16042 | -110.0765 | 1.1331 | 1.133 | 107.9499 | 3194.5882 | 115.0052 | 0.15643 | -110.0365 | 1.1331 | 1.1331 | 107.9369 | 3194.5129 | 115.0059 | 0.16048 | -111.2589 | 1.1331 | 1.1337 | 107.9508 | 3194.5457 | 115.0058 | 0.15479 | -111.7022 | 1.1331 | 1.1329 | 107.9644 | 3194.5326 | 115.0056 | [7 empty] | 115.0014 | 0.0039982 | 229.9966 | 0.50001 | -0.001992 | 0.00052652 | 0.00049776 | 8.194e-05 | 2.309e-06 | 0.000248 | 0.0018318 | 0.0024109 | 23.6518 | 2.0631 | 0.004974 | 1 | 459.9808 | 0.0048181 | 100001.8318 | 0 | 0 | 109.295 | 23.9888 | 0.32291
50 | FLUHIGH | 0.5 | 230 | 0 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9967 | 0.50001 | -0.0017606 | 1 | 0.0024727 | 0.0047466 | 2.079 | 100002.4727 | 2015-01-29 17:23 | 50.0002 | 0.44141 | WP- | 0 | VP+ | 0.1583 | -112.0648 | 1.1331 | 1.1329 | 107.9548 | 3194.5896 | 115.0057 | 0.15846 | -109.8901 | 1.1331 | 1.1329 | 107.9422 | 3194.5409 | 115.0056 | 0.15783 | -112.0334 | 1.1331 | 1.133 | 107.9634 | 3194.566 | 115.0057 | 0.15661 | -111.935 | 1.1331 | 1.133 | 107.9417 | 3194.572 | 115.0057 | 0.16061 | -110.9555 | 1.1331 | 1.1335 | 107.9485 | 3194.5317 | 115.0053 | [7 empty] | 115.0013 | 0.0035338 | 229.9967 | 0.50001 | -0.0017606 | 0.00037568 | 0.00050854 | 0.0001023 | 1.5424e-06 | 0.00025337 | 0.0024727 | 0.0022831 | 20.5876 | 2.079 | 0.0047466 | 1 | 459.9814 | 0.0030436 | 100002.4727 | 0 | 0 | 109.3 | 24.0016 | 0.32291
51 | FLUHIGH | 0.5 | 230 | -60 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9966 | 0.49999 | -59.9995 | 0.50001 | -0.0057521 | 0.0089687 | 2.0397 | 99994.2479 | 2015-01-29 17:27 | 50.0004 | 0.2207 | WP- | 0.37695 | VP+ | 0.17243 | 173.2139 | 1.1331 | 1.1333 | 107.9487 | 1597.1536 | 57.4955 | 0.17334 | 172.9302 | 1.1331 | 1.1328 | 107.9627 | 1597.0579 | 57.4954 | 0.17023 | 171.8288 | 1.1331 | 1.1329 | 107.9955 | 1597.1021 | 57.4951 | 0.16788 | 173.7577 | 1.1331 | 1.1335 | 107.9599 | 1597.1276 | 57.4949 | 0.17323 | 173.009 | 1.1331 | 1.1329 | 107.9628 | 1597.1313 | 57.4958 | [7 empty] | 57.4994 | 99.5898 | 229.9966 | 0.49999 | -59.9995 | 0.0004341 | 0.00047147 | 0.00011892 | 1.891e-06 | 0.00023407 | -0.0057521 | 0.004397 | 30.3757 | 2.0397 | 0.0089687 | 0.50001 | 459.9984 | 0.0039536 | 99994.2479 | 0 | 0 | 109.299 | 23.9991 | 0.32291
52 | FLUHIGH | 0.5 | 230 | -60 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9968 | 0.5 | -59.9992 | 0.50001 | -0.0063143 | 0.0088894 | 2.0406 | 99993.6857 | 2015-01-29 17:31 | 50.0002 | 0.2207 | WP- | 0.37695 | VP+ | 0.16668 | 172.664 | 1.1331 | 1.1334 | 107.9288 | 1597.192 | 57.4945 | 0.17082 | 171.7273 | 1.1331 | 1.133 | 107.9392 | 1597.1362 | 57.4954 | 0.16659 | 171.5637 | 1.1331 | 1.133 | 107.9477 | 1597.1171 | 57.4946 | 0.17122 | 171.3789 | 1.1331 | 1.1328 | 107.9516 | 1597.1423 | 57.4955 | 0.16994 | 173.7917 | 1.1331 | 1.133 | 107.9557 | 1597.0951 | 57.495 | [7 empty] | 57.5005 | 99.5907 | 229.9968 | 0.5 | -59.9992 | 0.00062782 | 0.00040538 | 0.00010159 | 1.9792e-06 | 0.00029754 | -0.0063143 | 0.0043563 | 30.0622 | 2.0406 | 0.0088894 | 0.50001 | 459.994 | 0.0041738 | 99993.6857 | 0 | 0 | 109.293 | 23.9836 | 0.32291
53 | FLUHIGH | 0.5 | 230 | -60 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9968 | 0.5 | -59.9994 | 0.50001 | -0.0068845 | 0.0086729 | 2.0456 | 99993.1155 | 2015-01-29 17:35 | 50.0002 | 0.2207 | WP- | 0.37695 | VP+ | 0.16817 | 171.9981 | 1.1331 | 1.133 | 107.9662 | 1597.124 | 57.495 | 0.16981 | 173.6116 | 1.1331 | 1.1329 | 107.9765 | 1597.1002 | 57.4948 | 0.17031 | 171.65 | 1.1331 | 1.133 | 107.9431 | 1597.1502 | 57.4957 | 0.16865 | 173.389 | 1.1331 | 1.1332 | 107.9366 | 1597.1149 | 57.4951 | 0.1715 | 171.7082 | 1.1331 | 1.1328 | 107.9537 | 1597.1028 | 57.4955 | [7 empty] | 57.5002 | 99.5907 | 229.9968 | 0.5 | -59.9994 | 0.00056196 | 0.00031341 | 4.1683e-05 | 1.226e-06 | 0.00028474 | -0.0068845 | 0.0042398 | 28.344 | 2.0456 | 0.0086729 | 0.50001 | 459.9945 | 0.0024081 | 99993.1155 | 0 | 0 | 109.285 | 23.963 | 0.32291
54 | FLUHIGH | 0.5 | 230 | 36.87 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9968 | 0.49999 | 36.8966 | 0.79972 | 0.0051806 | 0.0057979 | 2.0488 | 100005.1806 | 2015-01-29 17:39 | 50.0003 | 0.35254 | WP- | 0.26074 | VP- | 0.22772 | 146.4289 | 1.1331 | 1.1329 | 107.9429 | 2554.7246 | 91.9734 | 0.22581 | 147.4074 | 1.1331 | 1.1331 | 107.9709 | 2554.764 | 91.9731 | 0.22673 | 147.7855 | 1.1331 | 1.1336 | 107.9643 | 2554.7147 | 91.9735 | 0.22637 | 146.7299 | 1.1331 | 1.1333 | 107.9264 | 2554.7007 | 91.9726 | 0.22872 | 146.8737 | 1.1331 | 1.1336 | 107.9549 | 2554.7316 | 91.9733 | [7 empty] | 91.9654 | 69.041 | 229.9968 | 0.49999 | 36.8966 | 0.00045829 | 0.00017177 | 9.2706e-05 | 1.5778e-06 | 0.00016172 | 0.0051806 | 0.00283 | 27.3546 | 2.0488 | 0.0057979 | 0.79972 | 459.9993 | 0.0032678 | 100005.1806 | 0 | 0 | 109.284 | 23.9604 | 0.32291
55 | FLUHIGH | 0.5 | 230 | 36.87 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9967 | 0.5 | 36.8967 | 0.79972 | 0.0051181 | 0.0058418 | 2.0471 | 100005.1181 | 2015-01-29 17:44 | 50.0002 | 0.35254 | WP- | 0.26074 | VP- | 0.23045 | 146.6995 | 1.1331 | 1.1336 | 107.9564 | 2554.7054 | 91.9736 | 0.22645 | 147.6668 | 1.1331 | 1.1327 | 107.9563 | 2554.7715 | 91.9731 | 0.23014 | 146.8341 | 1.1331 | 1.1328 | 107.9696 | 2554.7366 | 91.9735 | 0.23046 | 146.9004 | 1.1331 | 1.1328 | 107.963 | 2554.7553 | 91.9733 | 0.22819 | 146.4924 | 1.1331 | 1.1333 | 107.9334 | 2554.7355 | 91.9733 | [7 empty] | 91.966 | 69.0416 | 229.9967 | 0.5 | 36.8967 | 0.00035576 | 0.00025872 | 0.00010187 | 1.7949e-06 | 0.00011445 | 0.0051181 | 0.0028537 | 27.851 | 2.0471 | 0.0058418 | 0.79972 | 459.9959 | 0.0040528 | 100005.1181 | 0 | 0 | 109.281 | 23.9527 | 0.32291
56 | FLUHIGH | 0.5 | 230 | 36.87 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.997 | 0.49999 | 36.8969 | 0.79972 | 0.0059269 | 0.0057596 | 2.0505 | 100005.9269 | 2015-01-29 17:48 | 50.0002 | 0.35254 | WP- | 0.26074 | VP- | 0.22665 | 147.6487 | 1.1331 | 1.1329 | 107.938 | 2554.7494 | 91.9728 | 0.22737 | 146.7029 | 1.1331 | 1.1327 | 107.9452 | 2554.7683 | 91.973 | 0.22969 | 147.5082 | 1.1331 | 1.1332 | 107.9381 | 2554.7796 | 91.9733 | 0.23083 | 146.9981 | 1.1331 | 1.1331 | 107.9353 | 2554.7269 | 91.9739 | 0.2262 | 147.7778 | 1.1331 | 1.1328 | 107.9455 | 2554.6939 | 91.9728 | [7 empty] | 91.9653 | 69.0417 | 229.997 | 0.49999 | 36.8969 | 0.00043208 | 0.00035271 | 4.0979e-05 | 2.0184e-06 | 0.00015102 | 0.0059269 | 0.0028088 | 26.8291 | 2.0505 | 0.0057596 | 0.79972 | 459.999 | 0.0041245 | 100005.9269 | 0 | 0 | 109.281 | 23.9527 | 0.32291
57 | FLUHIGH | 1 | 230 | 0 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9969 | 1 | -0.0013434 | 1 | 0.0030607 | 0.004757 | 2.0781 | 100003.0607 | 2015-01-29 17:52 | 50.0006 | 0.88281 | WP- | 0 | VP+ | 0.32395 | -111.013 | 1.1331 | 1.1329 | 107.9468 | 6389.1168 | 230.0097 | 0.3195 | -111.2515 | 1.1331 | 1.1329 | 107.9424 | 6389.1697 | 230.0091 | 0.32486 | -111.1047 | 1.1331 | 1.1332 | 107.9609 | 6389.187 | 230.0099 | 0.3163 | -110.9109 | 1.1331 | 1.1333 | 107.9507 | 6389.1575 | 230.0099 | 0.31901 | -110.9272 | 1.1331 | 1.1328 | 107.9647 | 6389.1032 | 230.0099 | [7 empty] | 230.0022 | 0.0053928 | 229.9969 | 1 | -0.0013434 | 0.00067333 | 0.00022515 | 4.1823e-05 | 3.0059e-06 | 5.6082e-05 | 0.0030607 | 0.0022892 | 20.7506 | 2.0781 | 0.004757 | 1 | 229.9916 | 0.0015906 | 100003.0607 | 0 | 0 | 109.281 | 23.9527 | 0.32291
58 | FLUHIGH | 1 | 230 | 0 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9968 | 1 | -0.0012171 | 1 | 0.0021322 | 0.0047163 | 2.0829 | 100002.1322 | 2015-01-29 17:56 | 49.9997 | 0.88281 | WP- | 0 | VP+ | 0.31949 | -111.016 | 1.1331 | 1.133 | 107.939 | 6389.1115 | 230.0094 | 0.31905 | -111.7954 | 1.1331 | 1.133 | 107.9211 | 6389.0917 | 230.0088 | 0.31822 | -110.9093 | 1.1331 | 1.1327 | 107.9651 | 6389.0649 | 230.0092 | 0.32245 | -111.3434 | 1.1331 | 1.1324 | 107.9491 | 6389.0666 | 230.0101 | 0.32017 | -111.3121 | 1.1331 | 1.1329 | 107.9607 | 6389.1243 | 230.0095 | [7 empty] | 230.0024 | 0.0048857 | 229.9968 | 1 | -0.0012171 | 0.0002529 | 0.00038738 | 8.0272e-05 | 1.1893e-06 | 9.65e-05 | 0.0021322 | 0.0022643 | 19.9633 | 2.0829 | 0.0047163 | 1 | 229.9911 | 0.00070195 | 100002.1322 | 0 | 0 | 109.281 | 23.9527 | 0.32291
59 | FLUHIGH | 1 | 230 | 0 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9969 | 1 | -0.0012966 | 1 | 0.0025232 | 0.0047128 | 2.0834 | 100002.5232 | 2015-01-29 18:00 | 50.0003 | 0.88281 | WP- | 0 | VP+ | 0.31972 | -111.5375 | 1.1331 | 1.133 | 107.9578 | 6389.1369 | 230.0096 | 0.32029 | -111.3112 | 1.1331 | 1.1328 | 107.9539 | 6389.0957 | 230.0093 | 0.31917 | -111.5521 | 1.1331 | 1.1326 | 107.9554 | 6389.1331 | 230.01 | 0.32387 | -110.7184 | 1.1331 | 1.133 | 107.9356 | 6389.0219 | 230.0099 | 0.32098 | -110.4922 | 1.1331 | 1.1326 | 107.9315 | 6389.1532 | 230.0104 | [7 empty] | 230.0021 | 0.005205 | 229.9969 | 1 | -0.0012966 | 0.00038604 | 0.00051856 | 7.0422e-05 | 1.7777e-06 | 0.00012918 | 0.0025232 | 0.0022621 | 19.8902 | 2.0834 | 0.0047128 | 1 | 229.9916 | 0.00098801 | 100002.5232 | 0 | 0 | 109.277 | 23.9424 | 0.32291
60 | FLUHIGH | 1 | 230 | -60 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9969 | 0.99999 | -59.9985 | 0.50002 | -0.0057302 | 0.0086367 | 2.0474 | 99994.2698 | 2015-01-29 18:04 | 49.9999 | 0.44141 | WP- | 0.75293 | VP+ | 0.22797 | -16.847 | 1.1331 | 1.1333 | 107.9598 | 3194.3003 | 114.992 | 0.2301 | -17.9546 | 1.1331 | 1.1332 | 107.9428 | 3194.2014 | 114.9909 | 0.2279 | -16.5117 | 1.1331 | 1.1334 | 107.9437 | 3194.4105 | 114.9916 | 0.22487 | -17.5124 | 1.1331 | 1.1332 | 107.9673 | 3194.3987 | 114.9918 | 0.22778 | -17.3601 | 1.1331 | 1.1339 | 107.9664 | 3194.3034 | 114.9909 | [7 empty] | 115.0022 | 199.1778 | 229.9969 | 0.99999 | -59.9985 | 0.00051797 | 0.00031506 | 6.9213e-05 | 1.6375e-06 | 0.00011538 | -0.0057302 | 0.0042183 | 27.7643 | 2.0474 | 0.0086367 | 0.50002 | 229.9997 | 0.00083724 | 99994.2698 | 0 | 0 | 109.274 | 23.9346 | 0.32291
61 | FLUHIGH | 1 | 230 | -60 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.997 | 1 | -59.9991 | 0.50001 | -0.0057348 | 0.0085931 | 2.0489 | 99994.2652 | 2015-01-29 18:09 | 50.0007 | 0.44141 | WP- | 0.75293 | VP+ | 0.22829 | -17.9854 | 1.1331 | 1.1332 | 107.9406 | 3194.219 | 114.9911 | 0.22678 | -16.9622 | 1.1331 | 1.1334 | 107.9601 | 3194.2815 | 114.9915 | 0.22426 | -18.0865 | 1.1331 | 1.133 | 107.949 | 3194.3337 | 114.9917 | 0.22849 | -18.2438 | 1.1331 | 1.1337 | 107.9432 | 3194.2835 | 114.991 | 0.22871 | -18.1888 | 1.1331 | 1.1326 | 107.9545 | 3194.3479 | 114.9913 | [7 empty] | 115.0011 | 199.1809 | 229.997 | 1 | -59.9991 | 0.00036376 | 0.00041378 | 0.00012495 | 1.3598e-06 | 0.00011129 | -0.0057348 | 0.0041939 | 27.3032 | 2.0489 | 0.0085931 | 0.50001 | 229.9977 | 0.00078383 | 99994.2652 | 0 | 0 | 109.267 | 23.9166 | 0.32291
62 | FLUHIGH | 1 | 230 | -60 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9969 | 0.99999 | -59.9988 | 0.50002 | -0.005391 | 0.0086049 | 2.0485 | 99994.609 | 2015-01-29 18:13 | 50.0002 | 0.44141 | WP- | 0.75293 | VP+ | 0.22279 | -17.6498 | 1.1331 | 1.133 | 107.9454 | 3194.2766 | 114.9915 | 0.22863 | -18.3177 | 1.1331 | 1.1328 | 107.9636 | 3194.3807 | 114.9914 | 0.22791 | -17.147 | 1.1331 | 1.1338 | 107.9277 | 3194.31 | 114.9913 | 0.22358 | -17.6673 | 1.1331 | 1.1333 | 107.9496 | 3194.2945 | 114.9914 | 0.22782 | -18.3767 | 1.1331 | 1.1326 | 107.9509 | 3194.2698 | 114.9909 | [7 empty] | 115.0012 | 199.1787 | 229.9969 | 0.99999 | -59.9988 | 0.00044438 | 0.00059097 | 4.0801e-05 | 2.0138e-06 | 0.00013687 | -0.005391 | 0.0042006 | 27.4339 | 2.0485 | 0.0086049 | 0.50002 | 229.9993 | 0.00096835 | 99994.609 | 0 | 0 | 109.265 | 23.9114 | 0.32291
63 | FLUHIGH | 1 | 230 | 36.87 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9969 | 0.99999 | 36.8967 | 0.79972 | 0.0053894 | 0.0057098 | 2.054 | 100005.3894 | 2015-01-29 18:17 | 50 | 0.70605 | WP- | 0.52246 | VP- | 0.42223 | -80.5311 | 1.1331 | 1.1327 | 107.9339 | 5109.4012 | 183.9435 | 0.41678 | -80.9628 | 1.1331 | 1.1327 | 107.9564 | 5109.5458 | 183.9448 | 0.42314 | -80.6278 | 1.1331 | 1.1332 | 107.9607 | 5109.4853 | 183.9447 | 0.41825 | -80.8345 | 1.1331 | 1.1332 | 107.9328 | 5109.4018 | 183.9444 | 0.42084 | -81.0612 | 1.1331 | 1.1326 | 107.9418 | 5109.4782 | 183.9445 | [7 empty] | 183.9307 | 138.0826 | 229.9969 | 0.99999 | 36.8967 | 0.0003988 | 0.00037812 | 7.4382e-05 | 2.3176e-06 | 6.3577e-05 | 0.0053894 | 0.0027798 | 25.8689 | 2.054 | 0.0057098 | 0.79972 | 229.9996 | 0.0013086 | 100005.3894 | 0 | 0 | 109.265 | 23.9114 | 0.32291
64 | FLUHIGH | 1 | 230 | 36.87 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9969 | 0.99999 | 36.8972 | 0.79971 | 0.0062852 | 0.0057025 | 2.0544 | 100006.2852 | 2015-01-29 18:21 | 50.0005 | 0.70605 | WP- | 0.52246 | VP- | 0.42495 | -81.0336 | 1.1331 | 1.1332 | 107.9578 | 5109.4567 | 183.9441 | 0.4171 | -81.2227 | 1.1331 | 1.1334 | 107.9481 | 5109.4212 | 183.9444 | 0.42494 | -80.5786 | 1.1331 | 1.1327 | 107.9447 | 5109.4957 | 183.9436 | 0.42197 | -80.5558 | 1.1331 | 1.1332 | 107.9486 | 5109.5056 | 183.9439 | 0.42238 | -81.2937 | 1.1331 | 1.1329 | 107.9734 | 5109.5196 | 183.9443 | [7 empty] | 183.9297 | 138.084 | 229.9969 | 0.99999 | 36.8972 | 0.00065622 | 0.00065867 | 0.0001202 | 3.244e-06 | 0.00011635 | 0.0062852 | 0.0027757 | 25.757 | 2.0544 | 0.0057025 | 0.79971 | 229.9996 | 0.0015746 | 100006.2852 | 0 | 0 | 109.265 | 23.9114 | 0.32291
65 | FLUHIGH | 1 | 230 | 36.87 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.997 | 0.99998 | 36.8966 | 0.79972 | 0.0067804 | 0.0056852 | 2.0555 | 100006.7804 | 2015-01-29 18:25 | 49.9999 | 0.70605 | WP- | 0.52246 | VP- | 0.41964 | -80.8767 | 1.1331 | 1.1334 | 107.9455 | 5109.5223 | 183.9444 | 0.42478 | -80.6824 | 1.1331 | 1.1328 | 107.9416 | 5109.5334 | 183.9436 | 0.42283 | -80.381 | 1.1331 | 1.133 | 107.94 | 5109.4797 | 183.9446 | 0.42092 | -80.4404 | 1.1331 | 1.1329 | 107.9274 | 5109.4864 | 183.945 | 0.42262 | -81.3265 | 1.1331 | 1.1333 | 107.9569 | 5109.5645 | 183.9438 | [7 empty] | 183.9302 | 138.0817 | 229.997 | 0.99998 | 36.8966 | 0.00048829 | 0.00063285 | 0.0001044 | 2.1008e-06 | 0.00014241 | 0.0067804 | 0.0027658 | 25.4775 | 2.0555 | 0.0056852 | 0.79972 | 230.0008 | 0.00093962 | 100006.7804 | 0 | 0 | 109.268 | 23.9192 | 0.32291
66 | FLUHIGH | 1 | 230 | 36.87 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  | ACal All done & RD31 reset 19.71°C 60.90%. | 229.9965 | 1 | 36.8673 | 0.80003 | 0.0076623 | 0.0056211 | 2.0603 | 100007.6623 | 2015-01-30 10:18 | 50.0006 | 0.70605 | WP- | 0.52148 | VP- | 0.30783 | -142.3377 | 1.1331 | 1.133 | 107.9321 | 5111.5721 | 184.018 | 0.30533 | -142.9257 | 1.1331 | 1.1328 | 107.9582 | 5111.6073 | 184.0182 | 0.30356 | -142.768 | 1.1331 | 1.1328 | 107.964 | 5111.6729 | 184.0184 | 0.30486 | -141.8883 | 1.1331 | 1.1333 | 107.9382 | 5111.5905 | 184.017 | 0.30496 | -142.9202 | 1.1331 | 1.1326 | 107.9763 | 5111.6179 | 184.0172 | [7 empty] | 184.0039 | 137.9897 | 229.9965 | 1 | 36.8673 | 0.00047542 | 0.00045962 | 7.6007e-05 | 2.3826e-06 | 7.4545e-05 | 0.0076623 | 0.0027283 | 24.2964 | 2.0603 | 0.0056211 | 0.80003 | 229.9961 | 0.0011498 | 100007.6623 | 0 | 0 | 109.147 | 23.6074 | 0.32293
67 | FLUHIGH | 1 | 230 | 36.87 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9968 | 0.99999 | 36.8672 | 0.80003 | 0.0076563 | 0.0056265 | 2.0598 | 100007.6563 | 2015-01-30 10:22 | 50.0004 | 0.70605 | WP- | 0.52148 | VP- | 0.30736 | -142.136 | 1.1331 | 1.1333 | 107.9596 | 5111.6296 | 184.0169 | 0.30588 | -142.3453 | 1.1331 | 1.1327 | 107.9425 | 5111.6255 | 184.0171 | 0.3068 | -141.7888 | 1.1331 | 1.1328 | 107.9475 | 5111.554 | 184.0165 | 0.30603 | -141.1886 | 1.1331 | 1.1326 | 107.969 | 5111.5331 | 184.0163 | 0.30647 | -141.4615 | 1.1331 | 1.1335 | 107.94 | 5111.5566 | 184.0165 | [7 empty] | 184.0028 | 137.9883 | 229.9968 | 0.99999 | 36.8672 | 0.00033527 | 0.00037496 | 0.00010516 | 2.0368e-06 | 4.5567e-05 | 0.0076563 | 0.0027315 | 24.4047 | 2.0598 | 0.0056265 | 0.80003 | 229.9984 | 0.0010989 | 100007.6563 | 0 | 0 | 109.147 | 23.6074 | 0.32293
68 | FLUHIGH | 1 | 230 | 36.87 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9966 | 1 | 36.8673 | 0.80003 | 0.0075201 | 0.0056166 | 2.0607 | 100007.5201 | 2015-01-30 10:26 | 50.0005 | 0.70605 | WP- | 0.52148 | VP- | 0.30589 | -142.7313 | 1.1331 | 1.1332 | 107.9699 | 5111.5697 | 184.0167 | 0.30343 | -142.4647 | 1.1331 | 1.1326 | 107.9578 | 5111.6114 | 184.0174 | 0.30835 | -142.2532 | 1.1331 | 1.1334 | 107.9455 | 5111.6019 | 184.0168 | 0.30609 | -141.8141 | 1.1331 | 1.1335 | 107.9713 | 5111.5773 | 184.0174 | 0.30817 | -142.069 | 1.1331 | 1.1332 | 107.9387 | 5111.5611 | 184.0161 | [7 empty] | 184.0032 | 137.9892 | 229.9966 | 1 | 36.8673 | 0.00038488 | 0.00038377 | 9.5535e-05 | 1.7227e-06 | 9.3047e-05 | 0.0075201 | 0.0027256 | 24.2037 | 2.0607 | 0.0056166 | 0.80003 | 229.9971 | 0.00093893 | 100007.5201 | 0 | 0 | 109.146 | 23.6048 | 0.32293
69 | FLUHIGH | 0.5 | 230 | 36.87 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9966 | 0.49999 | 36.8669 | 0.80003 | 0.0066877 | 0.0057261 | 2.0524 | 100006.6877 | 2015-01-30 10:31 | 50 | 0.35352 | WP- | 0.26074 | VP- | 0.40749 | -116.1357 | 1.1331 | 1.1325 | 107.9429 | 2555.7684 | 92.0092 | 0.40955 | -115.274 | 1.1331 | 1.1335 | 107.9468 | 2555.7572 | 92.0095 | 0.41258 | -115.7013 | 1.1331 | 1.1333 | 107.9307 | 2555.7674 | 92.0093 | 0.41087 | -116.0805 | 1.1331 | 1.1325 | 107.9505 | 2555.7639 | 92.0095 | 0.41195 | -115.9456 | 1.1331 | 1.1334 | 107.9443 | 2555.7651 | 92.0092 | [7 empty] | 92.0014 | 68.9934 | 229.9966 | 0.49999 | 36.8669 | 0.00045507 | 0.00050571 | 9.0385e-05 | 2.2535e-06 | 0.00021019 | 0.0066877 | 0.00279 | 26.3177 | 2.0524 | 0.0057261 | 0.80003 | 459.9979 | 0.0045298 | 100006.6877 | 0 | 0 | 109.145 | 23.6022 | 0.32293
70 | FLUHIGH | 0.5 | 230 | 36.87 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9969 | 0.5 | 36.8666 | 0.80003 | 0.0058722 | 0.0057576 | 2.0506 | 100005.8722 | 2015-01-30 10:35 | 50.0005 | 0.35352 | WP- | 0.26074 | VP- | 0.40829 | -115.364 | 1.1331 | 1.133 | 107.9494 | 2555.7432 | 92.0091 | 0.41317 | -116.0267 | 1.1331 | 1.133 | 107.9416 | 2555.7559 | 92.0092 | 0.40672 | -115.4405 | 1.1331 | 1.1327 | 107.9778 | 2555.7626 | 92.0097 | 0.40672 | -115.3659 | 1.1331 | 1.1334 | 107.9485 | 2555.7532 | 92.0092 | 0.40844 | -115.8654 | 1.1331 | 1.1333 | 107.9521 | 2555.8182 | 92.0099 | [7 empty] | 92.0022 | 68.9934 | 229.9969 | 0.5 | 36.8666 | 0.00042289 | 0.00057541 | 6.5119e-05 | 1.0112e-06 | 0.00033181 | 0.0058722 | 0.0028078 | 26.8264 | 2.0506 | 0.0057576 | 0.80003 | 459.9963 | 0.0018737 | 100005.8722 | 0 | 0 | 109.146 | 23.6048 | 0.32293
71 | FLUHIGH | 0.5 | 230 | 36.87 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.997 | 0.49999 | 36.8666 | 0.80003 | 0.0071056 | 0.0058931 | 2.0456 | 100007.1056 | 2015-01-30 10:39 | 50.0003 | 0.35352 | WP- | 0.26074 | VP- | 0.40537 | -115.4528 | 1.1331 | 1.1329 | 107.9126 | 2555.7766 | 92.0097 | 0.41224 | -115.3223 | 1.1331 | 1.1336 | 107.9312 | 2555.7988 | 92.009 | 0.41318 | -115.8206 | 1.1331 | 1.1332 | 107.9771 | 2555.8116 | 92.0089 | 0.40759 | -115.9556 | 1.1331 | 1.1331 | 107.967 | 2555.8068 | 92.0098 | 0.41131 | -115.3926 | 1.1331 | 1.1328 | 107.9288 | 2555.7334 | 92.0091 | [7 empty] | 92.0017 | 68.9932 | 229.997 | 0.49999 | 36.8666 | 0.00064094 | 0.00046835 | 4.5967e-05 | 2.4009e-06 | 0.00028876 | 0.0071056 | 0.0028808 | 28.3326 | 2.0456 | 0.0058931 | 0.80003 | 459.9989 | 0.0050332 | 100007.1056 | 0 | 0 | 109.148 | 23.61 | 0.32293
72 | FLUHIGH | 0.5 | 230 | 0 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9968 | 0.50001 | -0.0020313 | 1 | 0.0033382 | 0.0049424 | 2.0642 | 100003.3382 | 2015-01-30 10:43 | 50.0009 | 0.44141 | WP- | 0 | VP+ | 0.16119 | -111.2229 | 1.1331 | 1.1332 | 107.9647 | 3194.6089 | 115.0055 | 0.15995 | -110.8139 | 1.1331 | 1.1333 | 107.9487 | 3194.5758 | 115.0054 | 0.15531 | -112.2064 | 1.1331 | 1.1334 | 107.9659 | 3194.5521 | 115.0057 | 0.15724 | -110.1472 | 1.1331 | 1.1334 | 107.9619 | 3194.5964 | 115.0056 | 0.1594 | -110.1814 | 1.1331 | 1.1331 | 107.9582 | 3194.5998 | 115.0057 | [7 empty] | 115.0013 | 0.0040771 | 229.9968 | 0.50001 | -0.0020313 | 0.00055871 | 0.00044407 | 6.8813e-05 | 2.314e-06 | 0.00022125 | 0.0033382 | 0.0023944 | 23.4152 | 2.0642 | 0.0049424 | 1 | 459.9822 | 0.0045318 | 100003.3382 | 0 | 0 | 109.149 | 23.6125 | 0.32293
73 | FLUHIGH | 0.5 | 230 | 0 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9968 | 0.50002 | -0.0020599 | 1 | 0.0020476 | 0.0048741 | 2.0675 | 100002.0476 | 2015-01-30 10:47 | 50.0004 | 0.44141 | WP- | 0 | VP+ | 0.15499 | -111.6966 | 1.1331 | 1.1333 | 107.9312 | 3194.6546 | 115.0057 | 0.15366 | -110.8882 | 1.1331 | 1.1332 | 107.9553 | 3194.5896 | 115.0059 | 0.15725 | -112.0738 | 1.1331 | 1.133 | 107.9419 | 3194.563 | 115.0063 | 0.15879 | -109.8244 | 1.1331 | 1.133 | 107.9518 | 3194.5586 | 115.0045 | 0.15481 | -110.7612 | 1.1331 | 1.1335 | 107.9256 | 3194.529 | 115.0055 | [7 empty] | 115.0025 | 0.0041346 | 229.9968 | 0.50002 | -0.0020599 | 0.00048656 | 0.00046611 | 4.5361e-05 | 2.021e-06 | 0.00023222 | 0.0020476 | 0.0023574 | 22.7058 | 2.0675 | 0.0048741 | 1 | 459.9772 | 0.0039635 | 100002.0476 | 0 | 0 | 109.146 | 23.6048 | 0.32293
74 | FLUHIGH | 1 | 230 | 90 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0040824 | 6 | 123 |  | phase angle set at wrong polarity | 229.9965 | 0.99998 | 90.0002 | -1 | 1.5512e+36 | 0.013798 | 2.0829 | 1.5512e+39 | 2015-01-30 11:05 | 50.0003 | 0 | WP+ | 0.87012 | VP- | 0.3473 | -38.6421 | 1.1331 | 1.1326 | 107.9247 | 9.91e+37 | -230.0161 | 0.34827 | -38.2332 | 1.1331 | 1.1332 | 107.968 | 9.91e+37 | -230.0157 | 0.34372 | -38.8579 | 1.1331 | 1.1332 | 107.939 | 9.91e+37 | -230.0158 | 0.34726 | -38.7031 | 1.1331 | 1.133 | 107.9336 | 9.91e+37 | -230.0157 | 0.34552 | -38.1205 | 1.1331 | 1.1328 | 107.9369 | 9.91e+37 | -230.0153 | [7 empty] | -0.00091113 | 229.9913 | 229.9965 | 0.99998 | 90.0002 | 0.00032168 | 0.0004374 | 4.6938e-05 | 1.8518e-06 | 8.0136e-05 | 1.5512e+36 | 0.0066243 | 19.957 | 2.0829 | 0.013798 | -1 | 230.0016 | 0.00093786 | 1.5512e+39 | 0 | 0 | 109.124 | 23.5481 | 0.32293
75 | FLUHIGH | 1 | 230 | -90 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0040824 | 6 | 123 |  |  | 229.9966 | 0.99997 | -90 | 1 | 0.002383 | 0.013811 | 2.0822 | 100002.383 | 2015-01-30 11:09 | 50.0003 | 0 | WP- | 0.87012 | VP+ | 0.21226 | 120.8171 | 1.1331 | 1.1333 | 107.9466 | 6388.7608 | 230.0108 | 0.21308 | 121.7758 | 1.1331 | 1.1328 | 107.9395 | 6388.8018 | 230.0096 | 0.20751 | 120.9868 | 1.1331 | 1.133 | 107.9647 | 6388.7426 | 230.01 | 0.21508 | 121.4775 | 1.1331 | 1.1332 | 107.9663 | 6388.7397 | 230.0106 | 0.2108 | 122.2841 | 1.1331 | 1.1332 | 107.9502 | 6388.7614 | 230.01 | [7 empty] | -3.7823e-05 | 229.9899 | 229.9966 | 0.99997 | -90 | 0.00046773 | 0.00067343 | 7.3159e-05 | 2.673e-06 | 0.00011652 | 0.002383 | 0.0066328 | 20.0689 | 2.0822 | 0.013811 | 1 | 230.0033 | 0.0012513 | 100002.383 | 0 | 0 | 109.121 | 23.5404 | 0.32293
76 | FLUHIGH | 1 | 230 | -90 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0040824 | 6 | 123 |  |  | 229.9968 | 0.99996 | -90 | 1 | 0.0024671 | 0.013795 | 2.083 | 100002.4671 | 2015-01-30 11:13 | 50.0004 | 0 | WP+ | 0.87012 | VP+ | 0.21029 | 120.9285 | 1.1331 | 1.1329 | 107.9654 | 6388.7203 | 230.0104 | 0.20638 | 120.3737 | 1.1331 | 1.1326 | 107.9711 | 6388.7379 | 230.0106 | 0.21111 | 121.7227 | 1.1331 | 1.1332 | 107.9448 | 6388.6931 | 230.0104 | 0.21245 | 120.4999 | 1.1331 | 1.1329 | 107.9489 | 6388.7179 | 230.0094 | 0.20768 | 121.0869 | 1.1331 | 1.1332 | 107.9537 | 6388.6959 | 230.0098 | [7 empty] | -0.00013612 | 229.988 | 229.9968 | 0.99996 | -90 | 0.00039765 | 0.00044424 | 7.4986e-05 | 2.0582e-06 | 9.9065e-05 | 0.0024671 | 0.0066224 | 19.9442 | 2.083 | 0.013795 | 1 | 230.0057 | 0.0011448 | 100002.4671 | 0 | 0 | 109.115 | 23.5249 | 0.32293
77 | FLUHIGH | 1 | 230 | -90 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0040824 | 6 | 123 |  |  | 229.9968 | 0.99995 | -89.9999 | 1 | 0.0029754 | 0.013798 | 2.0829 | 100002.9754 | 2015-01-30 11:18 | 50 | 0 | WP+ | 0.87012 | VP+ | 0.21266 | 121.7254 | 1.1331 | 1.1333 | 107.9565 | 6388.6802 | 230.0095 | 0.21368 | 121.2058 | 1.1331 | 1.133 | 107.9485 | 6388.7002 | 230.0095 | 0.21019 | 121.2336 | 1.1331 | 1.1325 | 107.9584 | 6388.6914 | 230.0091 | 0.20821 | 121.0067 | 1.1331 | 1.133 | 107.9604 | 6388.6683 | 230.0099 | 0.21444 | 120.9929 | 1.1331 | 1.1332 | 107.9889 | 6388.6878 | 230.009 | [7 empty] | 0.00027579 | 229.9858 | 229.9968 | 0.99995 | -89.9999 | 0.00039387 | 0.00039841 | 6.8622e-05 | 1.7478e-06 | 9.8124e-05 | 0.0029754 | 0.0066246 | 19.9697 | 2.0829 | 0.013798 | 1 | 230.0077 | 0.0009325 | 100002.9754 | 0 | 0 | 109.108 | 23.5069 | 0.32293
78 | FLUHIGH | 1 | 230 | -30 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0046547 | 6 | 123 |  |  | 229.9968 | 1 | -30 | 0.5 | 0.0023639 | 0.014751 | 2.0861 | 100002.3639 | 2015-01-30 11:22 | 50.0001 | 0.76465 | WP- | 0.43457 | VP+ | 0.30814 | -68.3628 | 1.1331 | 1.1331 | 107.9771 | 3194.5421 | 115.0178 | 0.31186 | -68.7492 | 1.1331 | 1.134 | 107.9378 | 3194.5599 | 115.0178 | 0.31323 | -68.3082 | 1.1331 | 1.133 | 107.9692 | 3194.5655 | 115.0183 | 0.31058 | -67.9074 | 1.1331 | 1.1333 | 107.9381 | 3194.5436 | 115.0191 | 0.30996 | -67.9931 | 1.1331 | 1.1331 | 107.9613 | 3194.5805 | 115.0179 | [7 empty] | 199.188 | 115.0014 | 229.9968 | 1 | -30 | 0.0005084 | 0.00030019 | 7.4382e-05 | 1.9033e-06 | 8.694e-05 | 0.0023639 | 0.007071 | 19.4763 | 2.0861 | 0.014751 | 0.5 | 229.991 | 0.00091525 | 100002.3639 | 0 | 0 | 109.101 | 23.4889 | 0.32293
79 | FLUHIGH | 1 | 230 | -30 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0046547 | 6 | 123 |  |  | 229.9968 | 1 | -30.0004 | 0.50001 | 0.00054935 | 0.014776 | 2.0849 | 100000.5493 | 2015-01-30 11:26 | 50.0004 | 0.76465 | WP- | 0.43457 | VP+ | 0.31005 | -67.4326 | 1.1331 | 1.133 | 107.9547 | 3194.5526 | 115.0186 | 0.31453 | -68.0769 | 1.1331 | 1.1326 | 107.9335 | 3194.5156 | 115.0187 | 0.31354 | -67.9471 | 1.1331 | 1.1331 | 107.951 | 3194.5555 | 115.019 | 0.31446 | -68.4729 | 1.1331 | 1.1325 | 107.9556 | 3194.5445 | 115.0179 | 0.31036 | -68.6835 | 1.1331 | 1.1332 | 107.9181 | 3194.5639 | 115.0178 | [7 empty] | 199.1876 | 115.003 | 229.9968 | 1 | -30.0004 | 0.00057833 | 0.00047023 | 3.33e-05 | 2.5395e-06 | 0.00010455 | 0.00054935 | 0.0070872 | 19.6522 | 2.0849 | 0.014776 | 0.50001 | 229.9906 | 0.001235 | 100000.5493 | 0 | 0 | 109.096 | 23.476 | 0.32293
80 | FLUHIGH | 1 | 230 | -30 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0046547 | 6 | 123 |  |  | 229.9968 | 1 | -30.0002 | 0.5 | 0.00035735 | 0.014762 | 2.0856 | 100000.3573 | 2015-01-30 11:30 | 50.0003 | 0.76465 | WP- | 0.43457 | VP+ | 0.30769 | -67.5922 | 1.1331 | 1.1327 | 107.9392 | 3194.5256 | 115.0192 | 0.31211 | -68.3598 | 1.1331 | 1.1328 | 107.9644 | 3194.5372 | 115.0175 | 0.31027 | -68.3049 | 1.1331 | 1.1329 | 107.9218 | 3194.5199 | 115.0187 | 0.31217 | -67.6616 | 1.1331 | 1.1328 | 107.9541 | 3194.5112 | 115.0189 | 0.30809 | -67.9319 | 1.1331 | 1.1332 | 107.9772 | 3194.5244 | 115.0189 | [7 empty] | 199.1887 | 115.0024 | 229.9968 | 1 | -30.0002 | 0.00048385 | 0.00040424 | 9.3984e-05 | 2.0944e-06 | 8.9115e-05 | 0.00035735 | 0.0070779 | 19.552 | 2.0856 | 0.014762 | 0.5 | 229.9899 | 0.0010299 | 100000.3573 | 0 | 0 | 109.094 | 23.4708 | 0.32293
81 | FLUHIGH | 1 | 230 | -126.87 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0042105 | 6 | 123 |  |  | 229.9969 | 0.99997 | -126.8721 | 0.79998 | -0.00047421 | 0.014003 | 2.0834 | 99999.5258 | 2015-01-30 11:34 | 50.0005 | 0.5293 | WP+ | 0.69629 | VP+ | 0.20375 | 171.0966 | 1.1331 | 1.1331 | 107.9593 | 5110.722 | 183.9985 | 0.20257 | 170.9386 | 1.1331 | 1.1333 | 107.9486 | 5110.7144 | 183.9979 | 0.20057 | 170.8665 | 1.1331 | 1.1328 | 107.9271 | 5110.7131 | 183.9984 | 0.20248 | 170.0608 | 1.1331 | 1.1333 | 107.9753 | 5110.7358 | 183.9992 | 0.20196 | 172.1342 | 1.1331 | 1.1327 | 107.9299 | 5110.6985 | 183.9987 | [7 empty] | -138.0012 | 183.9867 | 229.9969 | 0.99997 | -126.8721 | 0.0005733 | 0.00013126 | 9.9329e-05 | 1.6846e-06 | 0.00010783 | -0.00047421 | 0.0067212 | 19.8882 | 2.0834 | 0.014003 | 0.79998 | 230.0037 | 0.00089847 | 99999.5258 | 0 | 0 | 109.096 | 23.476 | 0.32293
82 | FLUHIGH | 1 | 230 | -126.87 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0042105 | 6 | 123 |  |  | 229.9969 | 0.99997 | -126.8725 | 0.79997 | 0.00096524 | 0.014004 | 2.0834 | 100000.9652 | 2015-01-30 11:39 | 50.0002 | 0.5293 | WP+ | 0.69629 | VP+ | 0.20303 | 170.612 | 1.1331 | 1.1334 | 107.9434 | 5110.7431 | 183.9983 | 0.20212 | 171.3882 | 1.1331 | 1.1336 | 107.9369 | 5110.7361 | 183.9996 | 0.2033 | 171.7117 | 1.1331 | 1.1326 | 107.964 | 5110.7517 | 183.9993 | 0.20158 | 171.3554 | 1.1331 | 1.1332 | 107.9635 | 5110.7483 | 183.9992 | 0.20593 | 171.375 | 1.1331 | 1.133 | 107.9546 | 5110.7461 | 183.9988 | [7 empty] | -138.0016 | 183.985 | 229.9969 | 0.99997 | -126.8725 | 0.00025352 | 0.00033407 | 7.3079e-05 | 1.2231e-06 | 8.0059e-05 | 0.00096524 | 0.0067217 | 19.8931 | 2.0834 | 0.014004 | 0.79997 | 230.0048 | 0.00069404 | 100000.9652 | 0 | 0 | 109.095 | 23.4734 | 0.32293
83 | FLUHIGH | 1 | 230 | -126.87 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0042105 | 6 | 123 |  |  | 229.9967 | 0.99996 | -126.8725 | 0.79997 | 0.00034864 | 0.014006 | 2.0832 | 100000.3486 | 2015-01-30 11:43 | 50 | 0.5293 | WP+ | 0.69629 | VP+ | 0.20454 | 171.2858 | 1.1331 | 1.133 | 107.953 | 5110.6904 | 183.9989 | 0.2034 | 171.2588 | 1.1331 | 1.1329 | 107.9215 | 5110.6981 | 183.9989 | 0.19978 | 171.4852 | 1.1331 | 1.1328 | 107.9572 | 5110.6906 | 183.9987 | 0.20422 | 171.1859 | 1.1331 | 1.1329 | 107.9335 | 5110.6874 | 183.9986 | 0.20135 | 171.3206 | 1.1331 | 1.1333 | 107.9584 | 5110.7134 | 183.9988 | [7 empty] | -138.0014 | 183.9844 | 229.9967 | 0.99996 | -126.8725 | 0.00011045 | 0.00049642 | 5.9539e-05 | 1.7749e-06 | 5.4089e-05 | 0.00034864 | 0.0067233 | 19.9126 | 2.0832 | 0.014006 | 0.79997 | 230.005 | 0.00080275 | 100000.3486 | 0 | 0 | 109.096 | 23.476 | 0.32293
84 | FLUHIGH | 6 | 230 | 0 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9966 | 6.0003 | -0.0020998 | 1 | 0.0032343 | 0.0047158 | 2.0831 | 100003.2343 | 2015-01-30 12:05 | 50.0003 | 0.5293 | WP- | 0 | VP+ | 0.048581 | 112.0767 | 1.1331 | 1.1325 | 107.9499 | 38336.234 | 1380.0852 | 0.048114 | 112.2147 | 1.1331 | 1.1331 | 107.9567 | 38336.0682 | 1380.0865 | 0.046958 | 111.908 | 1.1331 | 1.1327 | 107.9427 | 38336.1802 | 1380.0862 | 0.047869 | 112.1467 | 1.1331 | 1.1333 | 107.9524 | 38336.2039 | 1380.0887 | 0.045395 | 113.3982 | 1.1331 | 1.1332 | 107.9489 | 38335.8888 | 1380.0886 | [7 empty] | 1380.0555 | 0.050576 | 229.9966 | 6.0003 | -0.0020998 | 0.0023888 | 0.00091061 | 4.396e-05 | 1.0768e-05 | 3.7809e-05 | 0.0032343 | 0.0022639 | 19.9405 | 2.0831 | 0.0047158 | 1 | 38.3307 | 0.00016075 | 100003.2343 | 0 | 0 | 109.11 | 23.512 | 0.32293
85 | FLUHIGH | 6 | 230 | 0 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9966 | 6.0003 | -0.0020049 | 1 | 0.0028067 | 0.0047269 | 2.0815 | 100002.8067 | 2015-01-30 12:09 | 50.0002 | 0.5293 | WP- | 0 | VP+ | 0.046342 | 114.0481 | 1.1331 | 1.133 | 107.9602 | 38336.0015 | 1380.0863 | 0.045949 | 114.3461 | 1.1331 | 1.1331 | 107.9447 | 38335.8044 | 1380.0829 | 0.04473 | 114.7219 | 1.1331 | 1.1327 | 107.9478 | 38335.388 | 1380.0869 | 0.044816 | 114.0555 | 1.1331 | 1.1326 | 107.9514 | 38335.7194 | 1380.0845 | 0.046764 | 115.4757 | 1.1331 | 1.1333 | 107.9738 | 38335.4544 | 1380.0809 | [7 empty] | 1380.0455 | 0.04829 | 229.9966 | 6.0003 | -0.0020049 | 0.002285 | 0.0015741 | 0.00012178 | 7.5538e-06 | 6.5352e-05 | 0.0028067 | 0.0022709 | 20.1827 | 2.0815 | 0.0047269 | 1 | 38.331 | 8.5223e-05 | 100002.8067 | 0 | 0 | 109.121 | 23.5404 | 0.32293
86 | FLUHIGH | 6 | 230 | 0 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9967 | 6.0003 | -0.0019517 | 1 | 0.0030533 | 0.0047259 | 2.0817 | 100003.0533 | 2015-01-30 12:14 | 50.0002 | 0.5293 | WP- | 0 | VP+ | 0.045375 | 115.6983 | 1.1331 | 1.1331 | 107.9757 | 38335.6915 | 1380.0784 | 0.046637 | 116.0337 | 1.1331 | 1.133 | 107.9356 | 38336.1132 | 1380.0874 | 0.045237 | 114.2783 | 1.1331 | 1.1332 | 107.9577 | 38335.6568 | 1380.087 | 0.045868 | 115.4682 | 1.1331 | 1.133 | 107.9522 | 38335.525 | 1380.0863 | 0.045479 | 113.3762 | 1.1331 | 1.1332 | 107.9594 | 38335.7442 | 1380.0847 | [7 empty] | 1380.0447 | 0.047009 | 229.9967 | 6.0003 | -0.0019517 | 0.0010867 | 0.0019542 | 6.764e-05 | 3.9042e-06 | 8.1134e-05 | 0.0030533 | 0.0022702 | 20.1609 | 2.0817 | 0.0047259 | 1 | 38.331 | 5.4188e-05 | 100003.0533 | 0 | 0 | 109.126 | 23.5533 | 0.32293
87 | FLUHIGH | 6 | 230 | -60 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9967 | 6.0001 | -59.9997 | 0.5 | -0.0015804 | 0.0085117 | 2.0526 | 99998.4196 | 2015-01-30 12:18 | 50.0002 | 0.26465 | WP- | 0.45215 | VP+ | 0.13625 | 124.5398 | 1.1331 | 1.1325 | 107.9653 | 19166.816 | 689.9802 | 0.13538 | 124.3426 | 1.1331 | 1.1331 | 107.9575 | 19166.7058 | 689.9838 | 0.13564 | 124.6845 | 1.1331 | 1.1333 | 107.9273 | 19166.6348 | 689.9837 | 0.13533 | 125.1523 | 1.1331 | 1.1331 | 107.9546 | 19166.6377 | 689.9835 | 0.13565 | 124.7474 | 1.1331 | 1.1333 | 107.9793 | 19166.754 | 689.9813 | [7 empty] | 690.0125 | 1195.1217 | 229.9967 | 6.0001 | -59.9997 | 0.0012544 | 0.0003808 | 7.1996e-05 | 4.6131e-06 | 3.8818e-05 | -0.0015804 | 0.0041467 | 26.241 | 2.0526 | 0.0085117 | 0.5 | 38.3319 | 8.287e-05 | 99998.4196 | 0 | 0 | 109.125 | 23.5507 | 0.32293
88 | FLUHIGH | 6 | 230 | -60 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9968 | 6.0002 | -59.9998 | 0.5 | -0.0016061 | 0.0086501 | 2.0466 | 99998.3939 | 2015-01-30 12:22 | 50.0005 | 0.26465 | WP- | 0.45215 | VP+ | 0.13557 | 123.902 | 1.1331 | 1.1328 | 107.9275 | 19166.7149 | 689.9767 | 0.13486 | 124.1564 | 1.1331 | 1.1329 | 107.9566 | 19166.5307 | 689.9816 | 0.13583 | 124.4996 | 1.1331 | 1.1331 | 107.9714 | 19166.3694 | 689.9842 | 0.13575 | 124.1492 | 1.1331 | 1.133 | 107.9684 | 19167.3293 | 689.9853 | 0.13658 | 123.9988 | 1.1331 | 1.1331 | 107.9529 | 19167.248 | 689.9832 | [7 empty] | 690.0173 | 1195.1353 | 229.9968 | 6.0002 | -59.9998 | 0.0015396 | 0.0015044 | 9.4878e-05 | 5.6614e-06 | 5.2589e-05 | -0.0016061 | 0.0042266 | 28.0333 | 2.0466 | 0.0086501 | 0.5 | 38.3315 | 6.2409e-05 | 99998.3939 | 0 | 0 | 109.123 | 23.5455 | 0.32293
89 | FLUHIGH | 6 | 230 | -60 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9967 | 6.0002 | -59.9999 | 0.5 | -0.0033876 | 0.0086734 | 2.0458 | 99996.6124 | 2015-01-30 12:26 | 50.0002 | 0.26465 | WP- | 0.45215 | VP+ | 0.13261 | 124.5681 | 1.1331 | 1.1333 | 107.9382 | 19166.748 | 689.9829 | 0.13558 | 124.7805 | 1.1331 | 1.1334 | 107.9678 | 19166.2594 | 689.9839 | 0.13562 | 124.7837 | 1.1331 | 1.1333 | 107.9153 | 19165.9925 | 689.9846 | 0.1348 | 124.3344 | 1.1331 | 1.1334 | 107.9483 | 19166.2438 | 689.9833 | 0.1348 | 124.314 | 1.1331 | 1.1336 | 107.9577 | 19166.2593 | 689.9818 | [7 empty] | 690.0102 | 1195.1289 | 229.9967 | 6.0002 | -59.9999 | 0.0020287 | 0.0021258 | 7.6731e-05 | 6.8031e-06 | 9.653e-05 | -0.0033876 | 0.0042396 | 28.2779 | 2.0458 | 0.0086734 | 0.5 | 38.3318 | 8.9827e-05 | 99996.6124 | 0 | 0 | 109.122 | 23.543 | 0.32293
90 | FLUHIGH | 6 | 230 | 36.87 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9967 | 6 | 36.8668 | 0.80003 | 0.0069901 | 0.0056407 | 2.0588 | 100006.9901 | 2015-01-30 12:30 | 50 | 0.42383 | WP- | 0.31348 | VP- | 0.33753 | -60.004 | 1.1331 | 1.1324 | 107.9424 | 30669.7976 | 1104.1012 | 0.33808 | -59.6748 | 1.1331 | 1.1331 | 107.9595 | 30669.7608 | 1104.1083 | 0.33827 | -59.8882 | 1.1331 | 1.1332 | 107.9557 | 30669.7455 | 1104.1094 | 0.33886 | -59.6879 | 1.1331 | 1.1332 | 107.9266 | 30669.6152 | 1104.1047 | 0.33823 | -59.6935 | 1.1331 | 1.133 | 107.946 | 30669.9004 | 1104.1106 | [7 empty] | 1104.0343 | 827.9337 | 229.9967 | 6 | 36.8668 | 0.00077844 | 0.002082 | 3.3604e-05 | 4.5161e-06 | 7.7054e-05 | 0.0069901 | 0.0027398 | 24.6604 | 2.0588 | 0.0056407 | 0.80003 | 38.3326 | 5.3548e-05 | 100006.9901 | 0 | 0 | 109.121 | 23.5404 | 0.32293
91 | FLUHIGH | 6 | 230 | 36.87 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9966 | 6 | 36.8669 | 0.80003 | 0.0080184 | 0.005707 | 2.0537 | 100008.0184 | 2015-01-30 12:35 | 50.0003 | 0.42383 | WP- | 0.31348 | VP- | 0.33862 | -59.5172 | 1.1331 | 1.1334 | 107.9654 | 30669.7992 | 1104.1084 | 0.33887 | -59.9067 | 1.1331 | 1.133 | 107.9349 | 30670.2063 | 1104.0996 | 0.33815 | -59.3024 | 1.1331 | 1.133 | 107.9543 | 30670.0611 | 1104.1089 | 0.33847 | -59.8952 | 1.1331 | 1.1333 | 107.9337 | 30670.4236 | 1104.1042 | 0.33866 | -59.7647 | 1.1331 | 1.1328 | 107.9445 | 30669.8949 | 1104.1051 | [7 empty] | 1104.0342 | 827.9341 | 229.9966 | 6 | 36.8669 | 0.0021725 | 0.002155 | 7.1753e-05 | 2.7306e-06 | 0.00012304 | 0.0080184 | 0.0027789 | 25.9407 | 2.0537 | 0.005707 | 0.80003 | 38.3325 | 4.7221e-05 | 100008.0184 | 0 | 0 | 109.123 | 23.5455 | 0.32293
92 | FLUHIGH | 6 | 230 | 36.87 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9968 | 6 | 36.8666 | 0.80003 | 0.0070564 | 0.0057509 | 2.0512 | 100007.0564 | 2015-01-30 12:39 | 50.0001 | 0.42383 | WP- | 0.31348 | VP- | 0.3376 | -59.6727 | 1.1331 | 1.1333 | 107.9492 | 30669.3712 | 1104.1075 | 0.33862 | -59.4178 | 1.1331 | 1.1329 | 107.9331 | 30669.9258 | 1104.1038 | 0.34011 | -59.513 | 1.1331 | 1.1327 | 107.9313 | 30670.1407 | 1104.1049 | 0.33823 | -59.5844 | 1.1331 | 1.1333 | 107.9468 | 30669.6723 | 1104.1033 | 0.33918 | -59.7067 | 1.1331 | 1.1331 | 107.9326 | 30670.1943 | 1104.1024 | [7 empty] | 1104.0371 | 827.9299 | 229.9968 | 6 | 36.8666 | 0.0017898 | 0.0019657 | 4.7062e-05 | 8.7917e-06 | 6.313e-05 | 0.0070564 | 0.0028036 | 26.6347 | 2.0512 | 0.0057509 | 0.80003 | 38.3326 | 0.00011755 | 100007.0564 | 0 | 0 | 109.122 | 23.543 | 0.32293
93 | FLUHIGH | 10 | 230 | 0 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9969 | 10.0004 | -0.00072925 | 1 | 0.0021847 | 0.0049145 | 2.0671 | 100002.1847 | 2015-01-30 12:43 | 50.0006 | 0.88281 | WP- | 0 | VP+ | 0.32201 | -111.215 | 1.1331 | 1.1331 | 107.9585 | 63893.0039 | 2300.1199 | 0.32068 | -111.3107 | 1.1331 | 1.1324 | 107.9553 | 63891.3238 | 2300.1216 | 0.32263 | -111.0774 | 1.1331 | 1.1334 | 107.9649 | 63892.9682 | 2300.1145 | 0.32103 | -110.6516 | 1.1331 | 1.1332 | 107.9611 | 63890.6938 | 2300.1086 | 0.32179 | -110.932 | 1.1331 | 1.1331 | 107.9802 | 63892.2699 | 2300.1199 | [7 empty] | 2300.0636 | 0.029275 | 229.9969 | 10.0004 | -0.00072925 | 0.0013685 | 0.0029044 | 8.5962e-05 | 5.2555e-06 | 7.2349e-05 | 0.0021847 | 0.0023774 | 22.7919 | 2.0671 | 0.0049145 | 1 | 22.9988 | 3.5543e-05 | 100002.1847 | 0 | 0 | 109.123 | 23.5455 | 0.32293
94 | FLUHIGH | 10 | 230 | 0 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9969 | 10.0004 | -0.00095927 | 1 | 0.0016436 | 0.004711 | 2.0836 | 100001.6436 | 2015-01-30 12:47 | 50.0004 | 0.88281 | WP- | 0 | VP+ | 0.32261 | -110.8595 | 1.1331 | 1.1331 | 107.9655 | 63890.9572 | 2300.1262 | 0.3208 | -110.6519 | 1.1331 | 1.1331 | 107.9448 | 63891.911 | 2300.1172 | 0.32173 | -110.6634 | 1.1331 | 1.1329 | 107.9513 | 63891.6557 | 2300.1184 | 0.32328 | -110.4335 | 1.1331 | 1.1336 | 107.9503 | 63891.4678 | 2300.1184 | 0.32341 | -110.7081 | 1.1331 | 1.1333 | 107.9397 | 63891.6908 | 2300.1118 | [7 empty] | 2300.0575 | 0.038508 | 229.9969 | 10.0004 | -0.00095927 | 0.0029915 | 0.0017655 | 6.8813e-05 | 1.1143e-05 | 4.3979e-05 | 0.0016436 | 0.002261 | 19.8555 | 2.0836 | 0.004711 | 1 | 22.9988 | 5.0673e-05 | 100001.6436 | 0 | 0 | 109.125 | 23.5507 | 0.32293
95 | FLUHIGH | 10 | 230 | 0 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 |  |  | 229.9968 | 10.0004 | -0.00091681 | 1 | 0.0012753 | 0.0047014 | 2.0849 | 100001.2753 | 2015-01-30 12:51 | 50 | 0.88281 | WP- | 0 | VP+ | 0.32174 | -110.7451 | 1.1331 | 1.1333 | 107.9525 | 63891.6242 | 2300.1125 | 0.321 | -110.8442 | 1.1331 | 1.133 | 107.9981 | 63891.492 | 2300.1099 | 0.32099 | -110.7445 | 1.1331 | 1.133 | 107.9247 | 63891.0963 | 2300.1155 | 0.32189 | -110.7333 | 1.1331 | 1.1334 | 107.9255 | 63891.4482 | 2300.1226 | 0.32136 | -110.7355 | 1.1331 | 1.1332 | 107.94 | 63891.4134 | 2300.1145 | [7 empty] | 2300.0616 | 0.036804 | 229.9968 | 10.0004 | -0.00091681 | 0.00077793 | 0.0007948 | 4.2922e-05 | 3.1742e-06 | 1.9799e-05 | 0.0012753 | 0.0022549 | 19.6504 | 2.0849 | 0.0047014 | 1 | 22.9988 | 2.0359e-05 | 100001.2753 | 0 | 0 | 109.125 | 23.5507 | 0.32293
96 | FLUHIGH | 10 | 230 | -60 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9968 | 10.0001 | -59.9986 | 0.50002 | -0.006861 | 0.0085271 | 2.0517 | 99993.139 | 2015-01-30 12:56 | 50 | 0.44141 | WP- | 0.75293 | VP+ | 0.20866 | -19.3764 | 1.1331 | 1.1329 | 107.9328 | 31942.6848 | 1149.9739 | 0.21091 | -19.3199 | 1.1331 | 1.1333 | 107.9665 | 31943.7202 | 1149.9672 | 0.20911 | -19.0633 | 1.1331 | 1.1328 | 107.9458 | 31943.8152 | 1149.9689 | 0.20794 | -19.0732 | 1.1331 | 1.1327 | 107.945 | 31943.4363 | 1149.9753 | 0.20835 | -19.661 | 1.1331 | 1.1331 | 107.9686 | 31943.0259 | 1149.9729 | [7 empty] | 1150.039 | 1991.8131 | 229.9968 | 10.0001 | -59.9986 | 0.002997 | 0.0029396 | 0.00012594 | 1.1811e-05 | 5.2111e-05 | -0.006861 | 0.0041561 | 26.4998 | 2.0517 | 0.0085271 | 0.50002 | 22.9996 | 4.8387e-05 | 99993.139 | 0 | 0 | 109.124 | 23.5481 | 0.32293
97 | FLUHIGH | 10 | 230 | -60 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9969 | 10 | -59.9984 | 0.50002 | -0.0049452 | 0.0086205 | 2.048 | 99995.0548 | 2015-01-30 13:00 | 49.9997 | 0.44141 | WP- | 0.75293 | VP+ | 0.20866 | -18.9902 | 1.1331 | 1.1333 | 107.9464 | 31943.8997 | 1149.9706 | 0.2094 | -19.3886 | 1.1331 | 1.1328 | 107.9547 | 31944.8807 | 1149.9757 | 0.20864 | -19.6851 | 1.1331 | 1.1328 | 107.9378 | 31943.3561 | 1149.9725 | 0.20685 | -19.235 | 1.1331 | 1.1328 | 107.9507 | 31943.6265 | 1149.9692 | 0.20949 | -19.1589 | 1.1331 | 1.1328 | 107.9291 | 31944.1636 | 1149.9685 | [7 empty] | 1150.0403 | 1991.7995 | 229.9969 | 10 | -59.9984 | 0.0017025 | 0.0012676 | 8.5146e-05 | 5.3307e-06 | 5.1245e-05 | -0.0049452 | 0.0042093 | 27.5966 | 2.048 | 0.0086205 | 0.50002 | 22.9997 | 4.5734e-05 | 99995.0548 | 0 | 0 | 109.127 | 23.5558 | 0.32293
98 | FLUHIGH | 10 | 230 | -60 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 |  |  | 229.9969 | 10.0001 | -59.9987 | 0.50002 | -0.0063493 | 0.0085107 | 2.0525 | 99993.6506 | 2015-01-30 13:04 | 50.0005 | 0.44141 | WP- | 0.75293 | VP+ | 0.20898 | -19.5253 | 1.1331 | 1.1332 | 107.9335 | 31944.0031 | 1149.9685 | 0.20541 | -19.3876 | 1.1331 | 1.1329 | 107.9424 | 31943.6862 | 1149.9686 | 0.20787 | -19.7828 | 1.1331 | 1.1326 | 107.9692 | 31943.1558 | 1149.9666 | 0.20706 | -19.1711 | 1.1331 | 1.1327 | 107.9424 | 31943.4314 | 1149.9712 | 0.20776 | -19.2548 | 1.1331 | 1.1326 | 107.9457 | 31943.2548 | 1149.9696 | [7 empty] | 1150.0392 | 1991.8251 | 229.9969 | 10.0001 | -59.9987 | 0.0018992 | 0.0023467 | 6.1463e-05 | 7.6573e-06 | 6.2747e-05 | -0.0063493 | 0.0041465 | 26.2803 | 2.0525 | 0.0085107 | 0.50002 | 22.9995 | 4.5766e-05 | 99993.6506 | 0 | 0 | 109.127 | 23.5558 | 0.32293
99 | FLUHIGH | 10 | 230 | 36.87 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9968 | 10.0001 | 36.8678 | 0.80002 | 0.0055057 | 0.0056644 | 2.0569 | 100005.5057 | 2015-01-30 13:08 | 50.0009 | 0.70605 | WP- | 0.52148 | VP- | 0.30352 | -140.6413 | 1.1331 | 1.133 | 107.9444 | 51115.2638 | 1840.1517 | 0.30422 | -140.6903 | 1.1331 | 1.1329 | 107.9273 | 51115.3141 | 1840.147 | 0.30477 | -140.9686 | 1.1331 | 1.1328 | 107.9587 | 51114.6502 | 1840.155 | 0.3025 | -140.5732 | 1.1331 | 1.1333 | 107.9625 | 51114.5436 | 1840.1519 | 0.3044 | -140.4768 | 1.1331 | 1.1328 | 107.9462 | 51115.6249 | 1840.1515 | [7 empty] | 1840.0415 | 1379.9247 | 229.9968 | 10.0001 | 36.8678 | 0.0008623 | 0.002556 | 6.0532e-05 | 7.3565e-06 | 5.0649e-05 | 0.0055057 | 0.0027539 | 25.1222 | 2.0569 | 0.0056644 | 0.80002 | 22.9995 | 4.0709e-05 | 100005.5057 | 0 | 0 | 109.124 | 23.5481 | 0.32293
100 | FLUHIGH | 10 | 230 | 36.87 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9967 | 10 | 36.8674 | 0.80003 | 0.0049362 | 0.0056069 | 2.0615 | 100004.9362 | 2015-01-30 13:12 | 50.0002 | 0.70605 | WP- | 0.52148 | VP- | 0.30238 | -140.7607 | 1.1331 | 1.1334 | 107.961 | 51114.8658 | 1840.1511 | 0.30413 | -140.5887 | 1.1331 | 1.1331 | 107.9171 | 51114.6663 | 1840.1509 | 0.30414 | -140.5091 | 1.1331 | 1.1327 | 107.9663 | 51114.9059 | 1840.1494 | 0.30389 | -140.7156 | 1.1331 | 1.133 | 107.9531 | 51114.7604 | 1840.1511 | 0.30554 | -140.8764 | 1.1331 | 1.1332 | 107.9405 | 51114.5249 | 1840.1534 | [7 empty] | 1840.04 | 1379.9056 | 229.9967 | 10 | 36.8674 | 0.0016795 | 0.0018561 | 3.6946e-05 | 5.37e-06 | 5.7333e-05 | 0.0049362 | 0.0027198 | 24.0026 | 2.0615 | 0.0056069 | 0.80003 | 22.9996 | 3.4309e-05 | 100004.9362 | 0 | 0 | 109.123 | 23.5455 | 0.32293
101 | FLUHIGH | 10 | 230 | 36.87 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | watts | 0.0018651 | 6 | 123 |  |  | 229.9967 | 10 | 36.8674 | 0.80003 | 0.00586 | 0.0056255 | 2.0599 | 100005.86 | 2015-01-30 13:17 | 50.0004 | 0.70605 | WP- | 0.52148 | VP- | 0.30237 | -140.3986 | 1.1331 | 1.1334 | 107.9407 | 51115.3019 | 1840.1511 | 0.30412 | -140.4929 | 1.1331 | 1.1331 | 107.9491 | 51115.6002 | 1840.1537 | 0.30424 | -140.8832 | 1.1331 | 1.133 | 107.9404 | 51115.0244 | 1840.1561 | 0.30302 | -140.6866 | 1.1331 | 1.1334 | 107.952 | 51115.0975 | 1840.1473 | 0.30321 | -140.7133 | 1.1331 | 1.1329 | 107.9562 | 51115.1451 | 1840.1478 | [7 empty] | 1840.0406 | 1379.907 | 229.9967 | 10 | 36.8674 | 0.0014272 | 0.0025576 | 6.9024e-05 | 8.4304e-06 | 4.6416e-05 | 0.00586 | 0.002731 | 24.3871 | 2.0599 | 0.0056255 | 0.80003 | 22.9996 | 4.1018e-05 | 100005.86 | 0 | 0 | 109.128 | 23.5584 | 0.32293
102 | FLUHIGH | 6 | 230 | -90 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | vars | 0.0040824 | 6 | 123 |  |  | 229.9967 | 5.9999 | -90.0002 | 1 | 0.0029728 | 0.013797 | 2.0826 | 100002.9728 | 2015-01-30 13:22 | 50.0005 | 0 | WP+ | 0.52148 | VP+ | 0.27668 | 7.5564 | 1.1331 | 1.1334 | 107.9248 | 38333.2859 | 1380.0807 | 0.27435 | 7.2536 | 1.1331 | 1.1329 | 107.9156 | 38333.378 | 1380.0791 | 0.27522 | 7.5007 | 1.1331 | 1.1326 | 107.9329 | 38333.4129 | 1380.0851 | 0.27455 | 7.4282 | 1.1331 | 1.1334 | 107.9301 | 38333.1779 | 1380.0864 | 0.27438 | 7.4899 | 1.1331 | 1.1331 | 107.9253 | 38333.1127 | 1380.0864 | [7 empty] | -0.0038394 | 1379.9568 | 229.9967 | 5.9999 | -90.0002 | 0.0015521 | 0.0020275 | 7.7091e-05 | 7.8174e-06 | 6.4442e-05 | 0.0029728 | 0.0066247 | 20.0088 | 2.0826 | 0.013797 | 1 | 38.3334 | 0.00010362 | 100002.9728 | 0 | 0 | 109.137 | 23.5816 | 0.32293
103 | FLUHIGH | 6 | 230 | -90 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | vars | 0.0040824 | 6 | 123 |  |  | 229.9968 | 5.9999 | -90.0004 | 1 | 0.0021203 | 0.013799 | 2.0825 | 100002.1203 | 2015-01-30 13:26 | 50.0003 | 0 | WP- | 0.52148 | VP+ | 0.26795 | 8.0822 | 1.1331 | 1.1334 | 107.9677 | 38333.1228 | 1380.0824 | 0.26408 | 7.6907 | 1.1331 | 1.1332 | 107.9426 | 38332.869 | 1380.0804 | 0.26798 | 7.6609 | 1.1331 | 1.133 | 107.932 | 38332.8464 | 1380.0782 | 0.26689 | 7.6184 | 1.1331 | 1.1331 | 107.9524 | 38332.8559 | 1380.08 | 0.26693 | 8.1981 | 1.1331 | 1.1332 | 107.9373 | 38332.8789 | 1380.0813 | [7 empty] | -0.0084425 | 1379.9557 | 229.9968 | 5.9999 | -90.0004 | 0.0028742 | 0.0028887 | 9.3316e-05 | 1.2413e-05 | 0.00011933 | 0.0021203 | 0.0066262 | 20.0274 | 2.0825 | 0.013799 | 1 | 38.3335 | 0.00018198 | 100002.1203 | 0 | 0 | 109.141 | 23.5919 | 0.32293
104 | FLUHIGH | 6 | 230 | -90 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | vars | 0.0040824 | 6 | 123 |  |  | 229.9967 | 5.9999 | -90 | 1 | 0.0015831 | 0.013795 | 2.0827 | 100001.5831 | 2015-01-30 13:31 | 50.0004 | 0 | WP+ | 0.52148 | VP+ | 0.27317 | 7.958 | 1.1331 | 1.1328 | 107.9439 | 38332.5305 | 1380.0791 | 0.27556 | 7.6966 | 1.1331 | 1.1333 | 107.9318 | 38332.5261 | 1380.084 | 0.27388 | 7.4706 | 1.1331 | 1.1331 | 107.9275 | 38332.3553 | 1380.0858 | 0.27385 | 8.005 | 1.1331 | 1.1329 | 107.9254 | 38332.4151 | 1380.0881 | 0.2739 | 7.57 | 1.1331 | 1.1325 | 107.9669 | 38332.3921 | 1380.082 | [7 empty] | 0.00049525 | 1379.9461 | 229.9967 | 5.9999 | -90 | 0.0021219 | 0.0025302 | 8.1173e-05 | 9.7783e-06 | 8.8106e-05 | 0.0015831 | 0.0066237 | 19.9966 | 2.0827 | 0.013795 | 1 | 38.3337 | 0.00012712 | 100001.5831 | 0 | 0 | 109.142 | 23.5945 | 0.32293
105 | FLUHIGH | 6 | 230 | -30 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | vars | 0.0046547 | 6 | 123 |  |  | 229.9968 | 6.0003 | -30.0002 | 0.5 | 0.00022777 | 0.014768 | 2.0854 | 100000.2278 | 2015-01-30 13:35 | 50.0004 | 0.45898 | WP- | 0.26074 | VP+ | 0.2891 | -80.2608 | 1.1331 | 1.1333 | 107.9288 | 19167.4755 | 690.0967 | 0.28939 | -80.2695 | 1.1331 | 1.1333 | 107.9347 | 19167.49 | 690.1075 | 0.28802 | -80.3928 | 1.1331 | 1.1328 | 107.9301 | 19167.3125 | 690.1037 | 0.2898 | -80.4396 | 1.1331 | 1.1332 | 107.9488 | 19167.3307 | 690.0991 | 0.28963 | -80.3412 | 1.1331 | 1.1332 | 107.9375 | 19167.4321 | 690.1005 | [7 empty] | 1195.147 | 690.0251 | 229.9968 | 6.0003 | -30.0002 | 0.0011404 | 0.00092333 | 2.6812e-05 | 4.2386e-06 | 5.0598e-05 | 0.00022777 | 0.0070815 | 19.5805 | 2.0854 | 0.014768 | 0.5 | 38.3312 | 7.0393e-05 | 100000.2278 | 0 | 0 | 109.141 | 23.5919 | 0.32293
106 | FLUHIGH | 6 | 230 | -30 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | vars | 0.0046547 | 6 | 123 |  |  | 229.997 | 6.0003 | -30.0003 | 0.5 | -0.00046038 | 0.014779 | 2.0849 | 99999.5396 | 2015-01-30 13:39 | 50.0004 | 0.45898 | WP- | 0.26074 | VP+ | 0.28853 | -80.2561 | 1.1331 | 1.1325 | 107.953 | 19167.4615 | 690.1074 | 0.29102 | -79.9875 | 1.1331 | 1.133 | 107.9296 | 19167.3927 | 690.103 | 0.2896 | -80.3921 | 1.1331 | 1.1327 | 107.944 | 19167.2387 | 690.1091 | 0.28901 | -80.3694 | 1.1331 | 1.1328 | 107.9634 | 19167.3136 | 690.1071 | 0.28941 | -80.3172 | 1.1331 | 1.1334 | 107.9344 | 19167.3569 | 690.093 | [7 empty] | 1195.1483 | 690.0279 | 229.997 | 6.0003 | -30.0003 | 0.0010847 | 0.0023436 | 3.2637e-05 | 3.657e-06 | 0.00010149 | -0.00046038 | 0.0070889 | 19.6619 | 2.0849 | 0.014779 | 0.5 | 38.3312 | 5.5496e-05 | 99999.5396 | 0 | 0 | 109.138 | 23.5842 | 0.32293
107 | FLUHIGH | 6 | 230 | -30 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | vars | 0.0046547 | 6 | 123 |  |  | 229.9968 | 6.0002 | -30 | 0.5 | 0.0011629 | 0.014791 | 2.0844 | 100001.1629 | 2015-01-30 13:43 | 50.0002 | 0.45898 | WP- | 0.26074 | VP+ | 0.28806 | -80.6592 | 1.1331 | 1.133 | 107.9733 | 19167.4005 | 690.1045 | 0.28885 | -80.4097 | 1.1331 | 1.1331 | 107.9582 | 19167.3808 | 690.098 | 0.28833 | -80.5511 | 1.1331 | 1.1329 | 107.9393 | 19167.5157 | 690.0982 | 0.28955 | -80.0521 | 1.1331 | 1.1331 | 107.9536 | 19167.3545 | 690.1022 | 0.28895 | -80.4757 | 1.1331 | 1.1332 | 107.9398 | 19167.4903 | 690.0981 | [7 empty] | 1195.1477 | 690.0194 | 229.9968 | 6.0002 | -30 | 0.0012182 | 0.0023177 | 0.000102 | 4.139e-06 | 9.9079e-05 | 0.0011629 | 0.0070961 | 19.7399 | 2.0844 | 0.014791 | 0.5 | 38.3312 | 7.3512e-05 | 100001.1629 | 0 | 0 | 109.135 | 23.5765 | 0.32293
108 | FLUHIGH | 6 | 230 | -126.87 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | vars | 0.0042105 | 6 | 123 |  | 19.71, 20.24 | 229.9967 | 6 | -126.8721 | 0.79998 | -0.0006426 | 0.014015 | 2.083 | 99999.3574 | 2015-01-30 13:47 | 50.0004 | 0.31738 | WP+ | 0.41797 | VP+ | 0.25916 | -164.1744 | 1.1331 | 1.1329 | 107.9449 | 30664.9989 | 1104.022 | 0.2578 | -164.165 | 1.1331 | 1.1331 | 107.936 | 30664.975 | 1104.0193 | 0.25897 | -164.1232 | 1.1331 | 1.133 | 107.9463 | 30664.8756 | 1104.019 | 0.2601 | -163.7856 | 1.1331 | 1.133 | 107.9474 | 30664.9516 | 1104.027 | 0.25966 | -164.1378 | 1.1331 | 1.1329 | 107.9569 | 30664.8603 | 1104.026 | [7 empty] | -828.0248 | 1103.9447 | 229.9967 | 6 | -126.8721 | 0.0018273 | 0.0012902 | 9.4231e-05 | 3.4939e-06 | 8.5304e-05 | -0.0006426 | 0.0067285 | 19.952 | 2.083 | 0.014015 | 0.79998 | 38.3331 | 6.6559e-05 | 99999.3574 | 0 | 0 | 109.137 | 23.5816 | 0.32293
109 | FLUHIGH | 6 | 230 | -126.87 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | vars | 0.0042105 | 6 | 123 |  |  | 229.9967 | 5.9999 | -126.8726 | 0.79997 | 0.0010238 | 0.014018 | 2.0829 | 100001.0238 | 2015-01-30 13:52 | 50.0001 | 0.31738 | WP+ | 0.41797 | VP+ | 0.26061 | -164.0799 | 1.1331 | 1.133 | 107.927 | 30664.9893 | 1104.0204 | 0.26188 | -164.4246 | 1.1331 | 1.1336 | 107.9628 | 30664.9485 | 1104.0181 | 0.26009 | -164.3567 | 1.1331 | 1.133 | 107.932 | 30664.9936 | 1104.0219 | 0.26186 | -164.5633 | 1.1331 | 1.1335 | 107.9467 | 30664.8919 | 1104.0189 | 0.26295 | -164.3041 | 1.1331 | 1.133 | 107.9334 | 30664.8862 | 1104.0264 | [7 empty] | -828.0249 | 1103.9266 | 229.9967 | 5.9999 | -126.8726 | 0.001706 | 0.002676 | 8.263e-05 | 9.8837e-06 | 8.0877e-05 | 0.0010238 | 0.0067302 | 19.9719 | 2.0829 | 0.014018 | 0.79997 | 38.3335 | 0.00013249 | 100001.0238 | 0 | 0 | 109.139 | 23.5868 | 0.32293
110 | FLUHIGH | 6 | 230 | -126.87 | 230 | C(.2) | 2 | 50 | 10 | 5 | 100000 | vars | 0.0042105 | 6 | 123 |  |  | 229.9969 | 5.9999 | -126.8727 | 0.79997 | 0.00059537 | 0.014017 | 2.0829 | 100000.5954 | 2015-01-30 13:56 | 50.0001 | 0.31738 | WP+ | 0.41797 | VP+ | 0.26264 | -164.2327 | 1.1331 | 1.1334 | 107.9514 | 30664.8025 | 1104.0206 | 0.26234 | -164.201 | 1.1331 | 1.1331 | 107.969 | 30664.6469 | 1104.0175 | 0.26151 | -164.8327 | 1.1331 | 1.1329 | 107.9676 | 30664.6866 | 1104.0269 | 0.26251 | -164.1363 | 1.1331 | 1.1325 | 107.9311 | 30664.761 | 1104.0259 | 0.26273 | -164.3985 | 1.1331 | 1.1332 | 107.9562 | 30664.8235 | 1104.0208 | [7 empty] | -828.0278 | 1103.9242 | 229.9969 | 5.9999 | -126.8727 | 0.0033202 | 0.0011699 | 8.6603e-05 | 1.0758e-05 | 8.2335e-05 | 0.00059537 | 0.0067295 | 19.9631 | 2.0829 | 0.014017 | 0.79997 | 38.3335 | 0.00012916 | 100000.5954 | 0 | 0 | 109.138 | 23.5842 | 0.32293
111 | FLUHIGH | 50 | 230 | 0 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0032656 | 6 | 123 |  |  | 229.9963 | 50.0012 | -0.00042602 | 1 | 0.0034938 | 0.0047403 | 2.0798 | 100003.4938 | 2015-01-30 15:40 | 49.9999 | 0.44141 | WP- | 0 | VP- | 0.16389 | -113.6578 | 1.1331 | 1.1331 | 107.9492 | 319458.5754 | 11500.3389 | 0.16444 | -113.1636 | 1.1331 | 1.1332 | 107.9552 | 319459.0828 | 11500.3623 | 0.1654 | -112.8786 | 1.1331 | 1.1334 | 107.9584 | 319455.3218 | 11500.3037 | 0.16337 | -113.0212 | 1.1331 | 1.133 | 107.9502 | 319457.3882 | 11500.3281 | 0.16333 | -112.8298 | 1.1331 | 1.1333 | 107.9883 | 319459.2781 | 11500.335 | [7 empty] | 11500.0837 | 0.085508 | 229.9963 | 50.0012 | -0.00042602 | 0.023211 | 0.018331 | 0.00011873 | 8.6506e-05 | 9.1329e-05 | 0.0034938 | 0.0022792 | 20.4541 | 2.0798 | 0.0047403 | 1 | 4.5998 | 1.6084e-05 | 100003.4938 | 0 | 0 | 109.182 | 23.6976 | 0.32292
112 | FLUHIGH | 50 | 230 | -60 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0051637 | 6 | 123 |  |  | 229.9965 | 49.9989 | -59.9977 | 0.50003 | -0.0033493 | 0.0086117 | 2.0479 | 99996.6507 | 2015-01-30 15:44 | 49.9999 | 0.2207 | WP- | 0.37695 | VP+ | 0.18338 | 170.6739 | 1.1331 | 1.133 | 107.9247 | 159721.475 | 5749.96 | 0.18228 | 171.1734 | 1.1331 | 1.1334 | 107.9572 | 159723.0741 | 5749.9478 | 0.18205 | 171.6175 | 1.1331 | 1.133 | 107.9621 | 159721.4164 | 5749.9443 | 0.18286 | 170.6307 | 1.1331 | 1.1331 | 107.9646 | 159723.3645 | 5749.9502 | 0.18224 | 170.6675 | 1.1331 | 1.1329 | 107.9375 | 159720.4021 | 5749.9746 | [7 empty] | 5750.1827 | 9958.7012 | 229.9965 | 49.9989 | -59.9977 | 0.02848 | 0.013416 | 5.7144e-05 | 5.116e-05 | 0.00014003 | -0.0033493 | 0.0042051 | 27.6098 | 2.0479 | 0.0086117 | 0.50003 | 4.6 | 8.5822e-06 | 99996.6507 | 0 | 0 | 109.189 | 23.7156 | 0.32292
113 | FLUHIGH | 50 | 230 | 36.87 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0037302 | 6 | 123 |  |  | 229.9969 | 49.999 | 36.8981 | 0.7997 | 0.0072694 | 0.005743 | 2.0514 | 100007.2694 | 2015-01-30 15:48 | 50.0002 | 0.35254 | WP- | 0.26074 | VP- | 0.22596 | 148.2547 | 1.1331 | 1.1331 | 107.9284 | 255468.3535 | 9196.8389 | 0.22562 | 148.7486 | 1.1331 | 1.1335 | 107.9313 | 255472.9241 | 9196.8213 | 0.22555 | 148.0187 | 1.1331 | 1.1332 | 107.9276 | 255470.6008 | 9196.8467 | 0.22528 | 148.3561 | 1.1331 | 1.1332 | 107.9521 | 255470.975 | 9196.7871 | 0.2253 | 148.4791 | 1.1331 | 1.1331 | 107.9738 | 255474.0567 | 9196.874 | [7 empty] | 9196.3012 | 6904.2982 | 229.9969 | 49.999 | 36.8981 | 0.021046 | 0.0071589 | 4.8011e-05 | 5.625e-05 | 8.8512e-05 | 0.0072694 | 0.0027995 | 26.5722 | 2.0514 | 0.005743 | 0.7997 | 4.6 | 1.1583e-05 | 100007.2694 | 0 | 0 | 109.198 | 23.7388 | 0.32292
114 | FLUHIGH | 50 | 230 | 0 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0032656 | 6 | 123 |  |  | 229.9968 | 50.0007 | -0.00065946 | 1 | 0.0045806 | 0.0047635 | 2.077 | 100004.5806 | 2015-01-30 15:52 | 50.0003 | 0.44141 | WP- | 0 | VP+ | 0.16207 | -112.9181 | 1.1331 | 1.1327 | 107.9603 | 319459.5012 | 11500.3223 | 0.163 | -113.0426 | 1.1331 | 1.133 | 107.9534 | 319461.0099 | 11500.2998 | 0.16193 | -113.0756 | 1.1331 | 1.1333 | 107.9672 | 319457.6622 | 11500.3193 | 0.162 | -112.657 | 1.1331 | 1.133 | 107.939 | 319457.4114 | 11500.3271 | 0.16231 | -112.7049 | 1.1331 | 1.133 | 107.9327 | 319460.6173 | 11500.3125 | [7 empty] | 11500.0059 | 0.13236 | 229.9968 | 50.0007 | -0.00065946 | 0.01215 | 0.012108 | 8.6536e-05 | 4.1227e-05 | 6.0324e-05 | 0.0045806 | 0.0022934 | 20.9315 | 2.077 | 0.0047635 | 1 | 4.5999 | 7.4642e-06 | 100004.5806 | 0 | 0 | 109.202 | 23.7491 | 0.32292
115 | FLUHIGH | 50 | 230 | 0 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0032656 | 6 | 123 |  |  | 229.9967 | 50.0008 | -0.0007548 | 1 | 0.0039288 | 0.004763 | 2.0771 | 100003.9288 | 2015-01-30 15:56 | 50.0004 | 0.44141 | WP- | 0 | VP+ | 0.16183 | -112.9463 | 1.1331 | 1.1332 | 107.9283 | 319456.8487 | 11500.3545 | 0.16203 | -112.7353 | 1.1331 | 1.1327 | 107.979 | 319457.6144 | 11500.2861 | 0.16173 | -113.1614 | 1.1331 | 1.1333 | 107.9475 | 319455.8647 | 11500.3662 | 0.1626 | -112.3263 | 1.1331 | 1.1335 | 107.9477 | 319460.0959 | 11500.3438 | 0.16084 | -112.6158 | 1.1331 | 1.133 | 107.9685 | 319456.5698 | 11500.2998 | [7 empty] | 11500.0145 | 0.1515 | 229.9967 | 50.0008 | -0.0007548 | 0.014903 | 0.017042 | 0.00010186 | 5.9208e-05 | 8.4908e-05 | 0.0039288 | 0.0022931 | 20.9221 | 2.0771 | 0.004763 | 1 | 4.5999 | 1.2669e-05 | 100003.9288 | 0 | 0 | 109.201 | 23.7465 | 0.32292
116 | FLUHIGH | 50 | 230 | 0 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0032656 | 6 | 123 |  |  | 229.9967 | 50.0009 | -0.00080974 | 1 | 0.0033405 | 0.0047423 | 2.0796 | 100003.3405 | 2015-01-30 16:00 | 50.0006 | 0.44141 | WP- | 0 | VP+ | 0.16241 | -113.0977 | 1.1331 | 1.1326 | 107.9428 | 319452.9651 | 11500.3447 | 0.16211 | -112.2225 | 1.1331 | 1.1328 | 107.9331 | 319456.1415 | 11500.3525 | 0.16016 | -113.0483 | 1.1331 | 1.133 | 107.9396 | 319457.1489 | 11500.3096 | 0.16033 | -112.6504 | 1.1331 | 1.1335 | 107.9666 | 319456.2354 | 11500.3213 | 0.16116 | -112.5959 | 1.1331 | 1.1335 | 107.9213 | 319458.3593 | 11500.3096 | [7 empty] | 11500.038 | 0.16253 | 229.9967 | 50.0009 | -0.00080974 | 0.021993 | 0.020071 | 4.2684e-05 | 9.3723e-05 | 0.0001 | 0.0033405 | 0.0022804 | 20.4975 | 2.0796 | 0.0047423 | 1 | 4.5999 | 1.9073e-05 | 100003.3405 | 0 | 0 | 109.194 | 23.7285 | 0.32292
117 | FLUHIGH | 50 | 230 | -60 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0051637 | 6 | 123 |  |  | 229.9968 | 49.9993 | -59.9981 | 0.50003 | -0.0017066 | 0.0085902 | 2.0489 | 99998.2934 | 2015-01-30 16:05 | 50 | 0.2207 | WP- | 0.37695 | VP+ | 0.18052 | 171.2273 | 1.1331 | 1.1327 | 107.9482 | 159724.9162 | 5749.9346 | 0.1804 | 171.0717 | 1.1331 | 1.1338 | 107.9387 | 159722.3959 | 5749.9546 | 0.1818 | 171.1741 | 1.1331 | 1.1332 | 107.9623 | 159722.7467 | 5749.937 | 0.18285 | 171.1313 | 1.1331 | 1.1332 | 107.9389 | 159724.5555 | 5749.9351 | 0.17996 | 171.3861 | 1.1331 | 1.1332 | 107.9619 | 159726.0808 | 5749.9272 | [7 empty] | 5750.1671 | 9958.8273 | 229.9968 | 49.9993 | -59.9981 | 0.0048869 | 0.029311 | 6.1368e-05 | 9.9037e-05 | 7.6054e-05 | -0.0017066 | 0.0041927 | 27.3229 | 2.0489 | 0.0085902 | 0.50003 | 4.6 | 1.8035e-05 | 99998.2934 | 0 | 0 | 109.186 | 23.7079 | 0.32292
118 | FLUHIGH | 50 | 230 | -60 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0051637 | 6 | 123 |  |  | 229.9968 | 49.9992 | -59.9981 | 0.50003 | -0.0021244 | 0.0086206 | 2.0475 | 99997.8756 | 2015-01-30 16:09 | 50 | 0.2207 | WP- | 0.37695 | VP+ | 0.18134 | 171.4629 | 1.1331 | 1.1332 | 107.9631 | 159724.2726 | 5749.9316 | 0.18089 | 171.1674 | 1.1331 | 1.133 | 107.9487 | 159724.844 | 5749.9604 | 0.18279 | 171.0951 | 1.1331 | 1.1326 | 107.948 | 159721.3087 | 5749.9561 | 0.1811 | 171.3498 | 1.1331 | 1.133 | 107.9383 | 159724.4482 | 5749.9072 | 0.18319 | 171.2464 | 1.1331 | 1.133 | 107.9519 | 159720.5474 | 5749.9272 | [7 empty] | 5750.1532 | 9958.8072 | 229.9968 | 49.9992 | -59.9981 | 0.0074517 | 0.020782 | 3.7337e-05 | 7.262e-05 | 7.5778e-05 | -0.0021244 | 0.0042102 | 27.7236 | 2.0475 | 0.0086206 | 0.50003 | 4.6 | 1.5718e-05 | 99997.8756 | 0 | 0 | 109.178 | 23.6873 | 0.32292
119 | FLUHIGH | 50 | 230 | -60 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0051637 | 6 | 123 |  |  | 229.9967 | 49.9995 | -59.9982 | 0.50003 | -0.0024477 | 0.0086556 | 2.0462 | 99997.5523 | 2015-01-30 16:13 | 50.0001 | 0.2207 | WP- | 0.37695 | VP+ | 0.18243 | 171.1459 | 1.1331 | 1.1331 | 107.9537 | 159726.4505 | 5749.9453 | 0.18011 | 170.7979 | 1.1331 | 1.1333 | 107.948 | 159721.0633 | 5749.9248 | 0.18104 | 171.0147 | 1.1331 | 1.1332 | 107.9668 | 159723.064 | 5749.9746 | 0.18273 | 171.4048 | 1.1331 | 1.1332 | 107.934 | 159723.8752 | 5749.9404 | 0.18154 | 170.8922 | 1.1331 | 1.1332 | 107.941 | 159722.154 | 5749.9414 | [7 empty] | 5750.1803 | 9958.8702 | 229.9967 | 49.9995 | -59.9982 | 0.014671 | 0.019525 | 8.4476e-05 | 0.00010522 | 3.4425e-05 | -0.0024477 | 0.0042301 | 28.1499 | 2.0462 | 0.0086556 | 0.50003 | 4.6 | 2.2948e-05 | 99997.5523 | 0 | 0 | 109.168 | 23.6615 | 0.32292
120 | FLUHIGH | 50 | 230 | 36.87 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0037302 | 6 | 123 |  |  | 229.9967 | 49.9992 | 36.8983 | 0.7997 | 0.0069043 | 0.0058019 | 2.0486 | 100006.9043 | 2015-01-30 16:17 | 50.0001 | 0.35254 | WP- | 0.26074 | VP- | 0.22481 | 148.4472 | 1.1331 | 1.1334 | 107.9438 | 255469.455 | 9196.8135 | 0.22496 | 148.2971 | 1.1331 | 1.1332 | 107.9646 | 255472.1473 | 9196.8555 | 0.22626 | 148.2027 | 1.1331 | 1.1335 | 107.9456 | 255464.9864 | 9196.8496 | 0.22517 | 148.2793 | 1.1331 | 1.133 | 107.9389 | 255470.0976 | 9196.8594 | 0.22488 | 148.4959 | 1.1331 | 1.1332 | 107.9478 | 255473.1721 | 9196.8281 | [7 empty] | 9196.284 | 6904.3479 | 229.9967 | 49.9992 | 36.8983 | 0.016398 | 0.011878 | 8.7904e-05 | 7.2569e-05 | 2.7564e-05 | 0.0069043 | 0.0028321 | 27.4043 | 2.0486 | 0.0058019 | 0.7997 | 4.6 | 1.3158e-05 | 100006.9043 | 0 | 0 | 109.162 | 23.646 | 0.32292
121 | FLUHIGH | 50 | 230 | 36.87 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0037302 | 6 | 123 |  |  | 229.9969 | 49.9993 | 36.8985 | 0.7997 | 0.0069411 | 0.0057865 | 2.0493 | 100006.9411 | 2015-01-30 16:21 | 50.0003 | 0.35254 | WP- | 0.26074 | VP- | 0.22588 | 148.5694 | 1.1331 | 1.1329 | 107.941 | 255472.8865 | 9196.834 | 0.225 | 148.3217 | 1.1331 | 1.1335 | 107.9503 | 255464.8025 | 9196.8672 | 0.226 | 148.1412 | 1.1331 | 1.1326 | 107.9499 | 255470.2905 | 9196.8574 | 0.2247 | 148.2603 | 1.1331 | 1.1335 | 107.9664 | 255472.1308 | 9196.8262 | 0.22574 | 148.512 | 1.1331 | 1.1332 | 107.9529 | 255473.0588 | 9196.8486 | [7 empty] | 9196.3045 | 6904.3943 | 229.9969 | 49.9993 | 36.8985 | 0.016255 | 0.010908 | 7.3852e-05 | 4.7013e-05 | 7.6674e-05 | 0.0069411 | 0.0028237 | 27.208 | 2.0493 | 0.0057865 | 0.7997 | 4.6 | 9.5821e-06 | 100006.9411 | 0 | 0 | 109.159 | 23.6383 | 0.32293
122 | FLUHIGH | 50 | 230 | 36.87 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0037302 | 6 | 123 |  |  | 229.9969 | 49.9992 | 36.8984 | 0.7997 | 0.0063666 | 0.0056799 | 2.0556 | 100006.3666 | 2015-01-30 16:26 | 50.0001 | 0.35254 | WP- | 0.26074 | VP- | 0.22515 | 148.2993 | 1.1331 | 1.1329 | 107.9326 | 255470.0244 | 9196.8848 | 0.22488 | 148.6027 | 1.1331 | 1.1326 | 107.9435 | 255470.4117 | 9196.8213 | 0.22578 | 148.0629 | 1.1331 | 1.1329 | 107.9493 | 255469.6576 | 9196.8496 | 0.22518 | 148.3936 | 1.1331 | 1.1332 | 107.9564 | 255468.1869 | 9196.7764 | 0.22595 | 148.4162 | 1.1331 | 1.1329 | 107.9722 | 255467.198 | 9196.8682 | [7 empty] | 9196.302 | 6904.3654 | 229.9969 | 49.9992 | 36.8984 | 0.017276 | 0.0089993 | 9.8076e-05 | 6.3561e-05 | 5.394e-05 | 0.0063666 | 0.0027631 | 25.4407 | 2.0556 | 0.0056799 | 0.7997 | 4.6 | 1.3062e-05 | 100006.3666 | 0 | 0 | 109.158 | 23.6357 | 0.32293
123 | FLUHIGH | 100 | 230 | 0 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0032656 | 6 | 123 |  |  | 229.9965 | 100.0017 | -0.0004366 | 1 | 0.0050558 | 0.0047358 | 2.0805 | 100005.0558 | 2015-01-30 17:00 | 50.0002 | 0.88281 | WP- | 0 | VP+ | 0.32272 | -112.3244 | 1.1331 | 1.1335 | 107.9556 | 638915.1882 | 23000.6406 | 0.3241 | -112.4721 | 1.1331 | 1.133 | 107.9582 | 638922.7685 | 23000.6211 | 0.32617 | -112.0978 | 1.1331 | 1.1333 | 107.9479 | 638920.6228 | 23000.6387 | 0.32449 | -112.7112 | 1.1331 | 1.1328 | 107.9522 | 638927.9563 | 23000.627 | 0.32419 | -112.2523 | 1.1331 | 1.1329 | 107.9587 | 638923.3263 | 23000.6445 | [7 empty] | 23000.0282 | 0.17526 | 229.9965 | 100.0017 | -0.0004366 | 0.018674 | 0.025257 | 8.1424e-05 | 9.3835e-05 | 6.2918e-05 | 0.0050558 | 0.0022763 | 20.3544 | 2.0805 | 0.0047358 | 1 | 2.2999 | 5.9803e-06 | 100005.0558 | 0 | 0 | 109.184 | 23.7027 | 0.32292
124 | FLUHIGH | 100 | 230 | -60 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0051637 | 6 | 123 |  |  | 229.9966 | 99.9978 | -59.9978 | 0.50003 | -0.0017474 | 0.0086924 | 2.0458 | 99998.2526 | 2015-01-30 17:04 | 49.9999 | 0.44141 | WP- | 0.75293 | VP+ | 0.20991 | -16.624 | 1.1331 | 1.1328 | 107.9281 | 319444.1546 | 11499.835 | 0.20776 | -16.8576 | 1.1331 | 1.1337 | 107.9756 | 319437.5691 | 11499.8398 | 0.20825 | -16.7628 | 1.1331 | 1.1333 | 107.989 | 319455.8254 | 11499.8428 | 0.20751 | -16.4801 | 1.1331 | 1.1337 | 107.9429 | 319454.8422 | 11499.8262 | 0.20837 | -17.4198 | 1.1331 | 1.1327 | 107.929 | 319450.0269 | 11499.8213 | [7 empty] | 11500.3464 | 19917.4057 | 229.9966 | 99.9978 | -59.9978 | 0.022008 | 0.016644 | 7.2056e-05 | 7.5532e-05 | 5.8578e-05 | -0.0017474 | 0.0042489 | 28.271 | 2.0458 | 0.0086924 | 0.50003 | 2.3 | 4.9948e-06 | 99998.2526 | 0 | 0 | 109.189 | 23.7156 | 0.32292
125 | FLUHIGH | 100 | 230 | 36.87 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0037302 | 6 | 123 |  |  | 229.9966 | 99.9985 | 36.8979 | 0.79971 | 0.0074543 | 0.0056326 | 2.0593 | 100007.4543 | 2015-01-30 17:09 | 50.0005 | 0.70605 | WP- | 0.52246 | VP- | 0.42818 | -80.7955 | 1.1331 | 1.1328 | 107.9555 | 510948.4928 | 18393.752 | 0.42919 | -80.4381 | 1.1331 | 1.1335 | 107.956 | 510947.2595 | 18393.7051 | 0.42783 | -80.9494 | 1.1331 | 1.1328 | 107.9501 | 510944.7469 | 18393.7402 | 0.42717 | -80.6461 | 1.1331 | 1.1329 | 107.9368 | 510942.7767 | 18393.7617 | 0.42837 | -80.6974 | 1.1331 | 1.1326 | 107.9488 | 510949.5747 | 18393.7461 | [7 empty] | 18392.7055 | 13808.57 | 229.9966 | 99.9985 | 36.8979 | 0.012813 | 0.031864 | 3.1825e-05 | 9.1378e-05 | 6.4111e-05 | 0.0074543 | 0.0027351 | 24.5178 | 2.0593 | 0.0056326 | 0.79971 | 2.3 | 4.3953e-06 | 100007.4543 | 0 | 0 | 109.182 | 23.6976 | 0.32292
126 | FLUHIGH | 100 | 230 | 0 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0032656 | 6 | 123 |  |  | 229.9966 | 100.0016 | -0.00041061 | 1 | 0.0048737 | 0.0047043 | 2.0846 | 100004.8737 | 2015-01-30 17:13 | 50.0007 | 0.88281 | WP- | 0 | VP+ | 0.32396 | -112.3677 | 1.1331 | 1.1329 | 107.9589 | 638924.3443 | 23000.6758 | 0.32452 | -112.2684 | 1.1331 | 1.1325 | 107.9478 | 638918.0839 | 23000.5996 | 0.32468 | -112.4888 | 1.1331 | 1.1334 | 107.948 | 638920.3142 | 23000.6133 | 0.32573 | -112.143 | 1.1331 | 1.133 | 107.9324 | 638918.2544 | 23000.6387 | 0.32534 | -112.7522 | 1.1331 | 1.1334 | 107.9438 | 638923.4605 | 23000.6094 | [7 empty] | 23000.0311 | 0.16483 | 229.9966 | 100.0016 | -0.00041061 | 0.027728 | 0.026703 | 9.1664e-05 | 8.8273e-05 | 6.6521e-05 | 0.0048737 | 0.0022567 | 19.7097 | 2.0846 | 0.0047043 | 1 | 2.2999 | 3.3429e-06 | 100004.8737 | 0 | 0 | 109.175 | 23.6795 | 0.32292
127 | FLUHIGH | 100 | 230 | 0 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0032656 | 6 | 123 |  |  | 229.9967 | 100.0015 | -0.00034606 | 1 | 0.0043158 | 0.0047139 | 2.0832 | 100004.3158 | 2015-01-30 17:17 | 50.0001 | 0.88281 | WP- | 0 | VP+ | 0.32551 | -112.232 | 1.1331 | 1.1332 | 107.9537 | 638918.9036 | 23000.6074 | 0.32634 | -112.5649 | 1.1331 | 1.133 | 107.9533 | 638913.4135 | 23000.5703 | 0.32616 | -112.383 | 1.1331 | 1.1329 | 107.9559 | 638918.1977 | 23000.6484 | 0.32589 | -112.5201 | 1.1331 | 1.1329 | 107.9481 | 638919.1779 | 23000.6035 | 0.32571 | -112.616 | 1.1331 | 1.1327 | 107.9507 | 638914.4438 | 23000.625 | [7 empty] | 23000.0131 | 0.13892 | 229.9967 | 100.0015 | -0.00034606 | 0.018511 | 0.018051 | 7.4147e-05 | 5.0547e-05 | 4.4967e-05 | 0.0043158 | 0.0022628 | 19.9139 | 2.0832 | 0.0047139 | 1 | 2.2999 | 1.3703e-06 | 100004.3158 | 0 | 0 | 109.172 | 23.6718 | 0.32292
128 | FLUHIGH | 100 | 230 | 0 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0032656 | 6 | 123 |  |  | 229.9967 | 100.0015 | -0.00049329 | 1 | 0.0037675 | 0.0047715 | 2.0766 | 100003.7675 | 2015-01-30 17:21 | 50.0002 | 0.88281 | WP- | 0 | VP+ | 0.32441 | -112.0489 | 1.1331 | 1.133 | 107.9728 | 638910.0561 | 23000.6309 | 0.32725 | -112.5458 | 1.1331 | 1.1332 | 107.9597 | 638916.1867 | 23000.6309 | 0.32493 | -112.1496 | 1.1331 | 1.133 | 107.9917 | 638920.4162 | 23000.6113 | 0.32602 | -112.2032 | 1.1331 | 1.1333 | 107.9384 | 638904.9049 | 23000.6133 | 0.3258 | -112.1895 | 1.1331 | 1.133 | 107.9533 | 638914.2222 | 23000.5918 | [7 empty] | 23000.0072 | 0.19802 | 229.9967 | 100.0015 | -0.00049329 | 0.03348 | 0.024668 | 0.00013418 | 0.0001018 | 6.1452e-05 | 0.0037675 | 0.0022978 | 21.0035 | 2.0766 | 0.0047715 | 1 | 2.2999 | 4.0955e-06 | 100003.7675 | 0 | 0 | 109.17 | 23.6666 | 0.32292
129 | FLUHIGH | 100 | 230 | -60 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0051637 | 6 | 123 |  |  | 229.9966 | 99.9985 | -59.9981 | 0.50003 | -0.001315 | 0.0085952 | 2.0489 | 99998.685 | 2015-01-30 17:25 | 50.0004 | 0.44141 | WP- | 0.75293 | VP+ | 0.20479 | -17.2417 | 1.1331 | 1.1334 | 107.9341 | 319443.0698 | 11499.8086 | 0.20717 | -16.7066 | 1.1331 | 1.1327 | 107.942 | 319444.1277 | 11499.7861 | 0.20788 | -17.318 | 1.1331 | 1.1325 | 107.9496 | 319452.6204 | 11499.7471 | 0.20312 | -17.0007 | 1.1331 | 1.1335 | 107.9453 | 319453.679 | 11499.8438 | 0.20659 | -17.2797 | 1.1331 | 1.1335 | 107.9225 | 319451.9158 | 11499.7656 | [7 empty] | 11500.3182 | 19917.6044 | 229.9966 | 99.9985 | -59.9981 | 0.015864 | 0.044967 | 6.8004e-05 | 0.00015204 | 7.1089e-05 | -0.001315 | 0.0041951 | 27.3254 | 2.0489 | 0.0085952 | 0.50003 | 2.3 | 7.3444e-06 | 99998.685 | 0 | 0 | 109.168 | 23.6615 | 0.32292
130 | FLUHIGH | 100 | 230 | -60 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0051637 | 6 | 123 |  |  | 229.9967 | 99.998 | -59.9978 | 0.50003 | -0.0018054 | 0.0085497 | 2.0507 | 99998.1946 | 2015-01-30 17:30 | 50 | 0.44141 | WP- | 0.75293 | VP+ | 0.20485 | -16.6457 | 1.1331 | 1.1329 | 107.9627 | 319445.437 | 11499.7451 | 0.20602 | -17.1494 | 1.1331 | 1.1332 | 107.9573 | 319443.8301 | 11499.8359 | 0.20471 | -16.9141 | 1.1331 | 1.1331 | 107.9341 | 319448.8285 | 11499.8047 | 0.20486 | -16.9862 | 1.1331 | 1.1332 | 107.9456 | 319455.7262 | 11499.7988 | 0.2049 | -16.5667 | 1.1331 | 1.1327 | 107.9359 | 319450.2243 | 11499.7744 | [7 empty] | 11500.3648 | 19917.4485 | 229.9967 | 99.998 | -59.9978 | 0.020961 | 0.016583 | 6.7854e-05 | 5.4917e-05 | 4.7851e-05 | -0.0018054 | 0.0041692 | 26.7901 | 2.0507 | 0.0085497 | 0.50003 | 2.3 | 1.8121e-06 | 99998.1946 | 0 | 0 | 109.166 | 23.6563 | 0.32292
131 | FLUHIGH | 100 | 230 | -60 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0051637 | 6 | 123 |  |  | 229.9966 | 99.9982 | -59.9979 | 0.50003 | -0.0016416 | 0.0086253 | 2.0478 | 99998.3584 | 2015-01-30 17:34 | 50.0001 | 0.44141 | WP- | 0.75293 | VP+ | 0.20557 | -16.9441 | 1.1331 | 1.1335 | 107.954 | 319443.3015 | 11499.7734 | 0.20474 | -17.3338 | 1.1331 | 1.1326 | 107.9474 | 319456.669 | 11499.7822 | 0.20594 | -17.0044 | 1.1331 | 1.1326 | 107.9359 | 319448.2196 | 11499.8135 | 0.20529 | -16.3771 | 1.1331 | 1.133 | 107.9367 | 319446.3963 | 11499.7617 | 0.2046 | -16.9635 | 1.1331 | 1.133 | 107.9232 | 319449.7403 | 11499.8018 | [7 empty] | 11500.3479 | 19917.508 | 229.9966 | 99.9982 | -59.9979 | 0.02463 | 0.020457 | 7.4068e-05 | 9.0933e-05 | 5.0193e-05 | -0.0016416 | 0.004212 | 27.6471 | 2.0478 | 0.0086253 | 0.50003 | 2.3 | 4.2745e-06 | 99998.3584 | 0 | 0 | 109.162 | 23.646 | 0.32292
132 | FLUHIGH | 100 | 230 | 36.87 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0037302 | 6 | 123 |  |  | 229.9968 | 99.9988 | 36.8982 | 0.7997 | 0.0060953 | 0.0056727 | 2.0563 | 100006.0953 | 2015-01-30 17:38 | 50.0008 | 0.70605 | WP- | 0.52246 | VP- | 0.42967 | -80.4735 | 1.1331 | 1.1333 | 107.9535 | 510938.1118 | 18393.6172 | 0.42783 | -80.6061 | 1.1331 | 1.1331 | 107.9639 | 510938.7248 | 18393.5469 | 0.43003 | -80.628 | 1.1331 | 1.133 | 107.9681 | 510942.9156 | 18393.6152 | 0.42823 | -80.6466 | 1.1331 | 1.1329 | 107.9463 | 510933.2521 | 18393.6914 | 0.43028 | -80.3326 | 1.1331 | 1.133 | 107.9762 | 510946.0146 | 18393.6934 | [7 empty] | 18392.7118 | 13808.7188 | 229.9968 | 99.9988 | 36.8982 | 0.012449 | 0.023149 | 5.1037e-05 | 9.8656e-05 | 3.8946e-05 | 0.0060953 | 0.0027586 | 25.2649 | 2.0563 | 0.0056727 | 0.7997 | 2.3 | 5.6365e-06 | 100006.0953 | 0 | 0 | 109.16 | 23.6409 | 0.32292
133 | FLUHIGH | 100 | 230 | 36.87 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0037302 | 6 | 123 |  |  | 229.9968 | 99.9985 | 36.8979 | 0.79971 | 0.006004 | 0.0056212 | 2.0603 | 100006.004 | 2015-01-30 17:42 | 50.0006 | 0.70605 | WP- | 0.52246 | VP- | 0.4287 | -80.2964 | 1.1331 | 1.1329 | 107.9453 | 510942.1201 | 18393.6055 | 0.42963 | -80.2482 | 1.1331 | 1.1329 | 107.9741 | 510941.0766 | 18393.6348 | 0.43012 | -80.5248 | 1.1331 | 1.1325 | 107.9464 | 510936.4446 | 18393.5879 | 0.42925 | -80.4821 | 1.1331 | 1.133 | 107.9443 | 510935.7506 | 18393.6348 | 0.43036 | -80.4652 | 1.1331 | 1.1332 | 107.9421 | 510940.9362 | 18393.6309 | [7 empty] | 18392.7093 | 13808.6062 | 229.9968 | 99.9985 | 36.8979 | 0.016988 | 0.013762 | 2.2309e-05 | 4.6345e-05 | 4.7435e-05 | 0.006004 | 0.0027283 | 24.2883 | 2.0603 | 0.0056212 | 0.79971 | 2.3 | 2.468e-06 | 100006.004 | 0 | 0 | 109.158 | 23.6357 | 0.32293
134 | FLUHIGH | 100 | 230 | 36.87 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0037302 | 6 | 123 | Prime EN |  | 229.9967 | 99.9983 | 36.8978 | 0.79971 | 0.0061798 | 0.0056392 | 2.0588 | 100006.1798 | 2015-01-30 17:46 | 50.0004 | 0.70605 | WP- | 0.52246 | VP- | 0.43034 | -80.5246 | 1.1331 | 1.1322 | 107.949 | 510938.581 | 18393.6523 | 0.42787 | -80.3711 | 1.1331 | 1.1329 | 107.9327 | 510946.8591 | 18393.6328 | 0.43004 | -80.3948 | 1.1331 | 1.1327 | 107.9556 | 510940.9052 | 18393.6309 | 0.43025 | -80.3885 | 1.1331 | 1.1332 | 107.9529 | 510937.7347 | 18393.668 | 0.42954 | -80.1674 | 1.1331 | 1.1328 | 107.9629 | 510935.1631 | 18393.5684 | [7 empty] | 18392.6979 | 13808.5156 | 229.9967 | 99.9983 | 36.8978 | 0.018338 | 0.023687 | 8.009e-05 | 8.0593e-05 | 5.1889e-05 | 0.0061798 | 0.002739 | 24.6477 | 2.0588 | 0.0056392 | 0.79971 | 2.3 | 4.1947e-06 | 100006.1798 | 0 | 0 | 109.161 | 23.6435 | 0.32292
135 | FLUHIGH | 0.5 | 230 | 0 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 | 0.043468 | ACal All done & RD31 reset 19.55°C 67.95%. | 229.9964 | 0.50002 | -0.0017565 | 1 | 0.0022317 | 0.0048251 | 2.071 | 100002.2317 | 2015-02-02 11:40 | 50.0004 | 0.44141 | WP- | 0 | VP+ | 0.1571 | -111.7747 | 1.1331 | 1.1326 | 107.9268 | 3194.5934 | 115.0056 | 0.15518 | -111.1023 | 1.1331 | 1.1331 | 107.9633 | 3194.6155 | 115.0061 | 0.15448 | -112.2819 | 1.1331 | 1.1328 | 107.9447 | 3194.5663 | 115.0061 | 0.15562 | -110.5984 | 1.1331 | 1.1337 | 107.9325 | 3194.5824 | 115.006 | 0.15394 | -112.1845 | 1.1331 | 1.1331 | 107.9382 | 3194.5666 | 115.0058 |  | RD31 Ch1 Prime EN | [5 empty] | 115.0025 | 0.0035257 | 229.9964 | 0.50002 | -0.0017565 | 0.0003276 | 0.00037665 | 9.6414e-05 | 1.301e-06 | 0.00018765 | 0.0022317 | 0.0023298 | 22.0207 | 2.071 | 0.0048251 | 1 | 459.9757 | 0.0024722 | 100002.2317 | 0 | 0 | 109.168 | 23.6615 | 0.32292
136 | FLUHIGH | 0.5 | 230 | 0 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 | -0.070005 | Prime Points | 229.9967 | 0.50002 | -0.0019196 | 1 | 0.0014952 | 0.0048089 | 2.0724 | 100001.4952 | 2015-02-02 11:45 | 50.0006 | 0.44141 | WP- | 0 | VP+ | 0.15332 | -110.8374 | 1.1331 | 1.1334 | 107.9499 | 3194.5686 | 115.0063 | 0.15689 | -110.0932 | 1.1331 | 1.1334 | 107.924 | 3194.5461 | 115.0051 | 0.15751 | -111.8779 | 1.1331 | 1.1336 | 107.9489 | 3194.595 | 115.0056 | 0.15319 | -111.3779 | 1.1331 | 1.1333 | 107.9643 | 3194.5505 | 115.0058 | 0.15676 | -112.2643 | 1.1331 | 1.1329 | 107.9582 | 3194.5571 | 115.0059 |  | -0.1321 | [5 empty] | 115.0026 | 0.003853 | 229.9967 | 0.50002 | -0.0019196 | 0.00035951 | 0.0004957 | 8.1263e-05 | 1.6672e-06 | 0.00024696 | 0.0014952 | 0.0023204 | 21.7583 | 2.0724 | 0.0048089 | 1 | 459.9764 | 0.0036672 | 100001.4952 | 0 | 0 | 109.173 | 23.6744 | 0.32292
137 | FLUHIGH | 1 | 230 | -60 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 | -0.074219 |  | 229.9966 | 0.99999 | -59.9985 | 0.50002 | -0.0063851 | 0.008529 | 2.0516 | 99993.6149 | 2015-02-02 11:49 | 50.0003 | 0.44141 | WP- | 0.75293 | VP+ | 0.2096 | -16.9745 | 1.1331 | 1.1328 | 107.9699 | 3194.3223 | 114.9922 | 0.21337 | -16.6047 | 1.1331 | 1.1333 | 107.9691 | 3194.3698 | 114.9918 | 0.21314 | -18.8422 | 1.1331 | 1.1331 | 107.9352 | 3194.2748 | 114.9914 | 0.21375 | -17.0286 | 1.1331 | 1.133 | 107.9519 | 3194.3221 | 114.9925 | 0.21495 | -17.5091 | 1.1331 | 1.1333 | 107.9522 | 3194.2425 | 114.992 |  | -0.19397 | [5 empty] | 115.0024 | 199.1782 | 229.9966 | 0.99999 | -59.9985 | 0.00069189 | 0.00050509 | 8.4563e-05 | 2.2924e-06 | 0.00015274 | -0.0063851 | 0.0041572 | 26.5257 | 2.0516 | 0.008529 | 0.50002 | 229.9987 | 0.0010304 | 99993.6149 | 0 | 0 | 109.178 | 23.6873 | 0.32292
138 | FLUHIGH | 1 | 230 | -60 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 | -0.016984 |  | 229.9968 | 0.99999 | -59.9986 | 0.50002 | -0.0064599 | 0.0085753 | 2.0496 | 99993.5401 | 2015-02-02 11:53 | 50.0004 | 0.44141 | WP- | 0.75293 | VP+ | 0.21175 | -18.4404 | 1.1331 | 1.1332 | 107.9339 | 3194.3321 | 114.9913 | 0.21471 | -17.6126 | 1.1331 | 1.1324 | 107.9567 | 3194.3063 | 114.992 | 0.21331 | -17.2314 | 1.1331 | 1.133 | 107.9601 | 3194.2645 | 114.9922 | 0.21041 | -17.5561 | 1.1331 | 1.1329 | 107.9367 | 3194.336 | 114.992 | 0.21279 | -16.8937 | 1.1331 | 1.1332 | 107.9586 | 3194.2725 | 114.992 |  | 0.021367 | [5 empty] | 115.0023 | 199.1789 | 229.9968 | 0.99999 | -59.9986 | 0.00037003 | 0.00038119 | 6.7403e-05 | 1.7356e-06 | 8.7255e-05 | -0.0064599 | 0.0041839 | 27.099 | 2.0496 | 0.0085753 | 0.50002 | 229.9985 | 0.00091713 | 99993.5401 | 0 | 0 | 109.173 | 23.6744 | 0.32292
139 | FLUHIGH | 1 | 230 | -90 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0040824 | 6 | 123 | 0.60513 |  | 229.9967 | 0.99996 | -90 | 1 | 0.010819 | 0.013802 | 2.0826 | 100010.8192 | 2015-02-02 12:07 | 50.0001 | 0 | WP- | 0.87012 | VP+ | 0.21062 | 120.9148 | 1.1331 | 1.1327 | 107.9403 | 6389.2555 | 230.0104 | 0.20722 | 120.4564 | 1.1331 | 1.133 | 107.9676 | 6389.2529 | 230.0101 | 0.21289 | 120.8034 | 1.1331 | 1.1333 | 107.9663 | 6389.243 | 230.011 | 0.20747 | 120.5812 | 1.1331 | 1.1327 | 107.9314 | 6389.2177 | 230.011 | 0.2074 | 120.2833 | 1.1331 | 1.1326 | 107.922 | 6389.222 | 230.0106 |  | 0.010976 | [5 empty] | 5.0091e-05 | 229.9877 | 229.9967 | 0.99996 | -90 | 0.00029391 | 0.00052673 | 5.0146e-05 | 2.1443e-06 | 7.3223e-05 | 0.010819 | 0.0066273 | 20.0038 | 2.0826 | 0.013802 | 1 | 230.0057 | 0.0010345 | 100010.8192 | 0 | 0 | 109.166 | 23.6563 | 0.32292
140 | FLUHIGH | 1 | 230 | -90 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0040824 | 6 | 123 | 0.53744 |  | 229.9966 | 0.99995 | -89.9998 | 1 | 0.01039 | 0.013796 | 2.083 | 100010.3897 | 2015-02-02 12:11 | 50.0001 | 0 | WP+ | 0.87012 | VP+ | 0.20429 | 119.4435 | 1.1331 | 1.1326 | 107.9619 | 6389.1886 | 230.0113 | 0.20716 | 118.7185 | 1.1331 | 1.1332 | 107.9499 | 6389.1641 | 230.0094 | 0.20763 | 119.3169 | 1.1331 | 1.133 | 107.9355 | 6389.1233 | 230.0103 | 0.21101 | 119.35 | 1.1331 | 1.1333 | 107.9483 | 6389.1288 | 230.0102 | 0.2054 | 120.5782 | 1.1331 | 1.1337 | 107.9562 | 6389.106 | 230.0102 |  | 0.034971 | [5 empty] | 0.00060469 | 229.9852 | 229.9966 | 0.99995 | -89.9998 | 0.00059009 | 0.00041967 | 7.6179e-05 | 1.773e-06 | 0.00014701 | 0.01039 | 0.006623 | 19.9512 | 2.083 | 0.013796 | 1 | 230.008 | 0.00091674 | 100010.3897 | 0 | 0 | 109.164 | 23.6512 | 0.32292
141 | FLUHIGH | 1 | 230 | -30 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0046547 | 6 | 123 | 0.41772 |  | 229.9966 | 1 | -29.9997 | 0.5 | 0.0085251 | 0.01475 | 2.0862 | 100008.5251 | 2015-02-02 12:15 | 50.0005 | 0.76465 | WP- | 0.43457 | VP+ | 0.30735 | -67.6675 | 1.1331 | 1.1325 | 107.9496 | 3194.7445 | 115.0184 | 0.31191 | -67.1549 | 1.1331 | 1.1338 | 107.959 | 3194.7577 | 115.0191 | 0.3086 | -67.5377 | 1.1331 | 1.1331 | 107.9359 | 3194.7339 | 115.0184 | 0.30842 | -67.2479 | 1.1331 | 1.1331 | 107.9616 | 3194.7378 | 115.0188 | 0.30721 | -67.8338 | 1.1331 | 1.1328 | 107.9441 | 3194.7329 | 115.0182 |  | 0.058637 | [5 empty] | 199.1895 | 115.0009 | 229.9966 | 1 | -29.9997 | 0.00019748 | 0.00044837 | 4.9149e-05 | 6.7468e-07 | 0.00011525 | 0.0085251 | 0.0070702 | 19.469 | 2.0862 | 0.01475 | 0.5 | 229.9896 | 0.00036742 | 100008.5251 | 0 | 0 | 109.16 | 23.6409 | 0.32292
142 | FLUHIGH | 1 | 230 | -30 | 230 | ESI(1) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0046547 | 6 | 123 | 0.60196 |  | 229.9968 | 1 | -29.9994 | 0.49999 | 0.0094353 | 0.014762 | 2.0856 | 100009.4353 | 2015-02-02 12:19 | 49.9997 | 0.76465 | WP- | 0.43457 | VP+ | 0.31158 | -67.6293 | 1.1331 | 1.1329 | 107.9266 | 3194.734 | 115.019 | 0.30783 | -67.2153 | 1.1331 | 1.133 | 107.9281 | 3194.7326 | 115.0186 | 0.31327 | -67.5279 | 1.1331 | 1.133 | 107.9591 | 3194.7266 | 115.0182 | 0.30741 | -67.8749 | 1.1331 | 1.1336 | 107.9452 | 3194.7347 | 115.0184 | 0.30676 | -67.6425 | 1.1331 | 1.1329 | 107.9247 | 3194.7352 | 115.0192 |  | 0.21704 | [5 empty] | 199.1895 | 114.9995 | 229.9968 | 1 | -29.9994 | 0.0004286 | 0.00046569 | 7.4577e-05 | 1.3973e-06 | 0.00013592 | 0.0094353 | 0.0070779 | 19.5517 | 2.0856 | 0.014762 | 0.49999 | 229.9907 | 0.00075764 | 100009.4353 | 0 | 0 | 109.157 | 23.6331 | 0.32293
143 | FLUHIGH | 2 | 230 | -126.87 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0042105 | 6 | 123 | 0.62329 |  | 229.9965 | 2 | -126.8733 | 0.79996 | 0.0067795 | 0.014023 | 2.0822 | 100006.7795 | 2015-02-02 12:35 | 50.0002 | 0.21191 | WP+ | 0.27832 | VP+ | 0.11172 | 58.9034 | 1.1331 | 1.1338 | 107.9384 | 10222.1631 | 368.0102 | 0.1103 | 59.7886 | 1.1331 | 1.1332 | 107.9449 | 10222.1658 | 368.008 | 0.11044 | 59.4687 | 1.1331 | 1.1327 | 107.9676 | 10222.144 | 368.0063 | 0.11147 | 60.1789 | 1.1331 | 1.1332 | 107.9452 | 10222.1208 | 368.0081 | 0.1113 | 59.8752 | 1.1331 | 1.1338 | 107.9564 | 10222.0972 | 368.004 |  | -0.050402 | [5 empty] | -276.0132 | 367.972 | 229.9965 | 2 | -126.8733 | 0.00067464 | 0.00066377 | 0.00010117 | 3.2239e-06 | 8.9941e-05 | 0.0067795 | 0.006735 | 20.0812 | 2.0822 | 0.014023 | 0.79996 | 115.0001 | 0.00050224 | 100006.7795 | 0 | 0 | 109.163 | 23.6486 | 0.32292
144 | FLUHIGH | 2 | 230 | -126.87 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0042105 | 6 | 123 | 0.53129 |  | 229.9967 | 2 | -126.8736 | 0.79996 | 0.0061042 | 0.014027 | 2.082 | 100006.1042 | 2015-02-02 12:40 | 50.0001 | 0.21191 | WP+ | 0.27832 | VP+ | 0.11021 | 59.9242 | 1.1331 | 1.1333 | 107.9429 | 10221.9535 | 368.0044 | 0.11169 | 59.7326 | 1.1331 | 1.1331 | 107.9628 | 10222.043 | 368.012 | 0.11131 | 60.1692 | 1.1331 | 1.1336 | 107.9699 | 10221.9928 | 368.005 | 0.10897 | 60.2358 | 1.1331 | 1.133 | 107.9427 | 10221.9977 | 368.0038 | 0.11037 | 60.3605 | 1.1331 | 1.1336 | 107.9621 | 10222.0062 | 368.0065 |  | 0.0087753 | [5 empty] | -276.0142 | 367.9695 | 229.9967 | 2 | -126.8736 | 0.00090021 | 0.00092202 | 6.9402e-05 | 2.2126e-06 | 0.00013976 | 0.0061042 | 0.0067371 | 20.1061 | 2.082 | 0.014027 | 0.79996 | 115.0006 | 0.00022003 | 100006.1042 | 0 | 0 | 109.165 | 23.6538 | 0.32292
145 | FLUHIGH | 5 | 230 | -30 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0046547 | 6 | 123 | 0.44397 |  | 229.9968 | 5.0002 | -29.9997 | 0.5 | 0.0059107 | 0.014757 | 2.0858 | 100005.9107 | 2015-02-02 12:44 | 50.0008 | 0.76465 | WP- | 0.43457 | VP+ | 0.31659 | -67.2104 | 1.1331 | 1.1333 | 107.9547 | 15973.6497 | 575.084 | 0.31242 | -67.7918 | 1.1331 | 1.133 | 107.943 | 15973.5703 | 575.0911 | 0.31259 | -66.9624 | 1.1331 | 1.1326 | 107.9611 | 15973.5478 | 575.0953 | 0.31318 | -67.8614 | 1.1331 | 1.1333 | 107.9627 | 15973.6239 | 575.0828 | 0.31193 | -67.6967 | 1.1331 | 1.133 | 107.951 | 15973.5313 | 575.0878 |  | -0.040688 | [5 empty] | 995.9662 | 575.0151 | 229.9968 | 5.0002 | -29.9997 | 0.0012025 | 0.0025504 | 6.6621e-05 | 7.4821e-06 | 0.00011243 | 0.0059107 | 0.0070751 | 19.5214 | 2.0858 | 0.014757 | 0.5 | 45.9972 | 0.00016214 | 100005.9107 | 0 | 0 | 109.165 | 23.6538 | 0.32292
146 | FLUHIGH | 5 | 230 | -30 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | vars | 0.0046547 | 6 | 123 | 0.45489 |  | 229.9967 | 5.0002 | -29.9997 | 0.5 | 0.0065939 | 0.014751 | 2.0861 | 100006.5939 | 2015-02-02 12:48 | 50.0006 | 0.76465 | WP- | 0.43457 | VP+ | 0.31632 | -67.3538 | 1.1331 | 1.1334 | 107.9541 | 15973.7499 | 575.0896 | 0.31503 | -67.8126 | 1.1331 | 1.133 | 107.9422 | 15973.7164 | 575.0867 | 0.3127 | -67.3173 | 1.1331 | 1.1328 | 107.9634 | 15973.6783 | 575.0869 | 0.31156 | -67.1985 | 1.1331 | 1.1331 | 107.9514 | 15973.6621 | 575.0878 | 0.31393 | -67.3136 | 1.1331 | 1.1331 | 107.9508 | 15973.5838 | 575.0864 |  | -0.063033 | [5 empty] | 995.965 | 575.0145 | 229.9967 | 5.0002 | -29.9997 | 0.002428 | 0.0012692 | 0.00011162 | 8.5226e-06 | 7.486e-05 | 0.0065939 | 0.0070708 | 19.4743 | 2.0861 | 0.014751 | 0.5 | 45.9972 | 0.00015642 | 100006.5939 | 0 | 0 | 109.165 | 23.6538 | 0.32292
147 | FLUHIGH | 2 | 230 | 0 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0006 | 6 | 123 | -0.33121 |  | 229.9966 | 2.0001 | -0.0020081 | 1 | 0.0019802 | 0.0037458 | 2.0076 | 100001.9802 | 2015-02-02 13:09 | 50 | 0.35352 | WP- | 0 | VP+ | 0.348 | -104.795 | 1.1331 | 1.1334 | 107.9501 | 12778.4515 | 460.0312 | 0.34747 | -104.7617 | 1.1331 | 1.1327 | 107.9445 | 12778.4964 | 460.027 | 0.34796 | -104.8275 | 1.1331 | 1.1331 | 107.9389 | 12778.4147 | 460.0342 | 0.35013 | -104.9191 | 1.1331 | 1.1335 | 107.9505 | 12778.6985 | 460.0297 | 0.34776 | -104.661 | 1.1331 | 1.133 | 107.9605 | 12778.2287 | 460.0316 |  | -0.088956 | [5 empty] | 460.0154 | 0.016122 | 229.9966 | 2.0001 | -0.0020081 | 0.00093247 | 0.00061026 | 0.00010833 | 4.2195e-06 | 7.6006e-05 | 0.0019802 | 0.0018658 | 50.4626 | 2.0076 | 0.0037458 | 1 | 114.9927 | 0.00058837 | 100001.9802 | 0 | 0 | 109.162 | 23.646 | 0.32292
148 | FLUHIGH | 2 | 230 | 0 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 | -0.086937 |  | 229.9967 | 2.0001 | -0.0023884 | 1 | 0.0025207 | 0.0048318 | 2.0699 | 100002.5207 | 2015-02-02 13:13 | 50.0003 | 0.35352 | WP- | 0 | VP+ | 0.34944 | -104.7827 | 1.1331 | 1.133 | 107.9725 | 12778.6372 | 460.0287 | 0.34923 | -104.4804 | 1.1331 | 1.1329 | 107.9385 | 12778.3801 | 460.0306 | 0.35108 | -104.6451 | 1.1331 | 1.1337 | 107.9426 | 12778.4444 | 460.0264 | 0.35 | -104.6635 | 1.1331 | 1.1328 | 107.9672 | 12778.3948 | 460.0298 | 0.35004 | -104.1819 | 1.1331 | 1.1326 | 107.9567 | 12778.2809 | 460.028 |  | -0.006693 | [5 empty] | 460.0118 | 0.019176 | 229.9967 | 2.0001 | -0.0023884 | 0.00062767 | 0.0014769 | 6.9423e-05 | 2.9056e-06 | 0.00018395 | 0.0025207 | 0.0023344 | 22.2432 | 2.0699 | 0.0048318 | 1 | 114.9937 | 0.0004099 | 100002.5207 | 0 | 0 | 109.166 | 23.6563 | 0.32292
149 | FLUHIGH | 2 | 230 | -60 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 | 0.07858 |  | 229.9968 | 2 | -59.9998 | 0.5 | -0.0031341 | 0.0086862 | 2.0449 | 99996.8659 | 2015-02-02 13:18 | 49.9999 | 0.17676 | WP- | 0.30176 | VP+ | 0.2859 | -165.584 | 1.1331 | 1.1331 | 107.9477 | 6388.7599 | 229.9881 | 0.28629 | -165.454 | 1.1331 | 1.1333 | 107.9435 | 6388.6863 | 229.9938 | 0.28507 | -165.3276 | 1.1331 | 1.1334 | 107.9574 | 6388.6778 | 229.9881 | 0.28532 | -165.5973 | 1.1331 | 1.1331 | 107.9527 | 6388.7085 | 229.9884 | 0.28555 | -165.4114 | 1.1331 | 1.1329 | 107.9227 | 6388.4754 | 229.9936 |  | 0.16725 | [5 empty] | 229.999 | 398.3668 | 229.9968 | 2 | -59.9998 | 0.00060306 | 0.00068303 | 6.1888e-05 | 2.5018e-06 | 7.5275e-05 | -0.0031341 | 0.0042478 | 28.5673 | 2.0449 | 0.0086862 | 0.5 | 114.998 | 0.00026683 | 99996.8659 | 0 | 0 | 109.165 | 23.6538 | 0.32292
150 | FLUHIGH | 2 | 230 | -60 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 | -0.5263 |  | 229.9968 | 2 | -59.9997 | 0.5 | -0.0070747 | 0.0088786 | 2.0396 | 99992.9253 | 2015-02-02 13:22 | 50.0002 | 0.17676 | WP- | 0.30176 | VP+ | 0.28356 | -165.9128 | 1.1331 | 1.1324 | 107.9568 | 6388.6528 | 229.9915 | 0.28283 | -165.5625 | 1.1331 | 1.1329 | 107.9506 | 6388.4815 | 229.99 | 0.2841 | -166.0921 | 1.1331 | 1.1333 | 107.9419 | 6388.38 | 229.9933 | 0.28439 | -165.2471 | 1.1331 | 1.1329 | 107.9437 | 6388.5904 | 229.9904 | 0.28474 | -165.5433 | 1.1331 | 1.1334 | 107.9556 | 6388.4795 | 229.9908 |  | -0.26254 | [5 empty] | 230.0029 | 398.3719 | 229.9968 | 2 | -59.9997 | 0.0016146 | 0.00073461 | 5.7676e-05 | 3.0278e-06 | 0.0002074 | -0.0070747 | 0.0043531 | 30.4079 | 2.0396 | 0.0088786 | 0.5 | 114.9964 | 0.00044105 | 99992.9253 | 0 | 0 | 109.163 | 23.6486 | 0.32292
151 | FLUHIGH | 5 | 230 | 0 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 | -0.11597 |  | 229.9969 | 5.0002 | -0.00097672 | 1 | 0.0024377 | 0.0047223 | 2.0821 | 100002.4377 | 2015-02-02 13:26 | 50.0003 | 0.88281 | WP- | 0 | VP+ | 0.31839 | -110.6766 | 1.1331 | 1.1333 | 107.942 | 31946.5866 | 1150.0649 | 0.31899 | -110.6053 | 1.1331 | 1.1327 | 107.9511 | 31946.1941 | 1150.063 | 0.32109 | -110.4279 | 1.1331 | 1.1333 | 107.9223 | 31946.1379 | 1150.0642 | 0.31921 | -111.1738 | 1.1331 | 1.1328 | 107.9379 | 31945.9688 | 1150.0646 | 0.31943 | -110.7474 | 1.1331 | 1.1325 | 107.9665 | 31946.1889 | 1150.0685 |  | -0.19853 | [5 empty] | 1150.0357 | 0.019605 | 229.9969 | 5.0002 | -0.00097672 | 0.0015532 | 0.0015025 | 6.6204e-05 | 5.7729e-06 | 7.4856e-05 | 0.0024377 | 0.0022681 | 20.0896 | 2.0821 | 0.0047223 | 1 | 45.9973 | 0.00010324 | 100002.4377 | 0 | 0 | 109.159 | 23.6383 | 0.32293
152 | FLUHIGH | 5 | 230 | 0 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0016328 | 6 | 123 | -0.10931 |  | 229.9969 | 5.0002 | -0.001116 | 1 | 0.0020661 | 0.0047115 | 2.0835 | 100002.0661 | 2015-02-02 13:30 | 50.0001 | 0.88281 | WP- | 0 | VP+ | 0.31833 | -111.0715 | 1.1331 | 1.133 | 107.9661 | 31946.1437 | 1150.0671 | 0.31964 | -110.4926 | 1.1331 | 1.1329 | 107.9558 | 31946.0543 | 1150.0614 | 0.31499 | -110.226 | 1.1331 | 1.1328 | 107.9628 | 31946.0914 | 1150.0671 | 0.31805 | -110.4522 | 1.1331 | 1.1334 | 107.9562 | 31946.4303 | 1150.0619 | 0.32046 | -110.4013 | 1.1331 | 1.1327 | 107.905 | 31946.0287 | 1150.0625 |  | -0.22284 | [5 empty] | 1150.0376 | 0.022401 | 229.9969 | 5.0002 | -0.001116 | 0.0017887 | 0.0019255 | 0.00010081 | 8.0953e-06 | 9.5929e-05 | 0.0020661 | 0.0022613 | 19.8672 | 2.0835 | 0.0047115 | 1 | 45.9972 | 0.0001842 | 100002.0661 | 0 | 0 | 109.155 | 23.628 | 0.32293
153 | FLUHIGH | 5 | 230 | -60 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 | -0.17373 |  | 229.997 | 5 | -59.9985 | 0.50002 | -0.0060693 | 0.0086233 | 2.0479 | 99993.9307 | 2015-02-02 13:34 | 50.0001 | 0.44141 | WP- | 0.75293 | VP+ | 0.20976 | -19.5476 | 1.1331 | 1.1328 | 107.9471 | 15972.0692 | 574.9789 | 0.21225 | -19.0807 | 1.1331 | 1.1333 | 107.9475 | 15971.713 | 574.9796 | 0.20955 | -19.2198 | 1.1331 | 1.1329 | 107.9346 | 15971.4723 | 574.9825 | 0.21099 | -19.4437 | 1.1331 | 1.1332 | 107.9686 | 15972.0808 | 574.9855 | 0.21184 | -18.8704 | 1.1331 | 1.133 | 107.933 | 15971.7501 | 574.9817 |  | -0.077329 | [5 empty] | 575.0203 | 995.9055 | 229.997 | 5 | -59.9985 | 0.0009389 | 0.0013558 | 7.903e-05 | 6.2243e-06 | 3.4034e-05 | -0.0060693 | 0.0042109 | 27.6259 | 2.0479 | 0.0086233 | 0.50002 | 45.9992 | 0.00013138 | 99993.9307 | 0 | 0 | 109.153 | 23.6228 | 0.32293
154 | FLUHIGH | 5 | 230 | -60 | 230 | C(.2) | 1 | 50 | 10 | 5 | 100000 | watts | 0.0025819 | 6 | 123 | -0.16859 |  | 229.9969 | 5.0001 | -59.9986 | 0.50002 | -0.005637 | 0.0084902 | 2.0535 | 99994.363 | 2015-02-02 13:38 | 50.0006 | 0.44141 | WP- | 0.75293 | VP+ | 0.21067 | -18.5903 | 1.1331 | 1.1332 | 107.9477 | 15971.8457 | 574.9781 | 0.21044 | -19.6532 | 1.1331 | 1.1333 | 107.9541 | 15971.9252 | 574.981 | 0.21184 | -18.6224 | 1.1331 | 1.1328 | 107.9633 | 15972.1889 | 574.9792 | 0.21105 | -18.2103 | 1.1331 | 1.1328 | 107.9768 | 15972.0365 | 574.9818 | 0.20793 | -18.6645 | 1.1331 | 1.1329 | 107.9458 | 15971.9544 | 574.9823 |  | -0.12595 | [5 empty] | 575.0241 | 995.9151 | 229.9969 | 5.0001 | -59.9986 | 0.0015795 | 0.0020978 | 0.00011113 | 7.7673e-06 | 8.69e-05 | -0.005637 | 0.0041345 | 25.9957 | 2.0535 | 0.0084902 | 0.50002 | 45.9988 | 0.00015866 | 99994.363 | 0 | 0 | 109.149 | 23.6125 | 0.32293
155 | FLUHIGH | 50 | 230 | 0 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0032656 | 6 | 123 | 0.14568 |  | 229.9965 | 50.0006 | -0.00058062 | 1 | 0.0046293 | 0.0048084 | 2.0724 | 100004.6293 | 2015-02-02 13:57 | 50.0001 | 0.44141 | WP- | 0 | VP+ | 0.16272 | -112.7428 | 1.1331 | 1.133 | 107.9431 | 319460.0155 | 11500.2363 | 0.1619 | -113.0438 | 1.1331 | 1.1331 | 107.9311 | 319454.8736 | 11500.3115 | 0.16227 | -112.7996 | 1.1331 | 1.133 | 107.9345 | 319458.044 | 11500.251 | 0.16276 | -113.4211 | 1.1331 | 1.1332 | 107.9407 | 319456.1368 | 11500.2637 | 0.16195 | -113.0403 | 1.1331 | 1.1326 | 107.9616 | 319463.2474 | 11500.2363 |  | -0.096719 | [5 empty] | 11499.9723 | 0.11654 | 229.9965 | 50.0006 | -0.00058062 | 0.0082911 | 0.015497 | 1.892e-05 | 3.5911e-05 | 7.721e-05 | 0.0046293 | 0.0023202 | 21.7506 | 2.0724 | 0.0048084 | 1 | 4.5999 | 7.4561e-06 | 100004.6293 | 0 | 0 | 109.142 | 23.5945 | 0.32293
156 | FLUHIGH | 50 | 230 | 0 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0032656 | 6 | 123 | 0.16518 |  | 229.9967 | 50.0008 | -0.00054853 | 1 | 0.0041269 | 0.0047614 | 2.0772 | 100004.1269 | 2015-02-02 14:01 | 50.0007 | 0.44141 | WP- | 0 | VP+ | 0.16216 | -113.1289 | 1.1331 | 1.1332 | 107.9355 | 319458.5662 | 11500.252 | 0.16246 | -113.0748 | 1.1331 | 1.1332 | 107.9454 | 319460.2842 | 11500.2402 | 0.16136 | -113.1273 | 1.1331 | 1.1331 | 107.9387 | 319459.7621 | 11500.2549 | 0.16096 | -113.1003 | 1.1331 | 1.1332 | 107.9415 | 319456.3294 | 11500.2344 | 0.16167 | -113.0891 | 1.1331 | 1.1333 | 107.9491 | 319456.9048 | 11500.2666 |  | -0.10854 | [5 empty] | 11500.0267 | 0.1101 | 229.9967 | 50.0008 | -0.00054853 | 0.012129 | 0.0023458 | 2.37e-05 | 5.2371e-05 | 1.1687e-05 | 0.0041269 | 0.0022922 | 20.8907 | 2.0772 | 0.0047614 | 1 | 4.5999 | 1.0804e-05 | 100004.1269 | 0 | 0 | 109.144 | 23.5996 | 0.32293
157 | FLUHIGH | 50 | 230 | -60 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0051637 | 6 | 123 | -0.15291 |  | 229.9966 | 49.999 | -59.9977 | 0.50003 | -0.0030354 | 0.0086901 | 2.045 | 99996.9646 | 2015-02-02 14:05 | 50.0002 | 0.2207 | WP- | 0.37695 | VP+ | 0.18833 | 170.674 | 1.1331 | 1.1331 | 107.9633 | 159720.2497 | 5749.8975 | 0.18597 | 170.949 | 1.1331 | 1.1329 | 107.9525 | 159721.4626 | 5749.8877 | 0.18785 | 170.8247 | 1.1331 | 1.1334 | 107.9576 | 159721.4618 | 5749.9087 | 0.18577 | 170.6486 | 1.1331 | 1.1328 | 107.9464 | 159727.749 | 5749.8706 | 0.18648 | 171.14 | 1.1331 | 1.1328 | 107.9488 | 159723.4276 | 5749.8848 |  | -0.15921 | [5 empty] | 5750.1979 | 9958.7201 | 229.9966 | 49.999 | -59.9977 | 0.014334 | 0.024013 | 8.4803e-05 | 8.6263e-05 | 9.2506e-05 | -0.0030354 | 0.0042495 | 28.5356 | 2.045 | 0.0086901 | 0.50003 | 4.6 | 1.7569e-05 | 99996.9646 | 0 | 0 | 109.141 | 23.5919 | 0.32293
158 | FLUHIGH | 50 | 230 | -60 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0051637 | 6 | 123 | -0.063156 |  | 229.9967 | 49.9986 | -59.9975 | 0.50004 | -0.0026794 | 0.0087882 | 2.0423 | 99997.3206 | 2015-02-02 14:09 | 49.9998 | 0.2207 | WP- | 0.37695 | VP+ | 0.18806 | 171.0918 | 1.1331 | 1.1332 | 107.9236 | 159724.1852 | 5749.8813 | 0.18756 | 171.0517 | 1.1331 | 1.133 | 107.95 | 159722.0184 | 5749.8848 | 0.18815 | 170.5227 | 1.1331 | 1.1329 | 107.9385 | 159719.8619 | 5749.9053 | 0.18836 | 170.6837 | 1.1331 | 1.1325 | 107.9395 | 159722.513 | 5749.8872 | 0.18706 | 171.1719 | 1.1331 | 1.1329 | 107.95 | 159727.4496 | 5749.9023 |  | -0.14505 | [5 empty] | 5750.1895 | 9958.6261 | 229.9967 | 49.9986 | -59.9975 | 0.019239 | 0.009004 | 8.5299e-05 | 1.9711e-05 | 0.00010228 | -0.0026794 | 0.004303 | 29.4339 | 2.0423 | 0.0087882 | 0.50004 | 4.6001 | 6.1788e-06 | 99997.3206 | 0 | 0 | 109.135 | 23.5765 | 0.32293
159 | FLUHIGH | 50 | 230 | 36.87 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0037302 | 6 | 123 | 0.066995 |  | 229.9969 | 49.999 | 36.8981 | 0.7997 | 0.0072851 | 0.0056841 | 2.0553 | 100007.2851 | 2015-02-02 14:13 | 50 | 0.35254 | WP- | 0.26074 | VP- | 0.22411 | 148.3336 | 1.1331 | 1.1328 | 107.9596 | 255471.9272 | 9196.8457 | 0.22304 | 148.3253 | 1.1331 | 1.1335 | 107.9612 | 255467.5708 | 9196.793 | 0.22222 | 148.2975 | 1.1331 | 1.1324 | 107.9506 | 255471.9179 | 9196.8037 | 0.22192 | 147.8967 | 1.1331 | 1.133 | 107.9477 | 255472.3734 | 9196.8057 | 0.22187 | 148.0983 | 1.1331 | 1.1332 | 107.9445 | 255471.8317 | 9196.7725 |  | -0.13605 | [5 empty] | 9196.2905 | 6904.2968 | 229.9969 | 49.999 | 36.8981 | 0.01528 | 0.019576 | 8.7738e-05 | 5.2693e-05 | 0.00010387 | 0.0072851 | 0.0027656 | 25.5244 | 2.0553 | 0.0056841 | 0.7997 | 4.6 | 1.0963e-05 | 100007.2851 | 0 | 0 | 109.128 | 23.5584 | 0.32293
160 | FLUHIGH | 50 | 230 | 36.87 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0037302 | 6 | 123 | -0.03585 |  | 229.9967 | 49.9991 | 36.8983 | 0.7997 | 0.0067382 | 0.0056587 | 2.0574 | 100006.7382 | 2015-02-02 14:18 | 50.0001 | 0.35254 | WP- | 0.26074 | VP- | 0.223 | 148.1555 | 1.1331 | 1.1333 | 107.9783 | 255471.1339 | 9196.7725 | 0.22293 | 148.1163 | 1.1331 | 1.1333 | 107.9575 | 255469.3376 | 9196.7793 | 0.22312 | 148.3205 | 1.1331 | 1.1333 | 107.9505 | 255467.7614 | 9196.7793 | 0.2241 | 148.1631 | 1.1331 | 1.1333 | 107.9435 | 255470.8795 | 9196.7959 | 0.22321 | 148.2225 | 1.1331 | 1.1333 | 107.9244 | 255468.9384 | 9196.7959 |  | -0.12455 | [5 empty] | 9196.2863 | 6904.3471 | 229.9967 | 49.9991 | 36.8983 | 0.010058 | 0.01149 | 5.0923e-05 | 4.7197e-05 | 3.9716e-05 | 0.0067382 | 0.0027505 | 25.0034 | 2.0574 | 0.0056587 | 0.7997 | 4.6 | 7.9487e-06 | 100006.7382 | 0 | 0 | 109.121 | 23.5404 | 0.32293
161 | FLUHIGH | 100 | 230 | 0 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0032656 | 6 | 123 | 0.10513 |  | 229.9963 | 100.0019 | -0.00046615 | 1 | 0.0048111 | 0.0047109 | 2.0836 | 100004.8111 | 2015-02-02 14:59 | 49.9999 | 0.88281 | WP- | 0 | VP+ | 0.3232 | -112.5838 | 1.1331 | 1.1328 | 107.9995 | 638919.5054 | 23000.4922 | 0.32142 | -112.4801 | 1.1331 | 1.1327 | 107.9406 | 638926.3685 | 23000.5469 | 0.32318 | -112.2396 | 1.1331 | 1.1324 | 107.941 | 638916.7254 | 23000.4922 | 0.32307 | -112.2092 | 1.1331 | 1.1329 | 107.9163 | 638924.3386 | 23000.5469 | 0.3227 | -112.41 | 1.1331 | 1.1327 | 107.9483 | 638922.0689 | 23000.5312 |  | -0.14016 | [5 empty] | 23000.0783 | 0.18713 | 229.9963 | 100.0019 | -0.00046615 | 0.01973 | 0.020343 | 3.6827e-05 | 7.7294e-05 | 5.0677e-05 | 0.0048111 | 0.0022609 | 19.8518 | 2.0836 | 0.0047109 | 1 | 2.2999 | 3.6745e-06 | 100004.8111 | 0 | 0 | 109.136 | 23.579 | 0.32293
162 | FLUHIGH | 100 | 230 | 0 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0032656 | 6 | 123 | 0.1738 |  | 229.9966 | 100.0018 | -0.00043844 | 1 | 0.0045877 | 0.0047193 | 2.0825 | 100004.5877 | 2015-02-02 15:03 | 50.0006 | 0.88281 | WP- | 0 | VP+ | 0.32321 | -112.5239 | 1.1331 | 1.1335 | 107.9423 | 638920.1183 | 23000.5703 | 0.32459 | -112.4519 | 1.1331 | 1.1334 | 107.9594 | 638925.2835 | 23000.5059 | 0.32451 | -112.4228 | 1.1331 | 1.1329 | 107.9863 | 638915.9585 | 23000.5391 | 0.32387 | -112.1549 | 1.1331 | 1.1338 | 107.9165 | 638920.332 | 23000.5469 | 0.32319 | -112.3857 | 1.1331 | 1.1328 | 107.951 | 638920.217 | 23000.498 |  | -0.13598 | [5 empty] | 23000.0786 | 0.176 | 229.9966 | 100.0018 | -0.00043844 | 0.012367 | 0.017938 | 4.7247e-05 | 5.317e-05 | 4.4685e-05 | 0.0045877 | 0.0022662 | 20.0263 | 2.0825 | 0.0047193 | 1 | 2.2999 | 3.0884e-06 | 100004.5877 | 0 | 0 | 109.14 | 23.5893 | 0.32293
163 | FLUHIGH | 100 | 230 | -60 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0051637 | 6 | 123 | -0.15429 |  | 229.9969 | 99.9978 | -59.9976 | 0.50004 | -0.0026261 | 0.0084975 | 2.0532 | 99997.3739 | 2015-02-02 15:08 | 50.0001 | 0.44141 | WP- | 0.75293 | VP+ | 0.19759 | -16.3087 | 1.1331 | 1.1338 | 107.9221 | 319451.4752 | 11499.707 | 0.19738 | -17.1013 | 1.1331 | 1.1331 | 107.9578 | 319449.1044 | 11499.7119 | 0.19658 | -16.8445 | 1.1331 | 1.1324 | 107.957 | 319443.0673 | 11499.7217 | 0.19968 | -16.9004 | 1.1331 | 1.1334 | 107.9538 | 319445.4254 | 11499.7354 | 0.19789 | -16.6 | 1.1331 | 1.1327 | 107.9594 | 319450.7654 | 11499.7686 |  | -0.18236 | [5 empty] | 11500.4288 | 19917.386 | 229.9969 | 99.9978 | -59.9976 | 0.021237 | 0.025569 | 6.467e-05 | 9.5432e-05 | 5.8283e-05 | -0.0026261 | 0.0041387 | 26.0984 | 2.0532 | 0.0084975 | 0.50004 | 2.3 | 4.9486e-06 | 99997.3739 | 0 | 0 | 109.135 | 23.5765 | 0.32293
164 | FLUHIGH | 100 | 230 | -60 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0051637 | 6 | 123 | -0.084918 |  | 229.9967 | 99.9975 | -59.9976 | 0.50004 | -0.0025269 | 0.0084965 | 2.0532 | 99997.4731 | 2015-02-02 15:12 | 50 | 0.44141 | WP- | 0.75293 | VP+ | 0.19759 | -16.747 | 1.1331 | 1.1327 | 107.9552 | 319451.6866 | 11499.7656 | 0.19707 | -17.616 | 1.1331 | 1.1331 | 107.9485 | 319445.875 | 11499.6943 | 0.19584 | -16.6084 | 1.1331 | 1.1329 | 107.9391 | 319448.6181 | 11499.7207 | 0.19757 | -16.8593 | 1.1331 | 1.1327 | 107.9398 | 319444.2061 | 11499.7871 | 0.19831 | -17.0122 | 1.1331 | 1.1333 | 107.9494 | 319444.9287 | 11499.7266 |  | -0.060356 | [5 empty] | 11500.3849 | 19917.333 | 229.9967 | 99.9975 | -59.9976 | 0.023778 | 0.018697 | 7.1611e-05 | 6.5631e-05 | 6.7273e-05 | -0.0025269 | 0.0041382 | 26.0848 | 2.0532 | 0.0084965 | 0.50004 | 2.3 | 4.3155e-06 | 99997.4731 | 0 | 0 | 109.127 | 23.5558 | 0.32293
165 | FLUHIGH | 100 | 230 | 36.87 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0037302 | 6 | 123 | 0.22073 |  | 229.9969 | 99.998 | 36.8678 | 0.80002 | 0.0073462 | 0.0056674 | 2.0567 | 100007.3462 | 2015-02-02 15:16 | 50 | 0.70605 | WP- | 0.52148 | VP- | 0.31347 | -140.76 | 1.1331 | 1.1326 | 107.9265 | 511142.1577 | 18400.8535 | 0.31487 | -140.7157 | 1.1331 | 1.1331 | 107.9382 | 511151.5554 | 18400.8262 | 0.31274 | -140.6503 | 1.1331 | 1.133 | 107.9377 | 511140.1853 | 18400.8184 | 0.31363 | -140.9959 | 1.1331 | 1.1331 | 107.9325 | 511148.7721 | 18400.8672 | 0.31156 | -140.8668 | 1.1331 | 1.1337 | 107.9549 | 511145.7376 | 18400.8242 |  | 0.005891 | [5 empty] | 18399.8928 | 13798.8614 | 229.9969 | 99.998 | 36.8678 | 0.019194 | 0.019513 | 4.0908e-05 | 6.557e-05 | 5.8588e-05 | 0.0073462 | 0.0027556 | 25.1742 | 2.0567 | 0.0056674 | 0.80002 | 2.3 | 3.8225e-06 | 100007.3462 | 0 | 0 | 109.119 | 23.5352 | 0.32293
166 | FLUHIGH | 100 | 230 | 36.87 | 230 | ESI(1) | 100 | 50 | 10 | 5 | 100000 | watts | 0.0037302 | 6 | 123 | 0.36189 |  | 229.9966 | 99.9979 | 36.8678 | 0.80002 | 0.0080558 | 0.0056697 | 2.0565 | 100008.0558 | 2015-02-02 15:20 | 50 | 0.70605 | WP- | 0.52148 | VP- | 0.3135 | -140.9121 | 1.1331 | 1.1328 | 107.9619 | 511144.9582 | 18400.7949 | 0.31258 | -140.308 | 1.1331 | 1.134 | 107.9378 | 511154.2349 | 18400.7871 | 0.3136 | -141.07 | 1.1331 | 1.133 | 107.957 | 511143.3623 | 18400.8125 | 0.31248 | -140.6531 | 1.1331 | 1.1331 | 107.9431 | 511150.7465 | 18400.8555 | 0.31357 | -140.7376 | 1.1331 | 1.133 | 107.9455 | 511147.429 | 18400.8105 |  | 0.017307 | [5 empty] | 18399.851 | 13798.8289 | 229.9966 | 99.9979 | 36.8678 | 0.017083 | 0.033316 | 6.6795e-05 | 0.00013321 | 4.5185e-05 | 0.0080558 | 0.002757 | 25.2156 | 2.0565 | 0.0056697 | 0.80002 | 2.3 | 6.686e-06 | 100008.0558 | 0 | 0 | 109.115 | 23.5249 | 0.32293
[17 empty] | Temperature | Humidity
[17 empty] | 19.76 | 65.69
[17 empty] | 19.71 | 60.9
[17 empty] | 19.55 | 67.95
[16 empty] | avg uncorrected | 19.6733 | 64.8467
[16 empty] | correction | 0.53
[16 empty] | corrected avg | 20.2033
Note on BA7: Callum McKenzie: This uses the uncorrected C1 value since we haven't measured the temperature at this point. It shouldn't matter because the dials are 1000ppm vs. the C1 variation of 50ppm.
Note on BP7: Callum McKenzie: Clamps added 16/03/2006 Checked Laurie 6/4/2006
Note on BR7: Callum McKenzie: Clamps added 16/03/2006 Checked Laurie 6/4/2006
Note on E9: Laurie Christian: This is the setting for the source being used. The Tables pages looks at the Watt-bridge.

## RD31 Data
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80 | 81 | 82 | 83 | 84 | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 117 | 118 | 119 | 120 | 121 | 122 | 123 | 124 | 125 | 126 | 127 | 128 | 129 | 130 | 131 | 132 | 133 | 134 | 135 | 136 | 137 | 138 | 139 | 140 | 141 | 142 | 143 | 144 | 145 | 146 | 147 | 148 | 149 | 150 | 151 | 152 | 153 | 154 | 155 | 156 | 157 | 158 | 159 | 160 | 161 | 162 | 163 | 164 | 165 | 166 | 167 | 168 | 169 | 170 | 171 | 172 | 173 | 174 | 175 | 176 | 177 | 178 | 179 | 180 | 181 | 182 | 183 | 184 | 185 | 186 | 187 | 188 | 189 | 190 | 191 | 192 | 193 | 194 | 195 | 196 | 197 | 198 | 199 | 200 | 201 | 202
[204 empty] | Averages | [6 empty] | Averages | [6 empty] | Averages | [5 empty] | Averages |  |  |  | Error | [6 empty] | Error | [6 empty] | Error | [5 empty] | Error
| Reading 1 | [6 empty] | Reading 1 | [6 empty] | Reading 1 | [6 empty] | Reading 1 | [6 empty] | Reading 2 | [6 empty] | Reading 2 | [6 empty] | Reading 2 | [6 empty] | Reading 2 | [6 empty] | Reading 3 | [6 empty] | Reading 3 | [6 empty] | Reading 3 | [6 empty] | Reading 3 | [6 empty] | Reading 4 | [6 empty] | Reading 4 | [6 empty] | Reading 4 | [6 empty] | Reading 4 | [6 empty] | Reading 5 | [6 empty] | Reading 5 | [6 empty] | Reading 5 | [6 empty] | Reading 5
Row | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | Phase | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Power Factor | RD-31 Analog Sense | RD-31 Delta phase (Volts A & phase read) | RD-31 Volts Delta mode | RD-31 Watts Delta mode | RD-31 VA Delta mode | RD-31 Vars Delta mode | RD-31 Vars WYE Cross-connected | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Power Factor | RD-31 Analog Sense | RD-31 Delta phase (Volts A & phase read) | RD-31 Volts Delta mode | RD-31 Watts Delta mode | RD-31 VA Delta mode | RD-31 Vars Delta mode | RD-31 Vars WYE Cross-connected | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Power Factor | RD-31 Analog Sense | RD-31 Delta phase (Volts A & phase read) | RD-31 Volts Delta mode | RD-31 Watts Delta mode | RD-31 VA Delta mode | RD-31 Vars Delta mode | RD-31 Vars WYE Cross-connected | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Power Factor | RD-31 Analog Sense | RD-31 Delta phase (Volts A & phase read) | RD-31 Volts Delta mode | RD-31 Watts Delta mode | RD-31 VA Delta mode | RD-31 Vars Delta mode | RD-31 Vars WYE Cross-connected | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Voltage (V) | RD-31 Current (A) | RD-31 Watts | RD-31 VA | RD-31 Vars | RD-31 Freq (Hz) | RD-31 Phase (Non-IEC) (deg) | RD-31 Watts | RD-31 VA | RD-31 Vars
| Input Ch 1 | [6 empty] | Input Ch 2 | [6 empty] | Input Ch 3 | [6 empty] | Input Net (Total) | [6 empty] | Input Ch 1 | [6 empty] | Input Ch 2 | [6 empty] | Input Ch 3 | [6 empty] | Input Net (Total) | [6 empty] | Input Ch 1 | [6 empty] | Input Ch 2 | [6 empty] | Input Ch 3 | [6 empty] | Input Net (Total) | [6 empty] | Input Ch 1 | [6 empty] | Input Ch 2 | [6 empty] | Input Ch 3 | [6 empty] | Input Net (Total) | [6 empty] | Input Ch 1 | [6 empty] | Input Ch 2 | [6 empty] | Input Ch 3 | [6 empty] | Input Net (Total) | [7 empty] | Input Ch 1 | [14 empty] | Input Ch 2 | [14 empty] | Input Ch 3 | [14 empty] | Input Net (Total) | [14 empty] | Input Ch 1 | [6 empty] | Input Ch 2 | [6 empty] | Input Ch 3 | [6 empty] | Total |  |  | Input Ch 1 | [6 empty] | Input Ch 2 | [6 empty] | Input Ch 3 | [6 empty] | Total
12 | 230.0017 | 2.0001 | 460.0302 | 460.0309 | 0.029275 | 49.9999 | -0.0036462 | 0.017548 | 3.6845e-06 | 3.9337e-09 | 6.35e-08 | -2.4997e-09 | 74.7484 | 19.9983 | 0.019704 | 3.1416e-06 | 6.5668e-09 | 6.245e-08 | 4.7291e-09 | 74.7117 | -42.0665 | 3.4028e+38 | 3.4028e+38 | 460.0302 | 460.0309 | 0.029275 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | 460.0317 | 460.0309 | 0.028164 | 49.9999 | -0.0035078 | 0.01623 | 3.6201e-06 | 4.4836e-10 | 6.0945e-08 | -3.9168e-09 | 74.6873 | 72.3643 | 0.01791 | 3.0948e-06 | 3.1522e-09 | 5.5878e-08 | 6.6105e-10 | 74.7665 | -9.6838 | 3.4028e+38 | 3.4028e+38 | 460.0317 | 460.0309 | 0.028164 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 2.0001 | 460.0307 | 460.0304 | 0.028721 | 49.9999 | -0.0038125 | 0.014024 | 3.572e-06 | 9.9018e-10 | 5.0712e-08 | -9.3693e-10 | 74.695 | 21.1808 | 0.015317 | 3.1637e-06 | 2.9375e-09 | 4.7842e-08 | 9.0501e-10 | 74.7812 | -26.1007 | 3.4028e+38 | 3.4028e+38 | 460.0307 | 460.0304 | 0.028721 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 2.0001 | 460.0274 | 460.0271 | 0.025305 | 50 | -0.0031517 | 0.015623 | 3.6042e-06 | 3.5379e-09 | 5.596e-08 | -2.2343e-10 | 74.7452 | 3.8177 | 0.016692 | 3.1594e-06 | 3.2611e-09 | 5.3116e-08 | 3.3726e-09 | 74.6734 | -34.5764 | 3.4028e+38 | 3.4028e+38 | 460.0274 | 460.0271 | 0.025305 | 3.4028e+38 | 3.4028e+38 | 230.0024 | 2.0001 | 460.0338 | 460.0336 | 0.026823 | 50 | -0.0036338 | 0.016483 | 3.6496e-06 | 1.3614e-09 | 6.0184e-08 | -3.9926e-09 | 74.723 | 75.7768 | 0.017775 | 3.1087e-06 | 3.1291e-09 | 5.6109e-08 | 7.1577e-10 | 74.7484 | -14.0608 | 3.4028e+38 | 3.4028e+38 | 460.0338 | 460.0336 | 0.026823 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0022 | 2.0001 | 460.0298 | 460.0296 | 0.028989 | 50 | -0.0036105 | 1 | 3.4028e+38 | 0 | 230.0068 | 460.0393 | 460.039 | 0.044338 | 0.00037401 | 0.015479 | 3.6779e-06 | 1.3158e-09 | 5.5595e-08 | -2.1775e-09 | 74.5896 | 62.6843 | 0.019938 | 3.4028e+38 | -35.0072 | 0.020378 | 0 | 0 | 0 | -3.5688e-05 | 0.015001 | 3.1253e-06 | 2.0876e-09 | 4.8586e-08 | 5.6826e-10 | 74.6749 | -19.4965 | 0.04738 | 3.4028e+38 | -9.7146 | 229.0502 | -6.8124e-10 | 0.00074197 | 3.9205e-10 | -8.7143e-06 | 3.4028e+38 | 3.4028e+38 | 460.0317 | 460.0314 | 0.03205 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 460.0253 | 460.0253 | 0.039041 | -0.00083538 | 230.002 | 2.0001 | 460.0308 | 460.0306 | 0.027657 | 49.9999 | -0.0035504 | 0.015981 | 3.6261e-06 | 2.0543e-09 | 5.826e-08 | -2.3139e-09 | 74.7198 | 38.6276 | 0.01748 | 3.1336e-06 | 3.8093e-09 | 5.5079e-08 | 2.0767e-09 | 74.7362 | -25.2977 | 460.0308 | 460.0306 | 0.027657 | 23.8417 | 11.3649 | 35.315 | 34.9623 | 1021120.6164 |  | -0.001846 | -999930.5151 | -999998.1871 | -1000000 | -999999.9999 | -999999.8309 |  | 38.6293 | -999923.9995 | -999998.4333 | -1000000 | -999999.9999 | -999999.8482 |  | -25.296 | -666654.895 | -666655.0126 | -326293.1551
13 | 230.0015 | 2.0001 | 460.0314 | 460.0304 | 0.026308 | 50 | -0.0032765 | 0.017501 | 3.6404e-06 | 2.4702e-09 | 6.3752e-08 | -3.7925e-09 | 74.7765 | 31.3094 | 0.019581 | 3.151e-06 | 5.4104e-09 | 6.1811e-08 | 3.651e-09 | 74.7849 | -31.2209 | 3.4028e+38 | 3.4028e+38 | 460.0314 | 460.0304 | 0.026308 | 3.4028e+38 | 3.4028e+38 | 230.0025 | 2.0001 | 460.0297 | 460.0294 | 0.025199 | 49.9999 | -0.0027647 | 0.015429 | 3.6476e-06 | 2.1232e-09 | 5.4757e-08 | 6.5754e-11 | 74.7486 | -5.5792 | 0.016877 | 3.1041e-06 | 3.5946e-09 | 5.3848e-08 | 3.5316e-09 | 74.7107 | -46.717 | 3.4028e+38 | 3.4028e+38 | 460.0297 | 460.0294 | 0.025199 | 3.4028e+38 | 3.4028e+38 | 230.0024 | 2.0001 | 460.0281 | 460.0271 | 0.029195 | 50 | -0.0036361 | 0.014369 | 3.6177e-06 | 2.0267e-09 | 5.3031e-08 | -8.0045e-10 | 74.7765 | 45.6533 | 0.01525 | 3.0976e-06 | 6.2952e-10 | 4.724e-08 | 2.9804e-11 | 74.807 | -0.95941 | 3.4028e+38 | 3.4028e+38 | 460.0281 | 460.0271 | 0.029195 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 2.0001 | 460.0289 | 460.0286 | 0.027375 | 49.9999 | -0.0034425 | 0.017501 | 3.6067e-06 | 4.1282e-09 | 6.4335e-08 | -3.7516e-09 | 74.8013 | 50.5961 | 0.019577 | 3.095e-06 | 5.9723e-09 | 6.3032e-08 | 3.1726e-09 | 74.7074 | -34.4551 | 3.4028e+38 | 3.4028e+38 | 460.0289 | 460.0286 | 0.027375 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 2.0001 | 460.0291 | 460.0321 | 0.029149 | 50 | -0.0036305 | 0.017318 | 3.6124e-06 | 1.4612e-09 | 6.2487e-08 | -4.6215e-09 | 74.7931 | 73.5328 | 0.01934 | 3.1043e-06 | 5.121e-09 | 5.9513e-08 | 2.1148e-09 | 74.8221 | -20.9462 | 3.4028e+38 | 3.4028e+38 | 460.0291 | 460.0321 | 0.029149 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0018 | 2.0001 | 460.034 | 460.0337 | 0.027341 | 50 | -0.0037565 | 1 | 3.4028e+38 | 0 | 230.0071 | 460.0399 | 460.0397 | 0.048526 | 0.0004802 | 0.017074 | 3.605e-06 | 3.7729e-09 | 6.1552e-08 | -3.8296e-09 | 74.7109 | 57.3984 | 0.039371 | 3.4028e+38 | -64.0586 | 0.020593 | 0 | 0 | 0 | -5.2671e-05 | 0.018468 | 3.0869e-06 | 4.3652e-09 | 5.9332e-08 | 1.3157e-09 | 74.7778 | -16.6087 | 0.070946 | 3.4028e+38 | 3.5556 | 228.6779 | 7.569e-11 | 0.0007297 | 2.8819e-10 | -2.3167e-05 | 3.4028e+38 | 3.4028e+38 | 460.0305 | 460.0302 | 0.029682 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 460.0396 | 460.0396 | 0.051151 | 0.00072507 | 230.0021 | 2.0001 | 460.0294 | 460.0295 | 0.027445 | 49.9999 | -0.0033501 | [17 empty] | 23.2345 | 10.3521 | 33.618 | 33.8488 | 344898.9393 |  | -0.00080837
14 | 230.0022 | 2.0001 | 460.0299 | 460.0297 | 0.025393 | 49.9999 | -0.0035097 | 0.016905 | 3.6402e-06 | 3.8201e-09 | 6.225e-08 | -1.5332e-09 | 74.7995 | 7.1403 | 0.018688 | 3.1695e-06 | 5.7213e-09 | 6.0397e-08 | 3.7484e-09 | 74.816 | -38.6222 | 3.4028e+38 | 3.4028e+38 | 460.0299 | 460.0297 | 0.025393 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 2.0001 | 460.0293 | 460.029 | 0.025324 | 50 | -0.0032487 | 0.016772 | 3.6488e-06 | 1.9961e-09 | 6.0227e-08 | -3.6716e-09 | 74.7631 | 59.9313 | 0.018022 | 3.075e-06 | 3.5066e-09 | 5.851e-08 | 9.1376e-10 | 74.7671 | -11.8293 | 3.4028e+38 | 3.4028e+38 | 460.0293 | 460.029 | 0.025324 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 2.0001 | 460.0305 | 460.0302 | 0.026874 | 49.9999 | -0.0033495 | 0.016283 | 3.6164e-06 | 2.6784e-09 | 5.7136e-08 | -1.9421e-10 | 74.7395 | 2.4017 | 0.01758 | 3.0893e-06 | 3.2811e-09 | 5.6083e-08 | 3.1442e-09 | 74.8364 | -41.0759 | 3.4028e+38 | 3.4028e+38 | 460.0305 | 460.0302 | 0.026874 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | 460.0262 | 460.026 | 0.0272 | 50 | -0.0033877 | 0.016014 | 3.5982e-06 | 3.1286e-09 | 5.9581e-08 | -3.9862e-10 | 74.7416 | 6.0636 | 0.017239 | 3.1478e-06 | 3.5873e-09 | 5.6649e-08 | 3.4283e-09 | 74.7617 | -40.2049 | 3.4028e+38 | 3.4028e+38 | 460.0262 | 460.026 | 0.0272 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | 460.0291 | 460.0288 | 0.022543 | 49.9999 | -0.0034685 | 0.017486 | 3.6547e-06 | 3.769e-09 | 6.3844e-08 | -2.5395e-09 | 74.6996 | 40.7739 | 0.019562 | 3.2322e-06 | 5.3959e-09 | 6.2935e-08 | 3.2477e-09 | 74.7536 | -33.0416 | 3.4028e+38 | 3.4028e+38 | 460.0291 | 460.0288 | 0.022543 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0018 | 2.0001 | 460.0284 | 460.0282 | 0.029957 | 50 | -0.0037516 | 1 | 3.4028e+38 | 0 | 230.0036 | 460.0287 | 460.0284 | 0.046501 | 0.0001036 | 0.017636 | 3.5982e-06 | 4.5961e-09 | 6.4058e-08 | -1.9645e-09 | 74.7312 | 26.8955 | 0.066501 | 3.4028e+38 | 15.3693 | 0.020519 | 0 | 0 | 0 | -5.4217e-05 | 0.019492 | 3.1135e-06 | 5.9971e-09 | 6.0725e-08 | 4.1717e-09 | 74.7258 | -44.0174 | 0.060283 | 3.4028e+38 | 16.1447 | 229.7713 | 3.3308e-10 | 0.00071207 | -1.1134e-10 | -2.8308e-05 | 3.4028e+38 | 3.4028e+38 | 460.0276 | 460.0309 | 0.027156 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 460.0313 | 460.0313 | 0.04896 | -0.00025423 | 230.0022 | 2.0001 | 460.029 | 460.0288 | 0.025467 | 49.9999 | -0.0033928 | [17 empty] | 23.6301 | 12.3412 | 36.316 | 35.7507 | 691032.8597 |  | -0.0015171
15 | 230.0011 | 2.0001 | 229.9878 | 460.0289 | 398.4102 | 50 | -60.0033 | 0.016958 | 3.6098e-06 | 4.8559e-09 | 6.1801e-08 | 4.1919e-09 | 74.7941 | -41.7531 | 0.018409 | 3.1272e-06 | 1.9827e-09 | 5.9263e-08 | 4.676e-09 | 74.7568 | -60.38 | 3.4028e+38 | 3.4028e+38 | 229.9878 | 460.0289 | 398.4102 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 2.0001 | 229.9899 | 460.0264 | 398.4098 | 50 | -60.0036 | 0.017568 | 3.6064e-06 | 4.8647e-09 | 6.3343e-08 | 2.4164e-09 | 74.7935 | -27.382 | 0.019353 | 3.1258e-06 | 4.2364e-09 | 6.05e-08 | 5.0171e-09 | 74.775 | -60.0454 | 3.4028e+38 | 3.4028e+38 | 229.9899 | 460.0264 | 398.4098 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 2.0001 | 229.9857 | 460.0291 | 398.4125 | 49.9999 | -60.0034 | 0.014666 | 3.6013e-06 | 2.7275e-09 | 5.3324e-08 | 7.9755e-10 | 74.7674 | -14.4191 | 0.015744 | 3.0971e-06 | 1.5691e-09 | 5.0049e-08 | 1.4187e-09 | 74.8163 | -47.1343 | 3.4028e+38 | 3.4028e+38 | 229.9857 | 460.0291 | 398.4125 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 2.0001 | 229.9893 | 460.0306 | 398.4119 | 50 | -60.0038 | 0.015434 | 3.633e-06 | 2.0281e-09 | 5.5841e-08 | 2.6201e-09 | 74.8101 | -54.9106 | 0.016542 | 3.195e-06 | 2.7431e-09 | 5.0538e-08 | 2.4725e-09 | 74.7655 | -58.0361 | 3.4028e+38 | 3.4028e+38 | 229.9893 | 460.0306 | 398.4119 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 2.0001 | 229.9911 | 460.0296 | 398.4109 | 50 | -60.0034 | 0.01508 | 3.656e-06 | 2.7093e-09 | 5.4831e-08 | 3.2252e-09 | 74.7468 | -51.418 | 0.016076 | 3.0695e-06 | 1.419e-09 | 4.9655e-08 | 2.0926e-09 | 74.7725 | -49.6475 | 3.4028e+38 | 3.4028e+38 | 229.9911 | 460.0296 | 398.4109 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0023 | 2.0001 | 229.9902 | 460.0284 | 398.4112 | 49.9999 | -60.0036 | 0.49995 | 3.4028e+38 | 0 | 230.002 | 229.9825 | 460.0292 | 398.4108 | -0.00038091 | 0.014383 | 3.6204e-06 | 3.1739e-09 | 5.1899e-08 | 2.0883e-09 | 74.7007 | -39.8632 | 0.051786 | 3.4028e+38 | 38.8875 | 0.020468 | 0 | 0 | 0 | 4.3851e-05 | 0.015518 | 3.1782e-06 | 1.0113e-09 | 4.6803e-08 | 9.1307e-10 | 74.7631 | -29.1458 | 0.034651 | 3.4028e+38 | 43.8932 | 229.4548 | 1.3848e-09 | 0.00070858 | -9.9794e-11 | -2.9028e-05 | 3.4028e+38 | 3.4028e+38 | 229.9908 | 460.0313 | 398.4145 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 229.9868 | 229.9868 | 398.4159 | -0.00027308 | 230.0017 | 2.0001 | 229.9888 | 460.0289 | 398.4111 | 50 | -60.0035 | [17 empty] | 21.8849 | 34.386 | -47.3469 | 56.5506 | 89.934 |  | -0.0032561
16 | 230.0019 | 2.0001 | 229.9868 | 460.0276 | 398.4115 | 49.9999 | -60.0039 | 0.015007 | 3.612e-06 | 2.4736e-09 | 5.3789e-08 | 3.0355e-09 | 74.7409 | -63.175 | 0.016454 | 3.1691e-06 | 1.8834e-09 | 5.0253e-08 | 1.9161e-09 | 74.8222 | -74.423 | 3.4028e+38 | 3.4028e+38 | 229.9868 | 460.0276 | 398.4115 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 2.0001 | 229.9906 | 460.029 | 398.4133 | 49.9999 | -60.0038 | 0.015818 | 3.5647e-06 | 1.5279e-09 | 5.6364e-08 | 3.6519e-09 | 74.7399 | -48.0127 | 0.017001 | 3.1067e-06 | 1.8381e-09 | 5.283e-08 | 2.9264e-09 | 74.7651 | -58.8202 | 3.4028e+38 | 3.4028e+38 | 229.9906 | 460.029 | 398.4133 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 2.0001 | 229.9924 | 460.0319 | 398.414 | 49.9999 | -60.0034 | 0.014268 | 3.6245e-06 | 3.0528e-09 | 5.0611e-08 | 1.2237e-09 | 74.7874 | -28.1208 | 0.015416 | 3.1795e-06 | 9.4074e-10 | 4.7079e-08 | 8.8123e-10 | 74.7982 | -42.1181 | 3.4028e+38 | 3.4028e+38 | 229.9924 | 460.0319 | 398.414 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 2.0001 | 229.9925 | 460.0312 | 398.4128 | 50 | -60.0037 | 0.014437 | 3.6194e-06 | 2.3963e-09 | 5.2226e-08 | 1.8924e-09 | 74.7634 | -32.4783 | 0.015658 | 3.1938e-06 | 1.8002e-10 | 5.0102e-08 | 1.0126e-09 | 74.8527 | -78.9729 | 3.4028e+38 | 3.4028e+38 | 229.9925 | 460.0312 | 398.4128 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 2.0001 | 229.9922 | 460.0286 | 398.4098 | 49.9999 | -60.0039 | 0.017634 | 3.61e-06 | 5.9493e-09 | 6.4033e-08 | 3.7982e-09 | 74.7512 | -26.737 | 0.01956 | 3.2145e-06 | 2.6129e-09 | 6.4024e-08 | 5.0496e-09 | 74.802 | -51.7416 | 3.4028e+38 | 3.4028e+38 | 229.9922 | 460.0286 | 398.4098 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0017 | 2.0001 | 229.9864 | 460.0249 | 398.409 | 49.9999 | -60.0031 | 0.49995 | 3.4028e+38 | 0 | 230.0033 | 229.9773 | 460.0311 | 398.4111 | -0.0014298 | 0.017492 | 3.5884e-06 | 4.4387e-09 | 6.2785e-08 | 3.1569e-09 | 74.7218 | -29.6173 | 0.087359 | 3.4028e+38 | 14.899 | 0.020232 | 0 | 0 | 0 | 2.2855e-05 | 0.019093 | 3.2151e-06 | 2.7905e-09 | 5.989e-08 | 5.4657e-09 | 74.714 | -56.745 | 0.054734 | 3.4028e+38 | 24.209 | 229.3323 | 1.8052e-10 | 0.00071959 | 8.2721e-10 | -4.8192e-05 | 3.4028e+38 | 3.4028e+38 | 229.9937 | 460.0292 | 398.4104 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 229.9759 | 229.9759 | 398.4113 | -0.0017102 | 230.0019 | 2.0001 | 229.9909 | 460.0297 | 398.4123 | 49.9999 | -60.0037 | [17 empty] | 22.4853 | 31.5374 | -29.901 | 54.6609 | 85.5307 |  | -0.0030894
17 | 230.0022 | 2.0001 | 229.9883 | 460.0287 | 398.4124 | 49.9999 | -60.0033 | 0.015066 | 3.5898e-06 | 3.5578e-09 | 5.5076e-08 | 5.8204e-10 | 74.6907 | -8.5671 | 0.016252 | 3.0637e-06 | 1.2015e-09 | 5.0315e-08 | 1.1821e-09 | 74.7898 | -50.5126 | 3.4028e+38 | 3.4028e+38 | 229.9883 | 460.0287 | 398.4124 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 2.0001 | 229.9895 | 460.0285 | 398.4113 | 50 | -60.0039 | 0.017406 | 3.5551e-06 | 4.5714e-09 | 6.1844e-08 | 2.9703e-09 | 74.8013 | -28.4773 | 0.019494 | 3.1076e-06 | 2.7312e-09 | 6.0271e-08 | 4.5584e-09 | 74.7924 | -48.8613 | 3.4028e+38 | 3.4028e+38 | 229.9895 | 460.0285 | 398.4113 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 2.0001 | 229.9914 | 460.0295 | 398.4123 | 49.9999 | -60.0031 | 0.016031 | 3.6295e-06 | 5.3104e-09 | 5.5908e-08 | 1.4302e-09 | 74.7096 | -12.4061 | 0.017327 | 3.2101e-06 | 1.7108e-09 | 5.2549e-08 | 1.5757e-09 | 74.8541 | -42.6462 | 3.4028e+38 | 3.4028e+38 | 229.9914 | 460.0295 | 398.4123 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | 229.9852 | 460.026 | 398.4114 | 49.9999 | -60.0031 | 0.017715 | 3.6287e-06 | 5.657e-09 | 6.4269e-08 | 2.9473e-09 | 74.7484 | -28.5508 | 0.019166 | 3.1488e-06 | 4.7648e-09 | 6.0703e-08 | 6.0919e-09 | 74.736 | -52.889 | 3.4028e+38 | 3.4028e+38 | 229.9852 | 460.026 | 398.4114 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 2.0001 | 229.9893 | 460.0252 | 398.4082 | 49.9998 | -60.0034 | 0.017494 | 3.6579e-06 | 6.8103e-09 | 6.4907e-08 | 3.6178e-09 | 74.7708 | -19.6397 | 0.019469 | 3.0885e-06 | 2.7409e-09 | 6.2372e-08 | 4.217e-09 | 74.7993 | -58.5342 | 3.4028e+38 | 3.4028e+38 | 229.9893 | 460.0252 | 398.4082 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 2.0001 | 229.9927 | 460.0279 | 398.4076 | 49.9999 | -60.0032 | 0.49995 | 3.4028e+38 | 0 | 230.007 | 229.9868 | 460.0391 | 398.417 | -0.00052882 | 0.016257 | 3.6071e-06 | 4.9559e-09 | 5.8602e-08 | 4.7476e-10 | 74.7813 | -7.6872 | 0.062577 | 3.4028e+38 | 2.2794 | 0.020319 | 0 | 0 | 0 | 3.4871e-05 | 0.016036 | 3.1667e-06 | 2.3927e-09 | 4.9879e-08 | 1.0634e-09 | 74.7507 | -31.2598 | 0.031928 | 3.4028e+38 | 17.2251 | 229.8538 | -8.8782e-10 | 0.00070373 | 3.5173e-10 | -2.8456e-05 | 3.4028e+38 | 3.4028e+38 | 229.9914 | 460.0311 | 398.4141 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 229.9885 | 229.9878 | 398.4182 | -0.00023891 | 230.002 | 2.0001 | 229.9887 | 460.0276 | 398.4111 | 49.9999 | -60.0034 | [17 empty] | 22.9257 | 36.2637 | -39.3016 | 57.9757 | 93.0761 |  | -0.0029666
18 | 230.0019 | 2.0001 | 368.036 | 460.028 | -276.0038 | 49.9999 | 36.8671 | 0.0166 | 3.5729e-06 | 9.7886e-10 | 5.8512e-08 | -4.2416e-09 | 74.8354 | 68.4709 | 0.018135 | 3.0997e-06 | 5.8544e-09 | 5.8348e-08 | 2.5972e-09 | 74.7406 | -25.1838 | 3.4028e+38 | 3.4028e+38 | 368.036 | 460.028 | -276.0038 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | 368.0362 | 460.0255 | -275.999 | 49.9999 | 36.8677 | 0.016391 | 3.6294e-06 | 1.638e-09 | 5.898e-08 | -3.5013e-09 | 74.7957 | 73.4595 | 0.018065 | 3.0841e-06 | 4.8916e-09 | 5.7595e-08 | 3.2118e-09 | 74.7641 | -33.2885 | 3.4028e+38 | 3.4028e+38 | 368.0362 | 460.0255 | -275.9991 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 2.0001 | 368.0367 | 460.0291 | -276.0025 | 50 | 36.8673 | 0.016195 | 3.623e-06 | 5.5316e-10 | 5.9956e-08 | -3.4781e-09 | 74.7229 | 75.0865 | 0.017999 | 3.0829e-06 | 4.4148e-09 | 5.5238e-08 | 3.1819e-09 | 74.7684 | -32.7904 | 3.4028e+38 | 3.4028e+38 | 368.0367 | 460.0291 | -276.0026 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 2.0001 | 368.0398 | 460.0285 | -275.9995 | 49.9999 | 36.8672 | 0.014332 | 3.6021e-06 | 9.0672e-10 | 5.1776e-08 | -2.1278e-09 | 74.766 | 107.2534 | 0.015299 | 3.207e-06 | 2.4499e-09 | 4.8616e-08 | 1.7721e-10 | 74.7795 | -3.9376 | 3.4028e+38 | 3.4028e+38 | 368.0398 | 460.0285 | -275.9995 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | 368.035 | 460.0251 | -276.0001 | 49.9999 | 36.8672 | 0.015358 | 3.6151e-06 | 2.4924e-09 | 5.5692e-08 | -1.8265e-09 | 74.7345 | 43.3538 | 0.016641 | 3.2049e-06 | 3.6301e-09 | 5.1885e-08 | 2.6326e-09 | 74.7483 | -36.3463 | 3.4028e+38 | 3.4028e+38 | 368.035 | 460.0251 | -276.0001 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0021 | 2.0001 | 368.0378 | 460.0269 | -275.9992 | 49.9999 | 36.867 | 0.80004 | 3.4028e+38 | 0 | 230.0008 | 368.0467 | 460.0233 | -275.9787 | 0.00041738 | 0.016622 | 3.6724e-06 | 1.7206e-09 | 6.0532e-08 | -4.1072e-09 | 74.7284 | 65.8096 | 0.0014192 | 3.4028e+38 | -16.283 | 0.019537 | 0 | 0 | 0 | -7.9022e-05 | 0.018788 | 3.106e-06 | 4.8522e-09 | 5.9788e-08 | 2.6014e-09 | 74.8118 | -33.3489 | 0.079733 | 3.4028e+38 | 45.3687 | 229.1973 | 4.282e-10 | 0.00073805 | 6.9368e-10 | -4.161e-06 | 3.4028e+38 | 3.4028e+38 | 368.0399 | 460.0306 | -276.0021 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 368.0519 | 368.0519 | -275.9861 | 0.00031684 | 230.002 | 2.0001 | 368.0368 | 460.0272 | -276.001 | 49.9999 | 36.8673 | [17 empty] | 21.8643 | 30.5923 | 55.2901 | 53.4695 | 66.9984 |  | 0.00036665
19 | 230.0019 | 2.0001 | 368.0405 | 460.0304 | -276.0015 | 49.9999 | 36.8668 | 0.014487 | 3.6142e-06 | -7.0281e-10 | 5.2351e-08 | -1.946e-09 | 74.7221 | 79.3088 | 0.015294 | 3.1526e-06 | 2.4117e-09 | 4.828e-08 | -7.114e-10 | 74.7359 | 75.7823 | 3.4028e+38 | 3.4028e+38 | 368.0405 | 460.0304 | -276.0016 | 3.4028e+38 | 3.4028e+38 | 230.0013 | 2.0001 | 368.0407 | 460.031 | -276.0025 | 49.9999 | 36.8671 | 0.017645 | 3.6679e-06 | 4.8174e-11 | 6.4893e-08 | -5.3729e-09 | 74.6242 | 95.7092 | 0.019403 | 3.2188e-06 | 3.9735e-09 | 6.2235e-08 | 1.8555e-09 | 74.654 | -9.8408 | 3.4028e+38 | 3.4028e+38 | 368.0407 | 460.031 | -276.0025 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 2.0001 | 368.0383 | 460.0264 | -275.999 | 49.9999 | 36.867 | 0.016157 | 3.6362e-06 | 1.3105e-09 | 6.0215e-08 | -3.3398e-09 | 74.7056 | 69.4363 | 0.01772 | 3.098e-06 | 4.344e-09 | 5.7986e-08 | 2.5148e-09 | 74.8071 | -36.768 | 3.4028e+38 | 3.4028e+38 | 368.0383 | 460.0264 | -275.9991 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 2.0001 | 368.0387 | 460.0287 | -275.9995 | 49.9999 | 36.8668 | 0.01733 | 3.5997e-06 | -1.2341e-09 | 6.2356e-08 | -6.098e-09 | 74.7153 | 111.0342 | 0.019318 | 3.0682e-06 | 3.4957e-09 | 5.9208e-08 | 9.1227e-10 | 74.8079 | -11.0997 | 3.4028e+38 | 3.4028e+38 | 368.0387 | 460.0287 | -275.9995 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 2.0001 | 368.0371 | 460.0273 | -276.0008 | 50 | 36.8666 | 0.01432 | 3.6358e-06 | 1.2794e-09 | 5.2606e-08 | -1.3481e-09 | 74.7369 | 46.9374 | 0.01559 | 3.117e-06 | 2.3411e-09 | 4.8704e-08 | 1.0478e-09 | 74.7552 | -13.1532 | 3.4028e+38 | 3.4028e+38 | 368.0371 | 460.0273 | -276.0009 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 2.0001 | 368.0357 | 460.0279 | -276.0007 | 49.9999 | 36.8671 | 0.80004 | 3.4028e+38 | 0 | 230.002 | 368.0448 | 460.0273 | -275.9889 | -0.00027919 | 0.014438 | 3.5756e-06 | 9.236e-10 | 5.159e-08 | -1.5663e-09 | 74.7714 | 57.7857 | 0.01685 | 3.4028e+38 | -17.3821 | 0.019899 | 0 | 0 | 0 | -6.0759e-05 | 0.015962 | 3.2199e-06 | 2.8372e-09 | 4.9225e-08 | 1.2592e-09 | 74.8402 | -37.6303 | 0.058245 | 3.4028e+38 | 37.9416 | 229.3297 | 6.9399e-10 | 0.00072775 | 1.2589e-09 | 8.7036e-06 | 3.4028e+38 | 3.4028e+38 | 368.0384 | 460.0247 | -275.999 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 368.0436 | 368.0436 | -275.9837 | -0.00017227 | 230.0019 | 2.0001 | 368.0391 | 460.0288 | -276.0007 | 49.9999 | 36.8669 | [17 empty] | 22.8903 | 34.7943 | 55.3135 | 57.1222 | 77.8207 |  | 0.000459
20 | 230.0017 | 2.0001 | 368.0378 | 460.0258 | -275.9995 | 50 | 36.8671 | 0.014643 | 3.5725e-06 | 3.1955e-10 | 5.2312e-08 | -3.0416e-09 | 74.724 | 106.4664 | 0.015657 | 3.1113e-06 | 2.3928e-09 | 4.8705e-08 | -8.0639e-10 | 74.7856 | 17.3888 | 3.4028e+38 | 3.4028e+38 | 368.0378 | 460.0258 | -275.9995 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 2.0001 | 368.0352 | 460.0259 | -276.0015 | 50 | 36.8667 | 0.017409 | 3.6743e-06 | -1.3727e-09 | 6.3592e-08 | -5.8779e-09 | 74.7882 | 97.507 | 0.018946 | 3.1997e-06 | 3.7076e-09 | 5.9514e-08 | -2.0942e-10 | 74.7605 | -14.0231 | 3.4028e+38 | 3.4028e+38 | 368.0352 | 460.0259 | -276.0016 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 2.0001 | 368.0365 | 460.0291 | -276.0016 | 50 | 36.8672 | 0.014781 | 3.5828e-06 | -7.3409e-10 | 5.2963e-08 | -3.5661e-09 | 74.6474 | 95.3173 | 0.015917 | 3.1256e-06 | 1.5963e-09 | 4.967e-08 | -1.4102e-09 | 74.6353 | 44.7657 | 3.4028e+38 | 3.4028e+38 | 368.0365 | 460.0291 | -276.0016 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | 368.0427 | 460.0303 | -275.9981 | 49.9999 | 36.867 | 0.016501 | 3.6141e-06 | 3.8218e-10 | 6.0932e-08 | -3.4511e-09 | 74.7211 | 65.9725 | 0.018351 | 3.1114e-06 | 4.1742e-09 | 5.7447e-08 | 2.8235e-09 | 74.7583 | -29.7635 | 3.4028e+38 | 3.4028e+38 | 368.0427 | 460.0303 | -275.9981 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 2.0001 | 368.0365 | 460.0255 | -275.9969 | 49.9999 | 36.8667 | 0.016214 | 3.6723e-06 | -2.0016e-10 | 6.009e-08 | -6.1555e-09 | 74.7843 | 103.5269 | 0.017977 | 3.1566e-06 | 2.8448e-09 | 5.5858e-08 | -4.3686e-10 | 74.8076 | -7.7089 | 3.4028e+38 | 3.4028e+38 | 368.0364 | 460.0255 | -275.9969 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0018 | 2.0001 | 368.0382 | 460.0289 | -275.9963 | 50 | 36.8663 | 0.80004 | 3.4028e+38 | 0 | 230.0041 | 368.0536 | 460.0311 | -275.9803 | 0.0013409 | 0.017674 | 3.6002e-06 | 8.2163e-10 | 6.3389e-08 | -5.1418e-09 | 74.7051 | 62.6234 | 0.034904 | 3.4028e+38 | -26.5565 | 0.019669 | 0 | 0 | 0 | -7.6559e-05 | 0.017861 | 3.0804e-06 | 4.2861e-09 | 5.5106e-08 | 2.4378e-09 | 74.7414 | -34.0943 | 0.065238 | 3.4028e+38 | 8.9126 | 229.4536 | 1.1978e-09 | 0.00072784 | 7.0624e-10 | 7.2974e-06 | 3.4028e+38 | 3.4028e+38 | 368.0406 | 460.0272 | -275.9995 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 368.0394 | 368.0394 | -275.9936 | -0.00024495 | 230.0018 | 2.0001 | 368.0377 | 460.0273 | -275.9995 | 50 | 36.867 | [17 empty] | 22.2836 | 29.2033 | 47.5948 | 50.2785 | 70.4613 |  | 0.00055169
21 | 230.0017 | 5.0002 | 1150.071 | 1150.0686 | 0.054722 | 49.9999 | -0.0028737 | 0.01773 | 3.8546e-06 | 1.1238e-08 | 6.83e-08 | -5.8588e-09 | 74.7064 | 29.5891 | 0.019678 | 3.1935e-06 | 1.1789e-08 | 6.292e-08 | 7.804e-09 | 74.7567 | -34.9646 | 3.4028e+38 | 3.4028e+38 | 1150.071 | 1150.0686 | 0.054722 | 3.4028e+38 | 3.4028e+38 | 230.002 | 5.0003 | 1150.0698 | 1150.0691 | 0.060261 | 50 | -0.0030022 | 0.014649 | 3.864e-06 | 3.223e-09 | 5.6837e-08 | -5.5629e-09 | 74.6449 | 75.8089 | 0.016031 | 3.2655e-06 | 4.6003e-09 | 4.9915e-08 | 9.525e-10 | 74.7691 | -17.6725 | 3.4028e+38 | 3.4028e+38 | 1150.0698 | 1150.0691 | 0.060261 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 5.0002 | 1150.068 | 1150.0658 | 0.06234 | 49.9999 | -0.0030281 | 0.016685 | 3.8352e-06 | 1.0098e-08 | 6.3989e-08 | -1.3117e-09 | 74.645 | 2.0352 | 0.018716 | 3.3211e-06 | 8.1666e-09 | 6.2141e-08 | 7.6791e-09 | 74.7377 | -39.9096 | 3.4028e+38 | 3.4028e+38 | 1150.068 | 1150.0658 | 0.06234 | 3.4028e+38 | 3.4028e+38 | 230.0024 | 5.0003 | 1150.0718 | 1150.0693 | 0.063619 | 49.9999 | -0.0030072 | 0.016508 | 3.7522e-06 | 1.029e-08 | 6.2205e-08 | -5.7708e-10 | 74.7189 | 3.1323 | 0.01809 | 3.1995e-06 | 6.6969e-09 | 5.939e-08 | 7.1633e-09 | 74.8042 | -47.5063 | 3.4028e+38 | 3.4028e+38 | 1150.0718 | 1150.0693 | 0.063619 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 5.0002 | 1150.0663 | 1150.0637 | 0.058451 | 50 | -0.0029624 | 0.016221 | 3.8452e-06 | 9.4615e-09 | 6.3088e-08 | -2.3508e-10 | 74.748 | 3.2092 | 0.01796 | 3.2002e-06 | 7.879e-09 | 5.7033e-08 | 7.7147e-09 | 74.8129 | -43.2338 | 3.4028e+38 | 3.4028e+38 | 1150.0663 | 1150.0637 | 0.058451 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0018 | 5.0002 | 1150.067 | 1150.0646 | 0.053148 | 50 | -0.0026478 | 1 | 3.4028e+38 | 0 | 230.001 | 1150.0635 | 1150.0594 | 0.094356 | -0.001818 | 0.016679 | 3.8276e-06 | 1.1632e-08 | 6.4666e-08 | 7.1856e-10 | 74.8062 | -1.5402 | 0.17509 | 3.4028e+38 | -48.729 | 0.020512 | 0 | 0 | 0 | -0.00011693 | 0.018759 | 3.2256e-06 | 7.5877e-09 | 6.2314e-08 | 7.583e-09 | 74.8319 | -43.266 | 0.14567 | 3.4028e+38 | 18.9705 | 229.1116 | 1.8374e-09 | 0.00075293 | 8.6762e-10 | -4.9299e-05 | 3.4028e+38 | 3.4028e+38 | 1150.069 | 1150.0665 | 0.056771 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1150.0702 | 1150.0702 | 0.10682 | 0.00018682 | 230.002 | 5.0002 | 1150.0694 | 1150.0673 | 0.059879 | 49.9999 | -0.0029747 | [17 empty] | 22.0518 | 6.7138 | 30.5893 | 28.7904 | 1354093.3534 |  | -0.0017075
22 | 230.002 | 5.0003 | 1150.0704 | 1150.068 | 0.060457 | 49.9999 | -0.0030119 | 0.01767 | 3.8944e-06 | 8.7169e-09 | 6.8738e-08 | -6.4461e-09 | 74.7626 | 33.9302 | 0.019288 | 3.2338e-06 | 9.3448e-09 | 6.2363e-08 | 5.6301e-09 | 74.8007 | -29.6024 | 3.4028e+38 | 3.4028e+38 | 1150.0704 | 1150.068 | 0.060457 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 5.0003 | 1150.0723 | 1150.0698 | 0.063866 | 49.9999 | -0.0030042 | 0.016912 | 3.839e-06 | 5.9172e-09 | 6.6119e-08 | -1.0376e-08 | 74.661 | 59.492 | 0.018481 | 3.2591e-06 | 7.9685e-09 | 6.0645e-08 | 2.8055e-09 | 74.6943 | -17.7996 | 3.4028e+38 | 3.4028e+38 | 1150.0723 | 1150.0698 | 0.063866 | 3.4028e+38 | 3.4028e+38 | 230.0025 | 5.0003 | 1150.0725 | 1150.0647 | 0.06243 | 49.9999 | -0.0031102 | 0.014421 | 3.8246e-06 | 3.3741e-09 | 5.5117e-08 | -2.4954e-09 | 74.6728 | 40.7054 | 0.015257 | 3.2484e-06 | 2.732e-09 | 4.7481e-08 | 1.1189e-09 | 74.843 | -19.1792 | 3.4028e+38 | 3.4028e+38 | 1150.0725 | 1150.0647 | 0.06243 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 5.0003 | 1150.0695 | 1150.0669 | 0.060138 | 49.9999 | -0.0029961 | 0.017294 | 3.8916e-06 | 1.0314e-08 | 6.8089e-08 | -4.5986e-09 | 74.7393 | 27.5021 | 0.019161 | 3.2768e-06 | 9.1037e-09 | 6.2567e-08 | 8.2763e-09 | 74.8197 | -46.3151 | 3.4028e+38 | 3.4028e+38 | 1150.0695 | 1150.0669 | 0.060138 | 3.4028e+38 | 3.4028e+38 | 230.002 | 5.0002 | 1150.0652 | 1150.0629 | 0.060147 | 49.9999 | -0.0028751 | 0.015913 | 3.7879e-06 | 8.4849e-09 | 6.1072e-08 | -5.817e-12 | 74.6916 | 0.039255 | 0.01727 | 3.221e-06 | 5.7828e-09 | 5.7337e-08 | 8.1872e-09 | 74.7821 | -48.4491 | 3.4028e+38 | 3.4028e+38 | 1150.0652 | 1150.0629 | 0.060147 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0024 | 5.0003 | 1150.0713 | 1150.0688 | 0.059307 | 49.9999 | -0.0029547 | 1 | 3.4028e+38 | 0 | 230.001 | 1150.0626 | 1150.0602 | 0.099173 | -0.0022843 | 0.017046 | 3.7746e-06 | 1.0321e-08 | 6.432e-08 | -1.2818e-09 | 74.7495 | 7.0594 | 0.15945 | 3.4028e+38 | -58.3292 | 0.02096 | 0 | 0 | 0 | -0.00012406 | 0.019198 | 3.3113e-06 | 1.0357e-08 | 6.3636e-08 | 8.3642e-09 | 74.6765 | -37.9564 | 0.17995 | 3.4028e+38 | -36.1213 | 229.5541 | 3.6115e-09 | 0.00075455 | 1.2307e-09 | -4.8393e-05 | 3.4028e+38 | 3.4028e+38 | 1150.0685 | 1150.066 | 0.055028 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1150.0846 | 1150.0846 | 0.10548 | 0.0019634 | 230.0021 | 5.0003 | 1150.07 | 1150.0665 | 0.061408 | 49.9999 | -0.0029995 | [17 empty] | 23.2634 | 9.3741 | 33.3449 | 30.2931 | 1459218.0237 |  | -0.0017555
23 | 230.0019 | 5.0003 | 1150.0731 | 1150.0677 | 0.064444 | 49.9999 | -0.0032106 | 0.015364 | 3.8716e-06 | 2.4318e-09 | 5.8158e-08 | -6.0973e-09 | 74.662 | 71.2682 | 0.016336 | 3.146e-06 | 5.0184e-09 | 5.1795e-08 | 7.3424e-10 | 74.723 | -6.3031 | 3.4028e+38 | 3.4028e+38 | 1150.0731 | 1150.0677 | 0.064444 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 5.0003 | 1150.0731 | 1150.0706 | 0.059834 | 49.9999 | -0.0029421 | 0.01785 | 3.9168e-06 | 9.7877e-09 | 6.925e-08 | -6.0551e-09 | 74.824 | 26.6226 | 0.019616 | 3.2602e-06 | 1.2081e-08 | 6.3985e-08 | 5.9406e-09 | 74.7639 | -27.7853 | 3.4028e+38 | 3.4028e+38 | 1150.0731 | 1150.0706 | 0.059834 | 3.4028e+38 | 3.4028e+38 | 230.002 | 5.0003 | 1150.0734 | 1150.0708 | 0.062052 | 49.9999 | -0.0033648 | 0.014444 | 3.8856e-06 | 3.8459e-09 | 5.6111e-08 | -4.5429e-09 | 74.7323 | 43.5863 | 0.015215 | 3.2531e-06 | 2.4673e-09 | 4.9512e-08 | 1.7588e-09 | 74.6438 | -25.424 | 3.4028e+38 | 3.4028e+38 | 1150.0734 | 1150.0708 | 0.062052 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 5.0002 | 1150.0682 | 1150.0679 | 0.060092 | 50 | -0.0029937 | 0.015878 | 3.8003e-06 | 8.2714e-09 | 5.9651e-08 | 1.2356e-10 | 74.7527 | -0.95011 | 0.016981 | 3.2004e-06 | 6.1763e-09 | 5.4528e-08 | 6.2953e-09 | 74.7731 | -50.1811 | 3.4028e+38 | 3.4028e+38 | 1150.0682 | 1150.0679 | 0.060092 | 3.4028e+38 | 3.4028e+38 | 230.002 | 5.0003 | 1150.071 | 1150.0685 | 0.060152 | 49.9999 | -0.0031566 | 0.014754 | 3.8716e-06 | 3.8509e-09 | 5.7862e-08 | -6.4419e-09 | 74.6352 | 67.3137 | 0.016406 | 3.1529e-06 | 4.9896e-09 | 5.171e-08 | 4.4895e-10 | 74.8522 | -4.9902 | 3.4028e+38 | 3.4028e+38 | 1150.071 | 1150.0685 | 0.060152 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0021 | 5.0002 | 1150.0676 | 1150.0704 | 0.060245 | 50 | -0.0030013 | 1 | 3.4028e+38 | 0 | 230.007 | 1150.0918 | 1150.0891 | 0.10619 | 0.0014977 | 0.01671 | 3.9322e-06 | 4.4157e-09 | 6.573e-08 | -1.0323e-08 | 74.691 | 65.4727 | 0.070326 | 3.4028e+38 | -45.4705 | 0.021332 | 0 | 0 | 0 | -0.00012508 | 0.019449 | 3.2295e-06 | 1.0288e-08 | 6.076e-08 | 6.2612e-09 | 74.7422 | -30.2759 | 0.16486 | 3.4028e+38 | 9.8207 | 229.6834 | 1.1237e-09 | 0.00072576 | -9.5493e-12 | -4.9671e-05 | 3.4028e+38 | 3.4028e+38 | 1150.0717 | 1150.0692 | 0.056298 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1150.0796 | 1150.0796 | 0.10447 | 0.00023839 | 230.002 | 5.0003 | 1150.0718 | 1150.0691 | 0.061315 | 49.9999 | -0.0031336 | [17 empty] | 22.1175 | 9.3788 | 33.1454 | 30.816 | 1505802.4006 |  | -0.0019145
24 | 230.0019 | 5.0003 | 574.9837 | 1150.0712 | 996.021 | 50 | -60.0029 | 0.014453 | 3.8896e-06 | 7.6038e-09 | 5.6793e-08 | 3.9762e-09 | 74.767 | -29.276 | 0.015174 | 3.1656e-06 | 2.2116e-09 | 4.7631e-08 | 1.9428e-09 | 74.7294 | -51.9682 | 3.4028e+38 | 3.4028e+38 | 574.9837 | 1150.0712 | 996.021 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 5.0003 | 574.9772 | 1150.0631 | 996.0174 | 50 | -60.0031 | 0.017479 | 3.7476e-06 | 1.3267e-08 | 6.5589e-08 | 1.1241e-08 | 74.7452 | -38.5639 | 0.019322 | 3.2782e-06 | 2.2937e-09 | 6.1571e-08 | 8.8843e-09 | 74.8452 | -77.349 | 3.4028e+38 | 3.4028e+38 | 574.9772 | 1150.0631 | 996.0174 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 5.0003 | 574.9869 | 1150.0682 | 996.0198 | 49.9999 | -60.0031 | 0.017374 | 3.8258e-06 | 1.2174e-08 | 6.659e-08 | 1.128e-08 | 74.7854 | -44.2348 | 0.019205 | 3.2197e-06 | 3.8604e-09 | 6.2686e-08 | 9.3318e-09 | 74.6677 | -64.2623 | 3.4028e+38 | 3.4028e+38 | 574.9869 | 1150.0682 | 996.0198 | 3.4028e+38 | 3.4028e+38 | 230.0013 | 5.0003 | 574.9821 | 1150.0677 | 996.0177 | 49.9999 | -60.0031 | 0.017194 | 3.7877e-06 | 1.1569e-08 | 6.507e-08 | 9.4954e-09 | 74.6651 | -41.6878 | 0.018776 | 3.194e-06 | 2.9336e-09 | 6.0608e-08 | 9.5282e-09 | 74.8077 | -70.9253 | 3.4028e+38 | 3.4028e+38 | 574.9821 | 1150.0677 | 996.0177 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 5.0003 | 574.9864 | 1150.0713 | 996.021 | 49.9999 | -60.0029 | 0.016232 | 3.8673e-06 | 1.0847e-08 | 6.128e-08 | 2.5696e-09 | 74.7097 | -14.3179 | 0.017172 | 3.1625e-06 | 3.7873e-09 | 5.5844e-08 | 4.6368e-09 | 74.763 | -54.6095 | 3.4028e+38 | 3.4028e+38 | 574.9864 | 1150.0713 | 996.021 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 5.0003 | 574.985 | 1150.071 | 996.0217 | 49.9999 | -60.003 | 0.49996 | 3.4028e+38 | 0 | 230.0072 | 574.9717 | 1150.0981 | 996.0475 | -0.0010382 | 0.016078 | 3.8928e-06 | 1.2914e-08 | 6.2624e-08 | 3.7521e-09 | 74.7559 | -18.1111 | 0.20363 | 3.4028e+38 | -10.0718 | 0.020462 | 0 | 0 | 0 | 7.1518e-05 | 0.017993 | 3.2393e-06 | 5.4301e-09 | 5.8349e-08 | 7.5422e-09 | 74.7655 | -53.4308 | 0.098357 | 3.4028e+38 | -73.8467 | 229.6887 | 1.4943e-09 | 0.00074562 | 2.2746e-09 | -9.4142e-05 | 3.4028e+38 | 3.4028e+38 | 574.985 | 1150.0687 | 996.0202 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 574.9548 | 574.9548 | 996.0413 | -0.0025455 | 230.0018 | 5.0003 | 574.9833 | 1150.0683 | 996.0194 | 49.9999 | -60.003 | [17 empty] | 21.8401 | 45.8915 | -63.9199 | 67.0321 | 112.2948 |  | -0.0044108
25 | 230.0018 | 5.0002 | 574.9804 | 1150.0646 | 996.0148 | 49.9999 | -60.003 | 0.01773 | 3.8192e-06 | 1.4079e-08 | 6.7682e-08 | 1.1041e-08 | 74.6117 | -37.5325 | 0.019515 | 3.2595e-06 | 5.4181e-09 | 6.3565e-08 | 8.9117e-09 | 74.897 | -66.0987 | 3.4028e+38 | 3.4028e+38 | 574.9804 | 1150.0646 | 996.0148 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 5.0003 | 574.9846 | 1150.0741 | 996.0239 | 49.9999 | -60.0033 | 0.014218 | 3.831e-06 | 6.2899e-09 | 5.4535e-08 | 4.8568e-09 | 74.6382 | -30.5646 | 0.01532 | 3.1245e-06 | 1.3767e-09 | 4.9073e-08 | 1.93e-09 | 74.743 | -88.8889 | 3.4028e+38 | 3.4028e+38 | 574.9846 | 1150.0741 | 996.0239 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 5.0003 | 574.9869 | 1150.0697 | 996.0189 | 49.9999 | -60.003 | 0.017649 | 3.8083e-06 | 1.5297e-08 | 6.6024e-08 | 6.1996e-09 | 74.707 | -21.6725 | 0.019148 | 3.1778e-06 | 6.2393e-09 | 6.0844e-08 | 9.5987e-09 | 74.7436 | -58.4022 | 3.4028e+38 | 3.4028e+38 | 574.9869 | 1150.0697 | 996.0189 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 5.0003 | 574.9854 | 1150.0725 | 996.0225 | 49.9999 | -60.003 | 0.014605 | 3.8845e-06 | 5.6714e-09 | 5.5806e-08 | 5.2955e-09 | 74.7697 | -42.618 | 0.015472 | 3.232e-06 | 1.2783e-09 | 4.9165e-08 | 3.291e-09 | 74.7655 | -68.7722 | 3.4028e+38 | 3.4028e+38 | 574.9854 | 1150.0725 | 996.0225 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 5.0003 | 574.9832 | 1150.0726 | 996.0215 | 50 | -60.003 | 0.014243 | 3.9435e-06 | 6.9527e-09 | 5.6052e-08 | 4.0431e-09 | 74.8103 | -32.1059 | 0.015224 | 3.2417e-06 | 2.099e-09 | 4.7964e-08 | 2.3064e-09 | 74.6219 | -41.6521 | 3.4028e+38 | 3.4028e+38 | 574.9832 | 1150.0726 | 996.0215 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0024 | 5.0003 | 574.9868 | 1150.0728 | 996.0217 | 50 | -60.003 | 0.49996 | 3.4028e+38 | 0 | 230.0033 | 574.965 | 1150.0732 | 996.0273 | -0.0005739 | 0.014194 | 3.8647e-06 | 5.2731e-09 | 5.5449e-08 | 5.5507e-09 | 74.7107 | -42.2977 | 0.10696 | 3.4028e+38 | -89.8158 | 0.020275 | 0 | 0 | 0 | 9.0563e-05 | 0.015445 | 3.2385e-06 | 7.3172e-10 | 4.9362e-08 | 3.0587e-09 | 74.7416 | -61.275 | 0.03396 | 3.4028e+38 | 31.6036 | 229.6642 | -4.7248e-10 | 0.00072106 | -3.9287e-10 | -8.4269e-05 | 3.4028e+38 | 3.4028e+38 | 574.9846 | 1150.0721 | 996.0232 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 574.9554 | 574.9554 | 996.0181 | -0.0018493 | 230.0021 | 5.0003 | 574.9841 | 1150.0707 | 996.0203 | 49.9999 | -60.003 | [17 empty] | 22.8087 | 40.9611 | -63.1179 | 63.347 | 105.7896 |  | -0.0042581
26 | 230.0024 | 5.0003 | 574.9862 | 1150.0751 | 996.0258 | 49.9999 | -60.003 | 0.015074 | 3.9412e-06 | 1.0215e-08 | 5.9703e-08 | 2.0425e-09 | 74.7329 | -14.5772 | 0.01618 | 3.1333e-06 | 3.3353e-09 | 5.0702e-08 | 3.7131e-09 | 74.8194 | -45.4599 | 3.4028e+38 | 3.4028e+38 | 574.9862 | 1150.0751 | 996.0258 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 5.0002 | 574.9783 | 1150.0641 | 996.0226 | 49.9999 | -60.0031 | 0.015552 | 3.8247e-06 | 8.021e-09 | 5.9504e-08 | 9.3876e-09 | 74.8084 | -47.3018 | 0.016965 | 3.26e-06 | 9.6623e-10 | 5.4726e-08 | 5.0727e-09 | 74.809 | -88.7401 | 3.4028e+38 | 3.4028e+38 | 574.9783 | 1150.0641 | 996.0226 | 3.4028e+38 | 3.4028e+38 | 230.0025 | 5.0003 | 574.9843 | 1150.0712 | 996.0197 | 49.9999 | -60.0032 | 0.014906 | 3.8228e-06 | 7.3667e-09 | 5.6923e-08 | 6.7406e-09 | 74.6117 | -45.3675 | 0.016173 | 3.1399e-06 | -3.9839e-10 | 5.1976e-08 | 3.5913e-09 | 74.6992 | -75.2208 | 3.4028e+38 | 3.4028e+38 | 574.9842 | 1150.0712 | 996.0197 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 5.0003 | 574.9848 | 1150.0708 | 996.0209 | 50 | -60.003 | 0.014755 | 3.9513e-06 | 8.5199e-09 | 5.9161e-08 | 1.6619e-09 | 74.6492 | -11.5509 | 0.01615 | 3.2654e-06 | 3.1247e-09 | 5.1383e-08 | 3.292e-09 | 74.7249 | -43.2819 | 3.4028e+38 | 3.4028e+38 | 574.9848 | 1150.0708 | 996.0209 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 5.0003 | 574.9809 | 1150.066 | 996.0164 | 50 | -60.003 | 0.017531 | 3.888e-06 | 1.4946e-08 | 6.8179e-08 | 5.4112e-09 | 74.6273 | -19.0611 | 0.019026 | 3.2925e-06 | 4.4566e-09 | 6.1156e-08 | 7.61e-09 | 74.7689 | -56.7222 | 3.4028e+38 | 3.4028e+38 | 574.9809 | 1150.066 | 996.0164 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 5.0003 | 574.989 | 1150.0695 | 996.0164 | 49.9999 | -60.0029 | 0.49996 | 3.4028e+38 | 0 | 230.005 | 574.9461 | 1150.0823 | 996.034 | -0.0034263 | 0.017477 | 3.8596e-06 | 1.4216e-08 | 6.7445e-08 | 9.645e-09 | 74.6309 | -31.7757 | 0.22898 | 3.4028e+38 | 21.5622 | 0.020901 | 0 | 0 | 0 | 6.4132e-05 | 0.01946 | 3.1227e-06 | 5.288e-09 | 6.083e-08 | 9.5769e-09 | 74.8285 | -68.4363 | 0.059602 | 3.4028e+38 | 20.5004 | 229.5087 | 1.7611e-09 | 0.00074541 | 1.0184e-09 | -9.8972e-05 | 3.4028e+38 | 3.4028e+38 | 574.9849 | 1150.068 | 996.0169 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 574.9407 | 574.9407 | 996.0236 | -0.0037249 | 230.002 | 5.0003 | 574.9829 | 1150.0694 | 996.0211 | 49.9999 | -60.0031 | [17 empty] | 22.8731 | 41.9069 | -61.5538 | 64.3554 | 108.1507 |  | -0.0042547
27 | 230.0016 | 5.0003 | 920.0944 | 1150.068 | -690.0043 | 50 | 36.8676 | 0.017359 | 3.9398e-06 | -4.6676e-10 | 6.828e-08 | -1.407e-08 | 74.6694 | 91.9001 | 0.019163 | 3.2395e-06 | 8.657e-09 | 6.1931e-08 | 9.9046e-10 | 74.7662 | -3.9443 | 3.4028e+38 | 3.4028e+38 | 920.0944 | 1150.068 | -690.0044 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 5.0002 | 920.0891 | 1150.0647 | -690.0042 | 49.9999 | 36.8674 | 0.015016 | 3.7853e-06 | 4.273e-09 | 5.6882e-08 | -3.4726e-09 | 74.7244 | 45.8754 | 0.016091 | 3.3187e-06 | 6.6723e-09 | 5.1636e-08 | 3.7816e-09 | 74.7203 | -29.5429 | 3.4028e+38 | 3.4028e+38 | 920.0891 | 1150.0647 | -690.0042 | 3.4028e+38 | 3.4028e+38 | 230.002 | 5.0002 | 920.0875 | 1150.0642 | -690.0065 | 49.9999 | 36.8676 | 0.014714 | 3.7844e-06 | 2.2062e-09 | 5.6076e-08 | -4.9716e-09 | 74.7853 | 56.2115 | 0.015869 | 3.1931e-06 | 4.8388e-09 | 5.0652e-08 | 1.6142e-09 | 74.7007 | -24.5179 | 3.4028e+38 | 3.4028e+38 | 920.0875 | 1150.0642 | -690.0066 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 5.0002 | 920.0872 | 1150.0656 | -690.011 | 50 | 36.8676 | 0.015045 | 3.8599e-06 | 2.9593e-09 | 5.7578e-08 | -4.4643e-09 | 74.8039 | 61.4623 | 0.016255 | 3.2834e-06 | 5.6063e-09 | 5.1635e-08 | 3.4778e-09 | 74.7592 | -34.4658 | 3.4028e+38 | 3.4028e+38 | 920.0872 | 1150.0656 | -690.011 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 5.0002 | 920.0852 | 1150.0625 | -690.002 | 49.9999 | 36.8673 | 0.014365 | 3.775e-06 | 1.8645e-09 | 5.5621e-08 | -4.316e-09 | 74.7175 | 76.3253 | 0.015674 | 3.2641e-06 | 4.6223e-09 | 5.0657e-08 | 2.3818e-09 | 74.7632 | -19.3394 | 3.4028e+38 | 3.4028e+38 | 920.0852 | 1150.0625 | -690.002 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0021 | 5.0002 | 920.0883 | 1150.0637 | -690.0013 | 49.9999 | 36.8673 | 0.80003 | 3.4028e+38 | 0 | 230.0031 | 920.105 | 1150.0708 | -689.9852 | -0.00043122 | 0.014277 | 3.861e-06 | -1.0249e-09 | 5.5152e-08 | -5.6802e-09 | 74.7614 | 97.0435 | -0.012596 | 3.4028e+38 | -73.7339 | 0.019724 | 0 | 0 | 0 | -0.00017367 | 0.015398 | 3.201e-06 | 3.8733e-09 | 4.9293e-08 | -1.5781e-09 | 74.7536 | 19.7047 | 0.085176 | 3.4028e+38 | 47.4425 | 229.7553 | 1.0471e-09 | 0.00072174 | -1.0099e-10 | 2.1847e-05 | 3.4028e+38 | 3.4028e+38 | 920.0891 | 1150.0641 | -690.0078 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 920.1196 | 920.1196 | -689.9799 | 0.0013642 | 230.0018 | 5.0002 | 920.0887 | 1150.065 | -690.0056 | 49.9999 | 36.8675 | [17 empty] | 21.9053 | 44.2239 | 69.1582 | 65.8497 | 82.4483 |  | 0.00040155
28 | 230.002 | 5.0003 | 920.0885 | 1150.0635 | -690.0009 | 49.9999 | 36.8672 | 0.017202 | 3.894e-06 | -3.2698e-09 | 6.7341e-08 | -1.3567e-08 | 74.6913 | 105.5007 | 0.018592 | 3.2584e-06 | 6.256e-09 | 6.1102e-08 | -1.2202e-09 | 74.7671 | 11.3802 | 3.4028e+38 | 3.4028e+38 | 920.0884 | 1150.0635 | -690.001 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 5.0003 | 920.0914 | 1150.0688 | -690.0083 | 49.9999 | 36.8674 | 0.016569 | 3.8949e-06 | -3.2858e-09 | 6.4476e-08 | -1.4528e-08 | 74.7173 | 105.3189 | 0.017854 | 3.251e-06 | 6.5777e-09 | 5.7696e-08 | -9.0937e-10 | 74.8144 | 15.0697 | 3.4028e+38 | 3.4028e+38 | 920.0914 | 1150.0688 | -690.0084 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 5.0002 | 920.0869 | 1150.0645 | -690.005 | 49.9999 | 36.8674 | 0.014386 | 3.8109e-06 | 1.3661e-09 | 5.4722e-08 | -3.9356e-09 | 74.7225 | 69.562 | 0.015525 | 3.2149e-06 | 3.5875e-09 | 4.9948e-08 | 7.5041e-10 | 74.7049 | -12.2755 | 3.4028e+38 | 3.4028e+38 | 920.0869 | 1150.0645 | -690.0051 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 5.0003 | 920.0882 | 1150.0654 | -690.0076 | 49.9999 | 36.8674 | 0.017627 | 3.9274e-06 | -1.3809e-12 | 6.9313e-08 | -1.4975e-08 | 74.7682 | 93.6473 | 0.0196 | 3.1694e-06 | 9.4416e-09 | 6.2207e-08 | 1.4821e-09 | 74.8417 | -9.1271 | 3.4028e+38 | 3.4028e+38 | 920.0881 | 1150.0654 | -690.0076 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 5.0002 | 920.0862 | 1150.0618 | -690.0038 | 49.9999 | 36.8675 | 0.015244 | 3.8731e-06 | -3.6202e-09 | 5.9017e-08 | -1.0999e-08 | 74.6571 | 114.5888 | 0.016442 | 3.1413e-06 | 3.6982e-09 | 5.4711e-08 | -1.6736e-09 | 74.7662 | 23.9097 | 3.4028e+38 | 3.4028e+38 | 920.0861 | 1150.0618 | -690.0039 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0014 | 5.0003 | 920.0912 | 1150.0645 | -690.0025 | 49.9999 | 36.8673 | 0.80003 | 3.4028e+38 | 0 | 230.0067 | 920.1354 | 1150.0878 | -689.9613 | 0.0033581 | 0.017065 | 3.9089e-06 | -2.8153e-09 | 6.8063e-08 | -1.5302e-08 | 74.727 | 103.6577 | -0.054632 | 3.4028e+38 | 27.8925 | 0.020023 | 0 | 0 | 0 | -0.0001972 | 0.01913 | 3.1421e-06 | 8.4735e-09 | 6.3688e-08 | 2.2965e-09 | 74.7385 | -13.9132 | 0.15152 | 3.4028e+38 | -36.5627 | 229.5651 | 1.2909e-09 | 0.00075563 | 4.3998e-10 | 3.9816e-06 | 3.4028e+38 | 3.4028e+38 | 920.0932 | 1150.0685 | -690.0047 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 920.1226 | 920.1226 | -689.9573 | 0.0019394 | 230.0018 | 5.0003 | 920.0882 | 1150.0648 | -690.0051 | 49.9999 | 36.8674 | [17 empty] | 21.8887 | 42.1076 | 67.3593 | 62.4397 | 75.0994 |  | 0.00012879
29 | 230.0017 | 5.0002 | 920.0851 | 1150.0627 | -690.0064 | 49.9998 | 36.8672 | 0.015852 | 3.8918e-06 | -4.4145e-09 | 6.1761e-08 | -1.2765e-08 | 74.7458 | 108.7364 | 0.017137 | 3.1758e-06 | 5.1223e-09 | 5.588e-08 | -2.8596e-09 | 74.843 | 31.183 | 3.4028e+38 | 3.4028e+38 | 920.0851 | 1150.0627 | -690.0065 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 5.0002 | 920.0887 | 1150.0665 | -690.0059 | 49.9999 | 36.8674 | 0.015875 | 3.8622e-06 | -5.2965e-09 | 6.1321e-08 | -1.382e-08 | 74.7837 | 111.9104 | 0.01711 | 3.1628e-06 | 4.423e-09 | 5.4733e-08 | -2.7409e-09 | 74.7197 | 34.838 | 3.4028e+38 | 3.4028e+38 | 920.0886 | 1150.0665 | -690.0059 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 5.0002 | 920.0891 | 1150.0681 | -690.001 | 50 | 36.8671 | 0.017031 | 3.9076e-06 | -6.0011e-10 | 6.6594e-08 | -1.3157e-08 | 74.7678 | 93.7476 | 0.018638 | 3.1775e-06 | 7.5234e-09 | 6.0009e-08 | -4.0843e-10 | 74.775 | -0.88788 | 3.4028e+38 | 3.4028e+38 | 920.0891 | 1150.0681 | -690.001 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 5.0002 | 920.0901 | 1150.0652 | -690.0053 | 50 | 36.8673 | 0.016411 | 3.8836e-06 | -3.314e-09 | 6.3778e-08 | -1.4828e-08 | 74.7266 | 105.8317 | 0.017535 | 3.1637e-06 | 5.8759e-09 | 5.7542e-08 | -1.662e-09 | 74.8195 | 15.2127 | 3.4028e+38 | 3.4028e+38 | 920.09 | 1150.0652 | -690.0054 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 5.0002 | 920.0874 | 1150.0647 | -690.0043 | 49.9999 | 36.8674 | 0.016668 | 3.9019e-06 | -2.0573e-09 | 6.4213e-08 | -1.3007e-08 | 74.6845 | 100.5388 | 0.018175 | 3.169e-06 | 7.0774e-09 | 5.8799e-08 | -1.1017e-09 | 74.7847 | 16.2221 | 3.4028e+38 | 3.4028e+38 | 920.0873 | 1150.0647 | -690.0043 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 5.0003 | 920.0934 | 1150.0675 | -690.0033 | 49.9999 | 36.8673 | 0.80003 | 3.4028e+38 | 0 | 230.0068 | 920.1353 | 1150.0903 | -689.9666 | 0.0033785 | 0.01698 | 3.925e-06 | -4.1677e-09 | 6.6644e-08 | -1.526e-08 | 74.682 | 101.4244 | -0.046316 | 3.4028e+38 | -26.7475 | 0.020561 | 0 | 0 | 0 | -0.00019517 | 0.018846 | 3.1963e-06 | 6.2835e-09 | 6.0196e-08 | -1.6139e-09 | 74.8245 | 3.4088 | 0.10698 | 3.4028e+38 | -26.3396 | 229.4803 | -4.3413e-11 | 0.00073313 | -7.5612e-10 | 8.0864e-06 | 3.4028e+38 | 3.4028e+38 | 920.0854 | 1150.0648 | -690.007 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 920.1354 | 920.1354 | -689.9578 | 0.0033921 | 230.0016 | 5.0002 | 920.0881 | 1150.0655 | -690.0046 | 49.9999 | 36.8673 | [17 empty] | 21.3711 | 39.3435 | 67.3237 | 62.6321 | 73.0082 |  | -1.5785e-05
30 | 230.0018 | 2.0001 | -0.026892 | 460.0289 | 460.0271 | 49.9999 | -90.0038 | 0.016231 | 3.5658e-06 | 1.3165e-09 | 5.7376e-08 | 5.2902e-09 | 74.8271 | -71.5692 | 0.017661 | 3.0926e-06 | 6.1779e-10 | 5.3971e-08 | 3.015e-09 | 74.6881 | -58.8884 | 3.4028e+38 | 3.4028e+38 | -0.026892 | 460.0289 | 460.0271 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | -0.029957 | 460.0282 | 460.029 | 50 | -90.0033 | 0.01437 | 3.6276e-06 | 2.5551e-09 | 5.1692e-08 | 2.3273e-09 | 74.807 | -36.7911 | 0.015402 | 3.0929e-06 | 1.0611e-09 | 4.6821e-08 | 1.358e-09 | 74.7769 | -65.6352 | 3.4028e+38 | 3.4028e+38 | -0.029957 | 460.0282 | 460.029 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | -0.02657 | 460.0281 | 460.0285 | 49.9999 | -90.003 | 0.016241 | 3.5969e-06 | 5.2674e-09 | 5.8451e-08 | 2.5935e-09 | 74.6501 | -27.9872 | 0.017642 | 3.2157e-06 | 2.6307e-09 | 5.5208e-08 | 1.9155e-09 | 74.6839 | -61.6122 | 3.4028e+38 | 3.4028e+38 | -0.02657 | 460.0281 | 460.0285 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | -0.024734 | 460.0294 | 460.0296 | 49.9999 | -90.0039 | 0.014356 | 3.5799e-06 | 1.1949e-09 | 5.2188e-08 | 2.6828e-09 | 74.7348 | -80.7356 | 0.015409 | 3.2006e-06 | 3.8692e-10 | 4.9254e-08 | 1.2981e-09 | 74.7 | -37.8554 | 3.4028e+38 | 3.4028e+38 | -0.024734 | 460.0294 | 460.0296 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 2.0001 | -0.030679 | 460.0248 | 460.0246 | 50 | -90.0038 | 0.017629 | 3.5655e-06 | 4.1122e-09 | 6.3276e-08 | 3.8877e-09 | 74.7973 | -47.0517 | 0.01965 | 3.1875e-06 | 1.5166e-09 | 6.2173e-08 | 4.0976e-09 | 74.8239 | -74.9655 | 3.4028e+38 | 3.4028e+38 | -0.030679 | 460.0248 | 460.0246 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 2.0001 | -0.026585 | 460.0273 | 460.0272 | 50 | -90.0029 | -5.0479e-05 | 3.4028e+38 | 0 | 230.0053 | -0.046927 | 460.0363 | 460.0262 | -0.0014865 | 0.017407 | 3.6195e-06 | 3.5925e-09 | 6.305e-08 | 4.0638e-09 | 74.6572 | -42.4083 | 0.071793 | 3.4028e+38 | -36.314 | 0.019978 | 0 | 0 | 0 | 5.2559e-05 | 0.018718 | 3.1565e-06 | 9.7294e-10 | 5.9019e-08 | 2.4776e-09 | 74.7043 | -57.5636 | 0.027265 | 3.4028e+38 | -4.2445 | 229.614 | -7.5238e-10 | 0.00072561 | 4.1792e-10 | -3.689e-05 | 3.4028e+38 | 3.4028e+38 | -0.024787 | 460.0277 | 460.0279 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -0.03963 | -0.03963 | 460.0296 | -0.0011776 | 230.0019 | 2.0001 | -0.027766 | 460.0279 | 460.0278 | 49.9999 | -90.0035 | [17 empty] | 22.4505 | 71.4417 | 8625186.0966 | 93.5048 | 93.2602 |  | -0.0031838
31 | 230.0019 | 2.0001 | -0.029944 | 460.0255 | 460.0253 | 50 | -90.0032 | 0.01751 | 3.5611e-06 | 4.6524e-09 | 6.2015e-08 | 3.6695e-09 | 74.7299 | -40.8639 | 0.019529 | 3.2151e-06 | 7.0953e-10 | 5.9972e-08 | 3.9059e-09 | 74.7852 | -61.9582 | 3.4028e+38 | 3.4028e+38 | -0.029944 | 460.0255 | 460.0253 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 2.0001 | -0.026301 | 460.0287 | 460.0288 | 49.9999 | -90.0033 | 0.017055 | 3.5838e-06 | 6.1259e-09 | 6.108e-08 | 2.7975e-09 | 74.6812 | -26.6534 | 0.018699 | 3.1225e-06 | 1.4334e-09 | 5.7577e-08 | 2.9541e-09 | 74.8549 | -47.5212 | 3.4028e+38 | 3.4028e+38 | -0.026301 | 460.0287 | 460.0288 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 2.0001 | -0.028607 | 460.0315 | 460.0323 | 49.9999 | -90.0039 | 0.015305 | 3.6399e-06 | 1.0224e-09 | 5.5185e-08 | 3.0983e-09 | 74.7498 | -79.5731 | 0.016372 | 3.086e-06 | 6.1283e-10 | 5.0217e-08 | 1.231e-09 | 74.7622 | -93.5249 | 3.4028e+38 | 3.4028e+38 | -0.028607 | 460.0315 | 460.0323 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | -0.028612 | 460.0299 | 460.0303 | 49.9999 | -90.0033 | 0.014344 | 3.6413e-06 | 1.1579e-09 | 5.0923e-08 | 2.8256e-09 | 74.8238 | -66.0431 | 0.015415 | 3.045e-06 | 6.6194e-10 | 4.8117e-08 | 1.4027e-09 | 74.714 | -78.7569 | 3.4028e+38 | 3.4028e+38 | -0.028612 | 460.0299 | 460.0303 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 2.0001 | -0.02699 | 460.0278 | 460.0282 | 49.9999 | -90.0037 | 0.016615 | 3.6004e-06 | 5.1711e-09 | 5.9788e-08 | 3.5594e-09 | 74.7151 | -36.4071 | 0.017927 | 3.1697e-06 | 7.2329e-10 | 5.53e-08 | 2.1771e-09 | 74.7782 | -59.6405 | 3.4028e+38 | 3.4028e+38 | -0.02699 | 460.0278 | 460.0282 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 2.0001 | -0.022471 | 460.0331 | 460.0329 | 50 | -90.0035 | -5.5978e-05 | 3.4028e+38 | 0 | 230.0068 | -0.039127 | 460.0374 | 460.0299 | -0.00077935 | 0.015802 | 3.6077e-06 | 3.9963e-09 | 5.7003e-08 | 3.2241e-09 | 74.77 | -33.8813 | 0.071574 | 3.4028e+38 | -68.9496 | 0.020425 | 0 | 0 | 0 | 6.6004e-05 | 0.016905 | 3.1011e-06 | 5.0793e-10 | 5.2397e-08 | 1.1458e-09 | 74.7577 | -43.0317 | 0.016257 | 3.4028e+38 | -49.3641 | 229.0397 | -1.7578e-11 | 0.00072627 | 1.9367e-10 | -2.9339e-05 | 3.4028e+38 | 3.4028e+38 | -0.025394 | 460.0307 | 460.0297 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -0.035888 | -0.035888 | 460.0323 | -0.00048963 | 230.0019 | 2.0001 | -0.028091 | 460.0287 | 460.029 | 49.9999 | -90.0035 | [17 empty] | 21.4791 | 71.9166 | -105778203.4723 | 94.2512 | 94.9617 |  | -0.0035124
32 | 230.0016 | 2.0001 | -0.027135 | 460.0287 | 460.0296 | 50 | -90.0034 | 0.016504 | 3.6074e-06 | 3.5747e-09 | 5.9066e-08 | 3.2387e-09 | 74.7803 | -37.9075 | 0.017902 | 3.0666e-06 | 1.4251e-09 | 5.5519e-08 | 2.7709e-09 | 74.7543 | -65.1993 | 3.4028e+38 | 3.4028e+38 | -0.027135 | 460.0287 | 460.0296 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 2.0001 | -0.029781 | 460.0289 | 460.0286 | 50 | -90.0034 | 0.014216 | 3.5961e-06 | 1.4724e-09 | 5.1128e-08 | 2.5374e-09 | 74.7139 | -62.5527 | 0.015457 | 3.0393e-06 | 4.2548e-10 | 4.8128e-08 | 6.8207e-10 | 74.7536 | -29.2429 | 3.4028e+38 | 3.4028e+38 | -0.029781 | 460.0289 | 460.0286 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | -0.026943 | 460.0283 | 460.0281 | 49.9999 | -90.0038 | 0.014441 | 3.6082e-06 | 2.5121e-09 | 5.2199e-08 | 1.9429e-09 | 74.7483 | -51.6693 | 0.015115 | 3.0676e-06 | -1.9471e-10 | 4.6332e-08 | 7.1398e-10 | 74.7347 | -43.1647 | 3.4028e+38 | 3.4028e+38 | -0.026943 | 460.0283 | 460.0281 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 2.0001 | -0.028689 | 460.0236 | 460.0236 | 49.9999 | -90.0037 | 0.016051 | 3.6151e-06 | 1.4591e-09 | 5.7623e-08 | 3.8672e-09 | 74.8263 | -82.0443 | 0.017381 | 3.1065e-06 | 2.47e-10 | 5.4007e-08 | 2.4019e-09 | 74.7995 | -79.1292 | 3.4028e+38 | 3.4028e+38 | -0.028689 | 460.0236 | 460.0236 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 2.0001 | -0.027328 | 460.0292 | 460.0292 | 49.9999 | -90.0035 | 0.015411 | 3.6145e-06 | 4.3077e-09 | 5.5708e-08 | 3.3714e-09 | 74.7017 | -43.4141 | 0.016783 | 3.1282e-06 | 1.2012e-09 | 4.9474e-08 | 1.3674e-09 | 74.698 | -53.8409 | 3.4028e+38 | 3.4028e+38 | -0.027328 | 460.0292 | 460.0292 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 2.0001 | -0.025571 | 460.0321 | 460.0333 | 49.9999 | -90.0037 | -6.4681e-05 | 3.4028e+38 | 0 | 230.0055 | -0.037046 | 460.0373 | 460.0322 | -0.00046371 | 0.015013 | 3.6071e-06 | 5.3619e-09 | 5.3857e-08 | 2.8179e-09 | 74.7228 | -44.1787 | 0.045644 | 3.4028e+38 | -34.9311 | 0.019913 | 0 | 0 | 0 | 7.1125e-05 | 0.015839 | 3.0483e-06 | 1.9679e-09 | 4.943e-08 | 9.2911e-10 | 74.7962 | -86.7268 | 0.0017522 | 3.4028e+38 | -58.4462 | 229.4889 | -7.7572e-10 | 0.00072018 | -6.6366e-10 | -2.5555e-05 | 3.4028e+38 | 3.4028e+38 | -0.026107 | 460.0298 | 460.03 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -0.038354 | -0.038871 | 460.0294 | -0.00014175 | 230.002 | 2.0001 | -0.027975 | 460.0277 | 460.0278 | 50 | -90.0036 | [17 empty] | 22.2485 | 91.7336 | -17200345.5123 | 111.154 | 111.3203 |  | -0.0037658
33 | 230.0019 | 2.0001 | 398.3788 | 460.0237 | 230.0368 | 50 | -30.0035 | 0.014819 | 3.571e-06 | 3.1457e-09 | 5.2924e-08 | 1.6859e-09 | 74.7565 | -41.4721 | 0.016043 | 3.2178e-06 | 1.9298e-09 | 5.0236e-08 | 2.7951e-09 | 74.7994 | -45.5309 | 3.4028e+38 | 3.4028e+38 | 398.3788 | 460.0237 | 230.0368 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | 398.3834 | 460.0291 | 230.0403 | 49.9999 | -30.0036 | 0.015522 | 3.6135e-06 | 3.1401e-09 | 5.404e-08 | 1.6839e-09 | 74.7637 | -31.7481 | 0.017068 | 3.1789e-06 | 3.5432e-09 | 5.2582e-08 | 3.0675e-09 | 74.8002 | -54.523 | 3.4028e+38 | 3.4028e+38 | 398.3834 | 460.0291 | 230.0403 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 2.0001 | 398.3843 | 460.0288 | 230.0374 | 49.9999 | -30.003 | 0.017223 | 3.5648e-06 | 3.9527e-09 | 6.1845e-08 | 6.4095e-10 | 74.8324 | -17.1159 | 0.019099 | 3.122e-06 | 3.0532e-09 | 5.9612e-08 | 4.0637e-09 | 74.6862 | -57.5509 | 3.4028e+38 | 3.4028e+38 | 398.3843 | 460.0288 | 230.0374 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | 398.3811 | 460.0279 | 230.0405 | 50 | -30.0034 | 0.014303 | 3.6042e-06 | 2.224e-09 | 5.156e-08 | 7.1298e-10 | 74.6837 | -20.0117 | 0.015352 | 3.1496e-06 | 8.1223e-10 | 4.8412e-08 | 8.6751e-10 | 74.6364 | -22.8761 | 3.4028e+38 | 3.4028e+38 | 398.3811 | 460.0279 | 230.0405 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 2.0001 | 398.3802 | 460.0278 | 230.0413 | 49.9999 | -30.0038 | 0.01469 | 3.6458e-06 | 3.7222e-09 | 5.1695e-08 | 1.8095e-09 | 74.7472 | -29.4449 | 0.01607 | 3.1074e-06 | 1.9714e-09 | 4.9963e-08 | 2.2098e-09 | 74.8198 | -39.0398 | 3.4028e+38 | 3.4028e+38 | 398.3802 | 460.0278 | 230.0413 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0022 | 2.0001 | 398.3839 | 460.029 | 230.0412 | 50 | -30.0038 | 0.866 | 3.4028e+38 | 0 | 230.0012 | 398.3799 | 460.0307 | 230.043 | -0.00071689 | 0.014548 | 3.602e-06 | 2.2186e-09 | 5.2117e-08 | 1.0691e-09 | 74.7866 | -29.5841 | 0.060679 | 3.4028e+38 | 33.6609 | 0.020099 | 0 | 0 | 0 | 4.5703e-06 | 0.01566 | 3.129e-06 | 2.3506e-09 | 4.896e-08 | 1.8866e-09 | 74.7515 | -49.7616 | 0.033321 | 3.4028e+38 | 63.6807 | 228.9907 | -1.1342e-10 | 0.00073605 | 1.5064e-11 | -2.5902e-05 | 3.4028e+38 | 3.4028e+38 | 398.3838 | 460.0257 | 230.0403 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 398.3719 | 398.3719 | 230.043 | -0.00075562 | 230.0018 | 2.0001 | 398.3816 | 460.0275 | 230.0393 | 49.9999 | -30.0034 | [17 empty] | 21.7997 | 18.0836 | 8.4384 | 39.9605 | 140.7254 |  | -0.0031041
34 | 230.0018 | 2.0001 | 398.3843 | 460.0289 | 230.0402 | 49.9999 | -30.0037 | 0.01654 | 3.5642e-06 | 5.3903e-09 | 5.9315e-08 | 1.4593e-09 | 74.664 | -19.8625 | 0.018277 | 3.121e-06 | 3.3625e-09 | 5.8636e-08 | 3.649e-09 | 74.8011 | -63.4859 | 3.4028e+38 | 3.4028e+38 | 398.3843 | 460.0289 | 230.0402 | 3.4028e+38 | 3.4028e+38 | 230.0012 | 2.0001 | 398.3854 | 460.0303 | 230.0388 | 49.9999 | -30.0035 | 0.016954 | 3.6503e-06 | 4.6172e-09 | 6.0712e-08 | -1.8227e-09 | 74.7376 | 6.8643 | 0.018523 | 3.1285e-06 | 3.3209e-09 | 5.7547e-08 | 1.8816e-09 | 74.7339 | -26.9118 | 3.4028e+38 | 3.4028e+38 | 398.3854 | 460.0303 | 230.0388 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 2.0001 | 398.3834 | 460.0314 | 230.0426 | 49.9999 | -30.0039 | 0.015435 | 3.5933e-06 | 3.7564e-09 | 5.5959e-08 | -1.0755e-09 | 74.701 | 8.4875 | 0.016706 | 3.167e-06 | 2.691e-09 | 5.1463e-08 | 8.4004e-10 | 74.8051 | -17.337 | 3.4028e+38 | 3.4028e+38 | 398.3834 | 460.0314 | 230.0426 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | 398.3849 | 460.0303 | 230.0352 | 49.9999 | -30.0032 | 0.016341 | 3.6353e-06 | 3.7901e-09 | 5.8251e-08 | -1.4879e-09 | 74.7644 | 42.4253 | 0.017471 | 3.167e-06 | 3.4104e-09 | 5.5487e-08 | 1.6354e-09 | 74.7638 | -43.2153 | 3.4028e+38 | 3.4028e+38 | 398.3849 | 460.0303 | 230.0352 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 2.0001 | 398.385 | 460.0294 | 230.0358 | 50 | -30.0042 | 0.015509 | 3.6006e-06 | 3.0574e-09 | 5.7099e-08 | -2.021e-09 | 74.686 | 21.7073 | 0.016649 | 3.0961e-06 | 2.7984e-09 | 5.3709e-08 | 5.7641e-10 | 74.6431 | -21.4496 | 3.4028e+38 | 3.4028e+38 | 398.385 | 460.0294 | 230.0358 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 2.0001 | 398.382 | 460.0291 | 230.0422 | 50 | -30.0038 | 0.86599 | 3.4028e+38 | 0 | 230.0071 | 398.3863 | 460.0411 | 230.0555 | 5.8728e-06 | 0.016251 | 3.6663e-06 | 4.3427e-09 | 5.96e-08 | -1.2478e-09 | 74.6619 | 25.425 | 0.064104 | 3.4028e+38 | -50.9183 | 0.020038 | 0 | 0 | 0 | -1.5766e-05 | 0.017892 | 3.1886e-06 | 2.8772e-09 | 5.695e-08 | 2.3269e-09 | 74.6629 | -31.5836 | 0.068152 | 3.4028e+38 | 7.1259 | 229.6711 | -5.6917e-10 | 0.00071893 | 3.0245e-10 | -3.2934e-05 | 3.4028e+38 | 3.4028e+38 | 398.383 | 460.0319 | 230.0455 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 398.3862 | 398.3862 | 230.0591 | 7.1581e-05 | 230.0018 | 2.0001 | 398.3846 | 460.03 | 230.0385 | 49.9999 | -30.0037 | [17 empty] | 21.6613 | 18.8242 | 9.5455 | 40.9933 | 138.9151 |  | -0.0035591
35 | 230.0018 | 2.0001 | 398.3814 | 460.0251 | 230.0394 | 50 | -30.0035 | 0.017149 | 3.5657e-06 | 4.9898e-09 | 6.0468e-08 | 1.6219e-09 | 74.7623 | -14.1619 | 0.01868 | 3.1947e-06 | 4.9712e-09 | 5.7755e-08 | 5.2759e-09 | 74.7757 | -46.7701 | 3.4028e+38 | 3.4028e+38 | 398.3814 | 460.0251 | 230.0394 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 2.0001 | 398.3833 | 460.0294 | 230.0396 | 49.9999 | -30.0035 | 0.017319 | 3.562e-06 | 4.8607e-09 | 6.0607e-08 | -7.6805e-10 | 74.7901 | 9.4718 | 0.018676 | 3.1543e-06 | 3.776e-09 | 5.8259e-08 | 3.263e-09 | 74.6571 | -35.8132 | 3.4028e+38 | 3.4028e+38 | 398.3833 | 460.0294 | 230.0396 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 2.0001 | 398.3846 | 460.0324 | 230.0434 | 49.9999 | -30.0037 | 0.014822 | 3.6462e-06 | 2.8998e-09 | 5.4047e-08 | -2.464e-10 | 74.7867 | 5.3857 | 0.015784 | 3.1587e-06 | 1.8777e-09 | 4.997e-08 | 6.185e-10 | 74.7548 | -36.0443 | 3.4028e+38 | 3.4028e+38 | 398.3846 | 460.0324 | 230.0434 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 2.0001 | 398.3868 | 460.0332 | 230.0412 | 49.9999 | -30.0035 | 0.017834 | 3.6176e-06 | 4.7675e-09 | 6.3539e-08 | -7.8995e-10 | 74.7938 | 10.0735 | 0.019432 | 3.12e-06 | 5.9334e-09 | 6.2511e-08 | 5.1102e-09 | 74.6553 | -40.7373 | 3.4028e+38 | 3.4028e+38 | 398.3868 | 460.0332 | 230.0412 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 2.0001 | 398.3828 | 460.0279 | 230.0373 | 49.9999 | -30.0031 | 0.017764 | 3.5714e-06 | 5.3812e-09 | 6.3093e-08 | -1.246e-09 | 74.6441 | 4.6116 | 0.019297 | 3.1218e-06 | 4.5985e-09 | 6.1019e-08 | 3.3541e-09 | 74.77 | -50.5763 | 3.4028e+38 | 3.4028e+38 | 398.3828 | 460.0279 | 230.0373 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0016 | 2.0001 | 398.3849 | 460.0279 | 230.04 | 49.9999 | -30.0037 | 0.866 | 3.4028e+38 | 0 | 230.0035 | 398.3777 | 460.0335 | 230.0511 | -0.00087594 | 0.01739 | 3.6121e-06 | 5.7269e-09 | 6.2824e-08 | 1.2628e-09 | 74.7885 | -18.7533 | 0.092567 | 3.4028e+38 | 4.9539 | 0.020307 | 0 | 0 | 0 | -1.4582e-05 | 0.01914 | 3.0489e-06 | 3.5376e-09 | 6.0239e-08 | 4.2121e-09 | 74.7217 | -48.3638 | 0.062005 | 3.4028e+38 | -30.1597 | 229.0957 | 1.5157e-10 | 0.00073105 | 5.3727e-10 | -3.4753e-05 | 3.4028e+38 | 3.4028e+38 | 398.3824 | 460.0308 | 230.041 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 398.3714 | 398.3714 | 230.0401 | -0.0010824 | 230.0018 | 2.0001 | 398.3838 | 460.0296 | 230.0402 | 49.9999 | -30.0035 | [17 empty] | 21.739 | 14.8394 | 3.0061 | 36.8537 | 146.9226 |  | -0.0034568
36 | 230.0016 | 5.0002 | -0.062986 | 1150.0656 | 1150.0663 | 50 | -90.0031 | 0.016583 | 3.8863e-06 | 4.6858e-09 | 6.3134e-08 | 1.2144e-08 | 74.7403 | -71.6019 | 0.017879 | 3.2095e-06 | -2.1281e-09 | 5.7481e-08 | 3.5555e-09 | 74.7966 | -110.4101 | 3.4028e+38 | 3.4028e+38 | -0.062986 | 1150.0656 | 1150.0663 | 3.4028e+38 | 3.4028e+38 | 230.002 | 5.0003 | -0.062064 | 1150.0684 | 1150.0653 | 50 | -90.0034 | 0.016634 | 3.8708e-06 | 5.0197e-09 | 6.3393e-08 | 1.23e-08 | 74.7068 | -65.6918 | 0.018205 | 3.1569e-06 | -2.2988e-09 | 5.7735e-08 | 5.0308e-09 | 74.6406 | -105.1457 | 3.4028e+38 | 3.4028e+38 | -0.062064 | 1150.0684 | 1150.0653 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 5.0003 | -0.060288 | 1150.069 | 1150.0695 | 50 | -90.0031 | 0.015433 | 3.8591e-06 | 9.3798e-09 | 5.963e-08 | 6.5296e-09 | 74.7281 | -35.6058 | 0.016726 | 3.139e-06 | 1.3932e-09 | 5.1802e-08 | 4.4065e-09 | 74.7012 | -66.2458 | 3.4028e+38 | 3.4028e+38 | -0.060288 | 1150.069 | 1150.0695 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 5.0003 | -0.065812 | 1150.0717 | 1150.0686 | 49.9999 | -90.0031 | 0.015351 | 3.8648e-06 | 3.1859e-09 | 6.0579e-08 | 8.999e-09 | 74.6252 | -74.461 | 0.016558 | 3.1495e-06 | -1.7529e-09 | 5.2944e-08 | 3.6049e-09 | 74.7254 | -110.9472 | 3.4028e+38 | 3.4028e+38 | -0.065812 | 1150.0717 | 1150.0686 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 5.0003 | -0.062139 | 1150.0656 | 1150.0665 | 50 | -90.0032 | 0.017204 | 3.8264e-06 | 8.5536e-09 | 6.5912e-08 | 1.4637e-08 | 74.6936 | -59.5513 | 0.018847 | 3.2399e-06 | -1.3852e-09 | 6.117e-08 | 5.1763e-09 | 74.7772 | -109.1347 | 3.4028e+38 | 3.4028e+38 | -0.062139 | 1150.0656 | 1150.0665 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0015 | 5.0002 | -0.064928 | 1150.0642 | 1150.066 | 49.9999 | -90.0033 | -5.5706e-05 | 3.4028e+38 | 0 | 230.0032 | -0.11675 | 1150.075 | 1150.0482 | -0.0044843 | 0.01778 | 3.7566e-06 | 1.2176e-08 | 6.9445e-08 | 1.3763e-08 | 74.6811 | -48.5005 | 0.15998 | 3.4028e+38 | -54.4723 | 0.021016 | 0 | 0 | 0 | 0.00014215 | 0.019444 | 3.1701e-06 | 1.1729e-09 | 6.159e-08 | 7.3154e-09 | 74.8158 | -85.928 | 0.0082335 | 3.4028e+38 | 8.5457 | 229.1731 | -1.3169e-10 | 0.00074261 | 1.0285e-09 | -9.0251e-05 | 3.4028e+38 | 3.4028e+38 | -0.058312 | 1150.0657 | 1150.0688 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -0.099844 | -0.099844 | 1150.0667 | -0.0032729 | 230.0019 | 5.0003 | -0.062658 | 1150.068 | 1150.0672 | 50 | -90.0032 | [17 empty] | 22.7719 | 69.3529 | -11365379.2547 | 92.1415 | 91.4563 |  | -0.0034933
37 | 230.0022 | 5.0003 | -0.060698 | 1150.0699 | 1150.0706 | 49.9999 | -90.0031 | 0.014835 | 3.9474e-06 | 8.9625e-09 | 5.8347e-08 | 5.8252e-09 | 74.6601 | -32.9119 | 0.016112 | 3.1129e-06 | 1.4973e-09 | 5.0173e-08 | 3.5242e-09 | 74.682 | -53.4978 | 3.4028e+38 | 3.4028e+38 | -0.060698 | 1150.0699 | 1150.0706 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 5.0003 | -0.063395 | 1150.0669 | 1150.0662 | 50 | -90.0033 | 0.017173 | 3.8794e-06 | 8.6774e-09 | 6.7072e-08 | 1.4656e-08 | 74.6746 | -62.18 | 0.01896 | 3.2028e-06 | -1.702e-09 | 6.0479e-08 | 5.6227e-09 | 74.7838 | -101.493 | 3.4028e+38 | 3.4028e+38 | -0.063395 | 1150.0669 | 1150.0662 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 5.0002 | -0.060201 | 1150.0684 | 1150.0682 | 49.9999 | -90.0031 | 0.017664 | 3.816e-06 | 1.2794e-08 | 6.6874e-08 | 1.1181e-08 | 74.737 | -42.2317 | 0.019312 | 3.2237e-06 | 2.0237e-09 | 6.1212e-08 | 8.9478e-09 | 74.7151 | -79.9351 | 3.4028e+38 | 3.4028e+38 | -0.060201 | 1150.0684 | 1150.0682 | 3.4028e+38 | 3.4028e+38 | 230.002 | 5.0003 | -0.062583 | 1150.0653 | 1150.0647 | 50 | -90.003 | 0.017597 | 3.8275e-06 | 1.0763e-08 | 6.7459e-08 | 1.4688e-08 | 74.6079 | -57.6479 | 0.019377 | 3.2198e-06 | -5.5093e-10 | 6.1825e-08 | 7.7755e-09 | 74.7521 | -87.2768 | 3.4028e+38 | 3.4028e+38 | -0.062583 | 1150.0653 | 1150.0647 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 5.0003 | -0.059181 | 1150.0706 | 1150.0688 | 50 | -90.0031 | 0.014955 | 3.8901e-06 | 3.4513e-09 | 5.6419e-08 | 8.4019e-09 | 74.7893 | -65.7902 | 0.015802 | 3.1629e-06 | -1.6234e-09 | 5.0939e-08 | 1.5339e-09 | 74.6731 | -135.1143 | 3.4028e+38 | 3.4028e+38 | -0.059181 | 1150.0706 | 1150.0688 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0018 | 5.0003 | -0.06718 | 1150.0702 | 1150.0702 | 49.9999 | -90.0031 | -5.4059e-05 | 3.4028e+38 | 0 | 230.0018 | -0.08491 | 1150.0674 | 1150.057 | -0.00075606 | 0.0144 | 3.8491e-06 | 3.2748e-09 | 5.6325e-08 | 7.3087e-09 | 74.7332 | -58.2197 | 0.08437 | 3.4028e+38 | 53.0224 | 0.02015 | 0 | 0 | 0 | 0.00016643 | 0.015438 | 3.2083e-06 | -4.1947e-10 | 4.964e-08 | 2.754e-09 | 74.6857 | -89.1646 | 0.00082948 | 3.4028e+38 | -14.5606 | 229.3472 | -7.1032e-10 | 0.00072101 | 4.404e-10 | -7.459e-05 | 3.4028e+38 | 3.4028e+38 | -0.066574 | 1150.0688 | 1150.0721 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -0.078802 | -0.078802 | 1150.0582 | -0.00073414 | 230.0018 | 5.0003 | -0.061212 | 1150.0682 | 1150.0677 | 49.9999 | -90.0031 | [17 empty] | 22.2534 | 66.0117 | -10433720.8505 | 88.1015 | 87.6496 |  | -0.0034452
38 | 230.0013 | 5.0003 | -0.062232 | 1150.0691 | 1150.0692 | 50 | -90.0032 | 0.016887 | 3.9235e-06 | 1.5218e-08 | 6.7409e-08 | 1.1968e-08 | 74.6826 | -39.7323 | 0.018463 | 3.184e-06 | -7.7071e-11 | 5.883e-08 | 7.0758e-09 | 74.7142 | -75.4326 | 3.4028e+38 | 3.4028e+38 | -0.062232 | 1150.0691 | 1150.0692 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 5.0003 | -0.063578 | 1150.0688 | 1150.0695 | 50 | -90.0032 | 0.014855 | 3.8662e-06 | 2.2987e-09 | 5.7394e-08 | 7.0094e-09 | 74.5988 | -61.5916 | 0.015915 | 3.1336e-06 | -1.5277e-09 | 4.9902e-08 | 1.7133e-09 | 74.915 | -101.5795 | 3.4028e+38 | 3.4028e+38 | -0.063578 | 1150.0688 | 1150.0695 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 5.0003 | -0.066729 | 1150.0687 | 1150.0697 | 49.9999 | -90.0032 | 0.016496 | 3.8342e-06 | 5.4722e-09 | 6.327e-08 | 1.3515e-08 | 74.5554 | -64.3823 | 0.017778 | 3.2352e-06 | -7.0655e-10 | 5.8623e-08 | 3.7543e-09 | 74.5881 | -104.6264 | 3.4028e+38 | 3.4028e+38 | -0.066729 | 1150.0687 | 1150.0697 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 5.0003 | -0.060364 | 1150.0736 | 1150.0737 | 50 | -90.0031 | 0.014435 | 3.8805e-06 | 5.6992e-09 | 5.4411e-08 | 6.2868e-09 | 74.6555 | -48.0201 | 0.015547 | 3.1323e-06 | 2.6927e-10 | 4.8946e-08 | 2.286e-09 | 74.773 | -83.282 | 3.4028e+38 | 3.4028e+38 | -0.060364 | 1150.0736 | 1150.0737 | 3.4028e+38 | 3.4028e+38 | 230.002 | 5.0003 | -0.057782 | 1150.0702 | 1150.0701 | 50 | -90.0033 | 0.014252 | 3.8786e-06 | 4.3831e-09 | 5.5289e-08 | 5.5604e-09 | 74.7514 | -50.9096 | 0.015527 | 3.095e-06 | 1.853e-10 | 4.7614e-08 | 1.8906e-09 | 74.7046 | -85.1388 | 3.4028e+38 | 3.4028e+38 | -0.057782 | 1150.0702 | 1150.0701 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0021 | 5.0003 | -0.063096 | 1150.0706 | 1150.0734 | 49.9999 | -90.0033 | -5.7838e-05 | 3.4028e+38 | 0 | 230.0045 | -0.081177 | 1150.0819 | 1150.0721 | -0.00018224 | 0.014516 | 3.8914e-06 | 7.5212e-09 | 5.6506e-08 | 5.8128e-09 | 74.6964 | -41.5672 | 0.12829 | 3.4028e+38 | 52.535 | 0.019779 | 0 | 0 | 0 | 0.00016384 | 0.01572 | 3.1625e-06 | 8.6805e-10 | 4.8826e-08 | 2.5645e-09 | 74.8141 | -65.9081 | 0.023486 | 3.4028e+38 | -58.6671 | 229.0887 | 4.0838e-10 | 0.00072514 | 1.0405e-09 | -7.3422e-05 | 3.4028e+38 | 3.4028e+38 | -0.063955 | 1150.0698 | 1150.0725 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -0.085746 | -0.082446 | 1150.0754 | -0.0013312 | 230.0019 | 5.0003 | -0.062137 | 1150.0701 | 1150.0704 | 50 | -90.0032 | [17 empty] | 21.7807 | 76.8186 | -15867683.4285 | 98.9104 | 99.2017 |  | -0.0034065
39 | 230.0019 | 5.0003 | 995.9528 | 1150.0614 | 575.0889 | 49.9999 | -30.0029 | 0.017681 | 3.8267e-06 | 1.3882e-08 | 6.7611e-08 | 1.5636e-09 | 74.5836 | -6.0989 | 0.019399 | 3.2356e-06 | 7.2306e-09 | 6.3166e-08 | 8.2997e-09 | 74.8078 | -42.0286 | 3.4028e+38 | 3.4028e+38 | 995.9528 | 1150.0614 | 575.0889 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 5.0003 | 995.9589 | 1150.0669 | 575.0872 | 49.9999 | -30.003 | 0.017582 | 3.8187e-06 | 1.2815e-08 | 6.7905e-08 | 7.8308e-10 | 74.7598 | -1.1782 | 0.019605 | 3.2309e-06 | 8.7389e-09 | 6.2149e-08 | 1.0196e-08 | 74.6422 | -45.4144 | 3.4028e+38 | 3.4028e+38 | 995.9589 | 1150.0669 | 575.0872 | 3.4028e+38 | 3.4028e+38 | 230.002 | 5.0002 | 995.9562 | 1150.0651 | 575.0851 | 50 | -30.003 | 0.015946 | 3.8458e-06 | 9.7165e-09 | 6.1482e-08 | 5.8927e-09 | 74.8177 | -32.4184 | 0.017015 | 3.3203e-06 | 4.3955e-09 | 5.6452e-08 | 7.5703e-09 | 74.7767 | -55.0477 | 3.4028e+38 | 3.4028e+38 | 995.9562 | 1150.0651 | 575.0851 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 5.0003 | 995.9623 | 1150.0714 | 575.0893 | 49.9999 | -30.0031 | 0.01453 | 3.9066e-06 | 5.4955e-09 | 5.669e-08 | 9.0593e-11 | 74.729 | -0.89523 | 0.015726 | 3.2e-06 | 2.7371e-09 | 4.9478e-08 | 2.1916e-09 | 74.8257 | -41.2087 | 3.4028e+38 | 3.4028e+38 | 995.9623 | 1150.0714 | 575.0893 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 5.0003 | 995.9619 | 1150.0706 | 575.0922 | 49.9999 | -30.0031 | 0.015997 | 3.9076e-06 | 9.9734e-09 | 6.423e-08 | -3.2077e-09 | 74.6977 | 18.3985 | 0.017388 | 3.1253e-06 | 7.3376e-09 | 5.7922e-08 | 4.7545e-09 | 74.7991 | -32.9601 | 3.4028e+38 | 3.4028e+38 | 995.9619 | 1150.0706 | 575.0922 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 5.0003 | 995.9611 | 1150.0695 | 575.0879 | 50 | -30.003 | 0.866 | 3.4028e+38 | 0 | 230.0071 | 995.9618 | 1150.0942 | 575.1323 | 0.00051985 | 0.016912 | 3.8733e-06 | 1.0808e-08 | 6.6173e-08 | -4.172e-09 | 74.6956 | 13.528 | 0.17384 | 3.4028e+38 | -32.1967 | 0.020167 | 0 | 0 | 0 | -3.2772e-05 | 0.01893 | 3.1551e-06 | 8.0936e-09 | 6.0053e-08 | 7.9082e-09 | 74.7633 | 0 | 0 | 3.4028e+38 | -52.8411 | 229.263 | 1.8564e-09 | 0.00074899 | 1.0213e-09 | -8.36e-05 | 3.4028e+38 | 3.4028e+38 | 995.9594 | 1150.0673 | 575.0853 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 995.9442 | 995.9472 | 575.1264 | -0.0011559 | 230.002 | 5.0003 | 995.9584 | 1150.0671 | 575.0886 | 49.9999 | -30.003 | [17 empty] | 22.488 | 4.3741 | -4.9288 | 24.8252 | 122.7614 |  | -0.003092
40 | 230.0022 | 5.0002 | 995.956 | 1150.0658 | 575.0917 | 49.9999 | -30.0032 | 0.015183 | 3.7983e-06 | 8.1463e-09 | 5.7561e-08 | 5.3966e-09 | 74.7453 | -31.139 | 0.016515 | 3.2909e-06 | 2.5029e-09 | 5.3469e-08 | 5.9487e-09 | 74.769 | -62.0676 | 3.4028e+38 | 3.4028e+38 | 995.956 | 1150.0658 | 575.0917 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 5.0003 | 995.9599 | 1150.0632 | 575.0832 | 49.9999 | -30.0029 | 0.017516 | 3.8504e-06 | 1.3088e-08 | 6.715e-08 | 4.2664e-09 | 74.717 | -16.0655 | 0.019221 | 3.2825e-06 | 7.041e-09 | 6.1991e-08 | 1.0486e-08 | 74.691 | -51.9995 | 3.4028e+38 | 3.4028e+38 | 995.9599 | 1150.0632 | 575.0832 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 5.0003 | 995.9635 | 1150.0717 | 575.0902 | 49.9999 | -30.0033 | 0.014716 | 3.9e-06 | 4.4423e-09 | 5.5897e-08 | -2.3388e-09 | 74.6614 | 15.7813 | 0.016056 | 3.277e-06 | 3.6932e-09 | 4.9608e-08 | 2.0689e-09 | 74.7361 | -28.249 | 3.4028e+38 | 3.4028e+38 | 995.9635 | 1150.0717 | 575.0901 | 3.4028e+38 | 3.4028e+38 | 230.002 | 5.0003 | 995.9568 | 1150.0652 | 575.0887 | 49.9999 | -30.0029 | 0.015286 | 3.8073e-06 | 8.7262e-09 | 5.9753e-08 | 5.9899e-09 | 74.7064 | -30.1662 | 0.016442 | 3.2462e-06 | 2.3808e-09 | 5.3154e-08 | 5.8119e-09 | 74.7135 | -70.4193 | 3.4028e+38 | 3.4028e+38 | 995.9568 | 1150.0652 | 575.0887 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 5.0003 | 995.9606 | 1150.0709 | 575.0911 | 50 | -30.0031 | 0.016971 | 3.861e-06 | 1.1057e-08 | 6.552e-08 | -3.1594e-09 | 74.7052 | 15.6398 | 0.018279 | 3.2805e-06 | 8.0778e-09 | 5.995e-08 | 6.1565e-09 | 74.6639 | -37.0147 | 3.4028e+38 | 3.4028e+38 | 995.9606 | 1150.0709 | 575.091 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 5.0003 | 995.9568 | 1150.0647 | 575.0854 | 50 | -30.0031 | 0.866 | 3.4028e+38 | 0 | 230.0076 | 995.9645 | 1150.0962 | 575.1277 | -0.00018247 | 0.016312 | 3.9532e-06 | 7.2581e-09 | 6.4761e-08 | -4.0585e-09 | 74.7606 | 25.5405 | 0.12472 | 3.4028e+38 | -8.6658 | 0.020551 | 0 | 0 | 0 | -2.6177e-05 | 0.017579 | 3.1484e-06 | 5.9309e-09 | 5.5039e-08 | 3.8523e-09 | 74.7171 | -30.4207 | 0.0988 | 3.4028e+38 | 37.4644 | 229.4368 | 6.7469e-10 | 0.00072825 | -7.023e-10 | -7.3057e-05 | 3.4028e+38 | 3.4028e+38 | 995.957 | 1150.0653 | 575.0874 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 995.9687 | 995.9687 | 575.1198 | 0.0001948 | 230.0019 | 5.0003 | 995.9594 | 1150.0674 | 575.089 | 49.9999 | -30.0031 | [17 empty] | 22.3533 | 8.6424 | -0.12612 | 29.4808 | 129.5459 |  | -0.0031927
41 | 230.0019 | 5.0003 | 995.9635 | 1150.0729 | 575.0887 | 50 | -30.0032 | 0.014431 | 3.8779e-06 | 4.5414e-09 | 5.6453e-08 | -1.5804e-09 | 74.6704 | 8.0025 | 0.015069 | 3.2885e-06 | 4.5349e-09 | 4.8453e-08 | 3.0905e-09 | 74.7862 | -49.5957 | 3.4028e+38 | 3.4028e+38 | 995.9635 | 1150.0729 | 575.0886 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 5.0003 | 995.9612 | 1150.0631 | 575.0859 | 49.9999 | -30.003 | 0.016951 | 3.8741e-06 | 1.3397e-08 | 6.6811e-08 | -1.032e-09 | 74.7017 | 4.7053 | 0.018761 | 3.2951e-06 | 7.7847e-09 | 5.9478e-08 | 5.683e-09 | 74.782 | -41.3601 | 3.4028e+38 | 3.4028e+38 | 995.9612 | 1150.0631 | 575.0858 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 5.0002 | 995.9573 | 1150.0659 | 575.0932 | 49.9999 | -30.0033 | 0.014347 | 3.8856e-06 | 5.3602e-09 | 5.5704e-08 | -2.6015e-10 | 74.6712 | 2.3753 | 0.015223 | 3.2475e-06 | 3.1493e-09 | 4.9397e-08 | 2.3724e-09 | 74.8091 | -36.4849 | 3.4028e+38 | 3.4028e+38 | 995.9573 | 1150.0659 | 575.0931 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 5.0003 | 995.9562 | 1150.0652 | 575.0876 | 49.9999 | -30.0031 | 0.017192 | 3.8111e-06 | 1.3115e-08 | 6.4914e-08 | 6.4269e-09 | 74.7453 | -25.0446 | 0.018682 | 3.2004e-06 | 7.0223e-09 | 6.0211e-08 | 9.7338e-09 | 74.7986 | -54.834 | 3.4028e+38 | 3.4028e+38 | 995.9562 | 1150.0652 | 575.0876 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 5.0003 | 995.96 | 1150.0774 | 575.0918 | 50 | -30.003 | 0.015555 | 3.8581e-06 | 8.1231e-09 | 6.0005e-08 | -4.0171e-09 | 74.6996 | 27.5929 | 0.016946 | 3.1716e-06 | 5.0845e-09 | 5.3731e-08 | 3.7253e-09 | 74.66 | -35.4146 | 3.4028e+38 | 3.4028e+38 | 995.96 | 1150.0774 | 575.0917 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0017 | 5.0003 | 995.9594 | 1150.0698 | 575.0933 | 49.9999 | -30.0033 | 0.866 | 3.4028e+38 | 0 | 230.0066 | 995.9657 | 1150.0911 | 575.1185 | -2.9686e-06 | 0.015514 | 3.8676e-06 | 5.7773e-09 | 5.872e-08 | -3.6903e-09 | 74.83 | 28.5022 | 0.10895 | 3.4028e+38 | -29.6367 | 0.020157 | 0 | 0 | 0 | -1.9115e-05 | 0.016475 | 3.1763e-06 | 5.0754e-09 | 5.11e-08 | 2.6525e-09 | 74.7543 | -28.1428 | 0.09704 | 3.4028e+38 | -43.296 | 229.7026 | 6.8448e-10 | 0.00072776 | 2.0412e-10 | -6.9125e-05 | 3.4028e+38 | 3.4028e+38 | 995.9595 | 1150.0696 | 575.0836 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 995.9677 | 995.9677 | 575.1221 | -0.00028342 | 230.0017 | 5.0003 | 995.9596 | 1150.0689 | 575.0894 | 50 | -30.0031 | [17 empty] | 21.559 | 6.6756 | -3.3212 | 29.4568 | 135.2816 |  | -0.0034713
42 | 230.002 | 2.0001 | -276.0438 | 460.0287 | 368.0076 | 50 | -126.874 | 0.017139 | 3.5764e-06 | -6.4275e-10 | 6.0899e-08 | 4.5283e-09 | 74.7633 | -103.8887 | 0.018804 | 3.1658e-06 | 2.7094e-11 | 5.702e-08 | 2.2414e-09 | 74.7129 | -110.6845 | 3.4028e+38 | 3.4028e+38 | -276.0438 | 460.0287 | 368.0076 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 2.0001 | -276.0498 | 460.0291 | 368.0031 | 49.9999 | -126.8738 | 0.016712 | 3.6045e-06 | -2.6803e-09 | 5.9854e-08 | 3.2095e-09 | 74.7555 | -115.8095 | 0.018176 | 3.1529e-06 | -8.1781e-10 | 5.7209e-08 | 1.2621e-09 | 74.7276 | -102.7368 | 3.4028e+38 | 3.4028e+38 | -276.0498 | 460.0291 | 368.0031 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 2.0001 | -276.0452 | 460.0303 | 368.0077 | 50.0001 | -126.8739 | 0.017723 | 3.6385e-06 | 2.4938e-09 | 6.2561e-08 | 5.0289e-09 | 74.737 | -69.5089 | 0.019349 | 3.1256e-06 | -2.5745e-10 | 6.2638e-08 | 2.2668e-09 | 74.7798 | -90.404 | 3.4028e+38 | 3.4028e+38 | -276.0452 | 460.0303 | 368.0077 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 2.0001 | -276.0441 | 460.0265 | 368.005 | 49.9999 | -126.8737 | 0.016291 | 3.6e-06 | -2.3736e-09 | 5.8804e-08 | 3.2646e-09 | 74.8124 | -132.55 | 0.017649 | 3.0983e-06 | -6.2317e-10 | 5.4662e-08 | 6.5405e-10 | 74.7979 | -156.3179 | 3.4028e+38 | 3.4028e+38 | -276.0441 | 460.0265 | 368.005 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | -276.047 | 460.0291 | 368.0057 | 50 | -126.8741 | 0.015789 | 3.6064e-06 | -1.7984e-09 | 5.6934e-08 | 3.04e-09 | 74.8733 | -113.8045 | 0.017078 | 3.1426e-06 | 9.1662e-10 | 5.3542e-08 | 9.5655e-10 | 74.785 | -141.4921 | 3.4028e+38 | 3.4028e+38 | -276.047 | 460.0291 | 368.0057 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 2.0001 | -276.0469 | 460.0312 | 368.0082 | 49.9999 | -126.8741 | -0.60006 | 3.4028e+38 | 0 | 230.0011 | -276.0625 | 460.028 | 367.9874 | -0.001134 | 0.017019 | 3.5629e-06 | -1.5636e-09 | 6.0678e-08 | 4.0786e-09 | 74.6891 | -106.2243 | -0.021213 | 3.4028e+38 | 8.025 | 0.020843 | 0 | 0 | 0 | 7.0618e-05 | 0.019325 | 3.1314e-06 | -1.1171e-09 | 6.0251e-08 | 2.268e-09 | 74.8155 | -97.6376 | -0.005183 | 3.4028e+38 | -14.5887 | 229.524 | -7.2824e-10 | 0.00072228 | 2.6517e-10 | -3.1767e-05 | 3.4028e+38 | 3.4028e+38 | -276.045 | 460.0263 | 368.0037 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -276.0633 | -276.0681 | 367.9894 | -0.001451 | 230.0018 | 2.0001 | -276.046 | 460.0287 | 368.0058 | 50 | -126.8739 | [17 empty] | 21.0336 | 77.8299 | 124.2447 | 98.7903 | 96.2696 |  | -0.00063297
43 | 230.0018 | 2.0001 | -276.0432 | 460.0268 | 368.0033 | 49.9999 | -126.8741 | 0.01634 | 3.6106e-06 | -2.6326e-09 | 5.8966e-08 | 3.548e-09 | 74.7181 | -128.8882 | 0.017596 | 3.0915e-06 | 6.9386e-10 | 5.44e-08 | 1.3377e-09 | 74.7233 | -62.1445 | 3.4028e+38 | 3.4028e+38 | -276.0433 | 460.0268 | 368.0033 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 2.0001 | -276.0534 | 460.0291 | 368.0071 | 49.9999 | -126.8739 | 0.014525 | 3.5237e-06 | 1.3608e-09 | 5.2145e-08 | 2.8875e-09 | 74.801 | -49.4158 | 0.015608 | 3.0543e-06 | 1.2434e-10 | 4.8493e-08 | 5.0711e-10 | 74.7938 | -114.7033 | 3.4028e+38 | 3.4028e+38 | -276.0534 | 460.0291 | 368.0071 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 2.0001 | -276.0472 | 460.0292 | 368.0089 | 50 | -126.8738 | 0.015664 | 3.5778e-06 | 2.6724e-09 | 5.7141e-08 | 4.0819e-09 | 74.8139 | -56.3583 | 0.016981 | 3.0764e-06 | -3.0531e-10 | 5.493e-08 | 1.1352e-09 | 74.7303 | -132.7348 | 3.4028e+38 | 3.4028e+38 | -276.0472 | 460.0292 | 368.0089 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | -276.0475 | 460.0294 | 368.0052 | 49.9999 | -126.874 | 0.017092 | 3.6109e-06 | -1.2557e-09 | 6.1776e-08 | 3.9954e-09 | 74.8176 | -99.4911 | 0.019005 | 3.0595e-06 | -2.3069e-10 | 5.8138e-08 | 7.457e-10 | 74.74 | -120.8802 | 3.4028e+38 | 3.4028e+38 | -276.0475 | 460.0294 | 368.0052 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 2.0001 | -276.0388 | 460.0272 | 368.0062 | 49.9999 | -126.8737 | 0.017708 | 3.6334e-06 | -1.708e-09 | 6.4131e-08 | 4.1037e-09 | 74.7252 | -93.1023 | 0.019698 | 3.1252e-06 | -7.107e-10 | 6.0042e-08 | 2.675e-09 | 74.7983 | -86.4893 | 3.4028e+38 | 3.4028e+38 | -276.0388 | 460.0272 | 368.0062 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0022 | 2.0001 | -276.0436 | 460.0276 | 368.0082 | 50 | -126.8738 | -0.60006 | 3.4028e+38 | 0 | 230.0016 | -276.0632 | 460.0295 | 367.9847 | -0.000951 | 0.017307 | 3.549e-06 | -6.1328e-10 | 6.0095e-08 | 3.4728e-09 | 74.7894 | -112.0122 | -0.028327 | 3.4028e+38 | 23.3991 | 0.019951 | 0 | 0 | 0 | 7.6425e-05 | 0.017938 | 3.0587e-06 | 5.8733e-10 | 5.5553e-08 | 1.5722e-09 | 74.7593 | -73.6357 | 0.0083101 | 3.4028e+38 | 31.2474 | 229.7215 | -3.9808e-10 | 0.00073281 | 1.525e-09 | -1.8687e-05 | 3.4028e+38 | 3.4028e+38 | -276.0468 | 460.029 | 368.0059 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -276.0529 | -276.0529 | 367.992 | 8.2044e-06 | 230.0019 | 2.0001 | -276.046 | 460.0283 | 368.0061 | 49.9999 | -126.8739 | [17 empty] | 21.6507 | 78.5032 | 126.9443 | 99.061 | 97.6167 |  | -0.00065458
44 | 230.0017 | 2.0001 | -276.0464 | 460.0307 | 368.0083 | 49.9999 | -126.8742 | 0.015422 | 3.5745e-06 | -1.383e-09 | 5.4254e-08 | 2.0575e-09 | 74.6239 | -141.4597 | 0.016444 | 3.1369e-06 | -2.0785e-10 | 5.147e-08 | 7.035e-10 | 74.7593 | -137.3803 | 3.4028e+38 | 3.4028e+38 | -276.0464 | 460.0307 | 368.0083 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 2.0001 | -276.0464 | 460.0296 | 368.0068 | 49.9999 | -126.8744 | 0.016981 | 3.5976e-06 | 3.4681e-09 | 6.0515e-08 | 5.6882e-09 | 74.8147 | -53.161 | 0.018815 | 3.1628e-06 | 1.4188e-09 | 5.8039e-08 | 1.6775e-09 | 74.7599 | -126.3479 | 3.4028e+38 | 3.4028e+38 | -276.0464 | 460.0296 | 368.0068 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 2.0001 | -276.0468 | 460.0301 | 368.0072 | 49.9999 | -126.874 | 0.017149 | 3.5879e-06 | -3.6737e-10 | 6.1492e-08 | 4.6359e-09 | 74.7023 | -106.7349 | 0.018833 | 3.1106e-06 | -1.327e-09 | 5.8596e-08 | 1.7948e-09 | 74.7307 | -121.2068 | 3.4028e+38 | 3.4028e+38 | -276.0468 | 460.0301 | 368.0072 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | -276.0472 | 460.0305 | 368.0071 | 49.9999 | -126.874 | 0.015344 | 3.6046e-06 | -1.359e-09 | 5.6923e-08 | 3.7701e-09 | 74.7829 | -115.9093 | 0.016994 | 3.1586e-06 | -4.214e-10 | 5.2142e-08 | 1.1393e-09 | 74.7886 | -150.3386 | 3.4028e+38 | 3.4028e+38 | -276.0472 | 460.0305 | 368.0071 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 2.0001 | -276.0455 | 460.0296 | 368.0077 | 50 | -126.8738 | 0.015706 | 3.606e-06 | -1.9787e-09 | 5.6868e-08 | 2.8862e-09 | 74.7547 | -134.6377 | 0.016974 | 3.1285e-06 | 3.6338e-10 | 5.3073e-08 | 4.4012e-10 | 74.8028 | -97.5041 | 3.4028e+38 | 3.4028e+38 | -276.0456 | 460.0296 | 368.0077 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0023 | 2.0001 | -276.0444 | 460.0311 | 368.0099 | 49.9999 | -126.8736 | -0.60006 | 3.4028e+38 | 0 | 230.0003 | -276.0623 | 460.027 | 367.9864 | -0.00084061 | 0.017038 | 3.5884e-06 | -2.8861e-09 | 6.25e-08 | 3.5285e-09 | 74.7044 | -96.0256 | -0.0067421 | 3.4028e+38 | 12.1844 | 0.01996 | 0 | 0 | 0 | 6.5613e-05 | 0.019415 | 3.0259e-06 | 3.1133e-10 | 5.9374e-08 | 2.6755e-09 | 74.8031 | -91.8123 | -0.0013054 | 3.4028e+38 | -5.9142 | 229.0096 | -2.8548e-10 | 0.00070823 | 9.4234e-10 | -2.717e-05 | 3.4028e+38 | 3.4028e+38 | -276.0484 | 460.0237 | 368.0023 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -276.0598 | -276.0598 | 367.9979 | -0.0010803 | 230.0018 | 2.0001 | -276.0465 | 460.0301 | 368.0074 | 49.9999 | -126.8741 | [17 empty] | 21.3434 | 77.0796 | 121.2405 | 99.6214 | 100.0294 |  | -0.00066293
45 | 230.0013 | 5.0002 | -690.1072 | 1150.063 | 920.0121 | 50 | -126.8738 | 0.017313 | 3.8177e-06 | -3.0414e-09 | 6.6164e-08 | 1.1653e-08 | 74.7012 | -97.7622 | 0.018837 | 3.1605e-06 | -4.6674e-09 | 6.0893e-08 | 8.2294e-10 | 74.7143 | -165.0789 | 3.4028e+38 | 3.4028e+38 | -690.1072 | 1150.063 | 920.0121 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 5.0003 | -690.1061 | 1150.0675 | 920.0167 | 49.9999 | -126.8736 | 0.014934 | 3.8711e-06 | -2.4005e-09 | 5.8288e-08 | 7.0827e-09 | 74.6087 | -111.6553 | 0.016126 | 3.1093e-06 | -3.5366e-09 | 5.0723e-08 | -2.4933e-10 | 74.7975 | 168.8968 | 3.4028e+38 | 3.4028e+38 | -690.1062 | 1150.0675 | 920.0166 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 5.0003 | -690.104 | 1150.0677 | 920.0182 | 50 | -126.8737 | 0.014419 | 3.8787e-06 | -8.0175e-10 | 5.6255e-08 | 6.8278e-09 | 74.6875 | -95.2098 | 0.015613 | 3.1698e-06 | -8.1302e-10 | 4.901e-08 | -9.2024e-10 | 74.8179 | 131.4603 | 3.4028e+38 | 3.4028e+38 | -690.1041 | 1150.0677 | 920.0181 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 5.0002 | -690.1064 | 1150.0707 | 920.0204 | 50 | -126.8735 | 0.014465 | 3.8794e-06 | -8.9414e-10 | 5.6126e-08 | 5.548e-09 | 74.6819 | -112.6514 | 0.015671 | 3.1651e-06 | -1.6738e-09 | 4.9513e-08 | -2.5255e-10 | 74.7185 | 170.6158 | 3.4028e+38 | 3.4028e+38 | -690.1064 | 1150.0707 | 920.0203 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 5.0003 | -690.1082 | 1150.0713 | 920.0204 | 50 | -126.8736 | 0.014198 | 3.887e-06 | 1.5482e-09 | 5.5978e-08 | 5.631e-09 | 74.767 | -85.6383 | 0.01542 | 3.1115e-06 | -6.4833e-10 | 4.8776e-08 | -6.153e-11 | 74.7294 | 169.9489 | 3.4028e+38 | 3.4028e+38 | -690.1083 | 1150.0713 | 920.0203 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0022 | 5.0003 | -690.1045 | 1150.0668 | 920.0159 | 49.9999 | -126.8735 | -0.60006 | 3.4028e+38 | 0 | 230.0018 | -690.1249 | 1150.0673 | 919.9954 | 0.00031161 | 0.014554 | 3.8992e-06 | -9.3277e-10 | 5.6794e-08 | 6.2839e-09 | 74.71 | -95.7278 | -0.011115 | 3.4028e+38 | 23.4134 | 0.019635 | 0 | 0 | 0 | 0.00020517 | 0.016093 | 3.1745e-06 | -2.6201e-09 | 4.9692e-08 | -2.8233e-11 | 74.7461 | 179.5475 | -0.071935 | 3.4028e+38 | 6.1055 | 229.0975 | -1.1409e-09 | 0.0007171 | -2.0959e-10 | -3.8707e-05 | 3.4028e+38 | 3.4028e+38 | -690.103 | 1150.0671 | 920.0154 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -690.129 | -690.129 | 919.9846 | 0.00053807 | 230.0017 | 5.0003 | -690.1064 | 1150.068 | 920.0176 | 50 | -126.8737 | [17 empty] | 21.428 | 67.8528 | 125.7449 | 89.8518 | 83.0117 |  | -0.0012333
46 | 230.0015 | 5.0002 | -690.1053 | 1150.0635 | 920.014 | 49.9999 | -126.8737 | 0.017189 | 3.8028e-06 | -2.4107e-09 | 6.6091e-08 | 1.2059e-08 | 74.6825 | -97.9799 | 0.018811 | 3.2023e-06 | -4.3868e-09 | 5.9038e-08 | 2.0362e-09 | 74.7313 | -154.7625 | 3.4028e+38 | 3.4028e+38 | -690.1053 | 1150.0635 | 920.014 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 5.0003 | -690.1088 | 1150.0698 | 920.0187 | 49.9999 | -126.8735 | 0.014228 | 3.8017e-06 | -1.2635e-11 | 5.4079e-08 | 6.481e-09 | 74.7599 | -90.0515 | 0.015239 | 3.1882e-06 | -2.1387e-09 | 4.8606e-08 | 4.0711e-10 | 74.7671 | -170.9847 | 3.4028e+38 | 3.4028e+38 | -690.1089 | 1150.0698 | 920.0187 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 5.0002 | -690.0995 | 1150.0647 | 920.0187 | 49.9999 | -126.8733 | 0.015833 | 3.8343e-06 | 8.1172e-09 | 6.0966e-08 | 9.6953e-09 | 74.6255 | -50.9387 | 0.017076 | 3.1118e-06 | -1.225e-09 | 5.3469e-08 | 3.1184e-09 | 74.8104 | -111.4463 | 3.4028e+38 | 3.4028e+38 | -690.0996 | 1150.0647 | 920.0187 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 5.0003 | -690.1077 | 1150.0665 | 920.0156 | 50 | -126.8736 | 0.016118 | 3.8202e-06 | -2.0743e-09 | 6.1557e-08 | 9.6128e-09 | 74.7142 | -100.7774 | 0.017854 | 3.1271e-06 | -3.4783e-09 | 5.5847e-08 | -9.6365e-10 | 74.7304 | 166.4254 | 3.4028e+38 | 3.4028e+38 | -690.1077 | 1150.0665 | 920.0156 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 5.0003 | -690.1036 | 1150.0663 | 920.0179 | 50 | -126.8735 | 0.014312 | 3.881e-06 | 2.3622e-09 | 5.5101e-08 | 5.6967e-09 | 74.734 | -69.9257 | 0.015422 | 3.0805e-06 | -1.639e-09 | 4.8891e-08 | 4.6956e-10 | 74.7898 | -162.8562 | 3.4028e+38 | 3.4028e+38 | -690.1036 | 1150.0663 | 920.0179 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0018 | 5.0003 | -690.1062 | 1150.068 | 920.0187 | 50 | -126.8739 | -0.60006 | 3.4028e+38 | 0 | 230.0052 | -690.1295 | 1150.0811 | 920.0044 | -0.0013049 | 0.01478 | 3.88e-06 | 6.7144e-09 | 5.8675e-08 | 9.434e-09 | 74.7762 | -50.9081 | 0.13059 | 3.4028e+38 | -68.6501 | 0.020295 | 0 | 0 | 0 | 0.00019774 | 0.016439 | 3.1254e-06 | -1.1835e-09 | 5.1376e-08 | 3.2808e-09 | 74.6131 | -117.5441 | -0.032758 | 3.4028e+38 | 63.2223 | 229.157 | -9.6743e-10 | 0.00072232 | 1.0889e-09 | -4.1739e-05 | 3.4028e+38 | 3.4028e+38 | -690.1037 | 1150.0676 | 920.0128 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -690.144 | -690.144 | 920.0045 | -0.0020336 | 230.0019 | 5.0003 | -690.105 | 1150.0662 | 920.017 | 49.9999 | -126.8735 | [17 empty] | 22.8657 | 69.9219 | 128.8158 | 92.0623 | 85.5506 |  | -0.0011485
47 | 230.0016 | 5.0002 | -690.1024 | 1150.0619 | 920.0138 | 49.9999 | -126.8736 | 0.017122 | 3.8065e-06 | 7.3488e-09 | 6.7752e-08 | 1.3263e-08 | 74.7029 | -61.6588 | 0.018814 | 3.1572e-06 | -3.4693e-09 | 6.0747e-08 | 3.0497e-09 | 74.7593 | -136.3269 | 3.4028e+38 | 3.4028e+38 | -690.1024 | 1150.0619 | 920.0138 | 3.4028e+38 | 3.4028e+38 | 230.002 | 5.0002 | -690.1022 | 1150.0647 | 920.0143 | 49.9999 | -126.8736 | 0.016579 | 3.7986e-06 | 8.1879e-09 | 6.2725e-08 | 1.1992e-08 | 74.7742 | -54.864 | 0.01798 | 3.127e-06 | -1.8691e-09 | 5.5986e-08 | 3.2349e-09 | 74.7361 | -130.8869 | 3.4028e+38 | 3.4028e+38 | -690.1022 | 1150.0647 | 920.0143 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 5.0002 | -690.1031 | 1150.0643 | 920.0128 | 49.9999 | -126.8736 | 0.017461 | 3.7729e-06 | 7.2984e-09 | 6.5894e-08 | 1.5003e-08 | 74.79 | -65.4386 | 0.019139 | 3.2115e-06 | -1.9698e-09 | 6.0352e-08 | 3.8167e-09 | 74.7412 | -128.3383 | 3.4028e+38 | 3.4028e+38 | -690.1031 | 1150.0643 | 920.0128 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 5.0002 | -690.1061 | 1150.0658 | 920.0191 | 50 | -126.8735 | 0.014907 | 3.8673e-06 | 5.9272e-09 | 5.8068e-08 | 7.9724e-09 | 74.7851 | -58.4119 | 0.015974 | 3.2069e-06 | 7.0937e-11 | 5.1257e-08 | 1.4687e-09 | 74.728 | -135.6483 | 3.4028e+38 | 3.4028e+38 | -690.1061 | 1150.0658 | 920.0191 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 5.0003 | -690.1052 | 1150.0681 | 920.0176 | 50 | -126.8735 | 0.0158 | 3.8909e-06 | -2.1464e-09 | 5.9418e-08 | 7.6556e-09 | 74.7061 | -110.7586 | 0.017359 | 3.0684e-06 | -3.0359e-09 | 5.2946e-08 | -1.0798e-09 | 74.6795 | 168.4058 | 3.4028e+38 | 3.4028e+38 | -690.1053 | 1150.0681 | 920.0175 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0021 | 5.0003 | -690.1055 | 1150.0659 | 920.0149 | 49.9999 | -126.8736 | -0.60006 | 3.4028e+38 | 0 | 230.0013 | -690.126 | 1150.0697 | 919.9974 | 0.00054378 | 0.014331 | 3.9151e-06 | 1.1306e-09 | 5.5292e-08 | 6.879e-09 | 74.5624 | -80.6665 | 0.032363 | 3.4028e+38 | 51.7009 | 0.020012 | 0 | 0 | 0 | 0.00020581 | 0.015362 | 3.2092e-06 | -1.5677e-10 | 4.7432e-08 | 3.7319e-10 | 74.8205 | -130.9898 | -0.017244 | 3.4028e+38 | 53.1621 | 229.8034 | -3.9082e-10 | 0.00070711 | 5.908e-10 | -3.3014e-05 | 3.4028e+38 | 3.4028e+38 | -690.1074 | 1150.0679 | 920.0167 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -690.1266 | -690.1266 | 919.9998 | -0.00019137 | 230.0019 | 5.0002 | -690.1038 | 1150.065 | 920.0155 | 49.9999 | -126.8736 | [17 empty] | 21.7979 | 66.7057 | 123.8638 | 88.0697 | 81.1289 |  | -0.0011615
48 | 230.0016 | 0.50002 | 115.0064 | 115.0063 | 0.0082116 | 50 | -0.0042356 | 0.017732 | 3.571e-06 | 4.1876e-10 | 6.2564e-08 | -4.1087e-09 | 74.7117 | 99.2635 | 0.019543 | 3.0946e-06 | 1.8622e-09 | 6.223e-08 | -1.019e-09 | 74.8184 | 29.9004 | 3.4028e+38 | 3.4028e+38 | 115.0064 | 115.0063 | 0.0082116 | 3.4028e+38 | 3.4028e+38 | 230.002 | 0.50002 | 115.0062 | 115.0065 | 0.0083993 | 50 | -0.0041481 | 0.015781 | 3.5708e-06 | 1.1086e-09 | 5.6159e-08 | -2.3055e-09 | 74.7161 | 102.7729 | 0.017324 | 3.0806e-06 | 4.2296e-10 | 5.3352e-08 | -2.4904e-09 | 74.7832 | 98.7424 | 3.4028e+38 | 3.4028e+38 | 115.0062 | 115.0065 | 0.0083993 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 0.50002 | 115.0056 | 115.0056 | 0.0088322 | 49.9999 | -0.004146 | 0.014562 | 3.5242e-06 | 7.7829e-10 | 5.211e-08 | -5.2431e-10 | 74.7791 | 63.4251 | 0.01593 | 3.0841e-06 | 7.105e-10 | 4.9162e-08 | -5.8158e-10 | 74.7417 | 35.1104 | 3.4028e+38 | 3.4028e+38 | 115.0056 | 115.0056 | 0.0088322 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 0.50002 | 115.0056 | 115.0055 | 0.0081789 | 50 | -0.0040747 | 0.016603 | 3.5682e-06 | -4.3056e-10 | 5.9251e-08 | -2.5832e-09 | 74.7411 | 87.7298 | 0.018345 | 3.1678e-06 | 3.7887e-09 | 5.7887e-08 | 4.5422e-10 | 74.8182 | 5.6867 | 3.4028e+38 | 3.4028e+38 | 115.0056 | 115.0055 | 0.0081789 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 0.50002 | 115.0059 | 115.0058 | 0.008963 | 50 | -0.0043081 | 0.014987 | 3.5827e-06 | -1.4149e-09 | 5.4808e-08 | -2.3363e-09 | 74.8128 | 118.7622 | 0.016276 | 3.04e-06 | 4.6102e-10 | 5.1461e-08 | -1.5943e-09 | 74.7207 | 121.0033 | 3.4028e+38 | 3.4028e+38 | 115.0059 | 115.0058 | 0.008963 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0018 | 0.50002 | 115.006 | 115.0059 | 0.0087544 | 49.9999 | -0.0042214 | 1 | 3.4028e+38 | 0 | 230.0069 | 115.0088 | 115.0088 | 0.012217 | 0.00011293 | 0.015554 | 3.5571e-06 | -4.3918e-10 | 5.5628e-08 | -2.8193e-09 | 74.7698 | 116.823 | -0.026242 | 3.4028e+38 | -40.0771 | 0.020361 | 0 | 0 | 0 | -4.7585e-05 | 0.017069 | 3.1544e-06 | 2.6912e-10 | 5.3946e-08 | -1.8249e-09 | 74.6991 | 65.0568 | 0.015681 | 3.4028e+38 | 43.7263 | 229.4432 | -4.1499e-10 | 0.00071011 | 2.1211e-10 | 2.1731e-05 | 3.4028e+38 | 3.4028e+38 | 115.0061 | 115.006 | 0.0086611 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 115.0086 | 115.0086 | 0.012829 | 6.0859e-05 | 230.0019 | 0.50002 | 115.0059 | 115.0059 | 0.008517 | 50 | -0.0041825 | [17 empty] | 23.8126 | 11.7515 | 34.7248 | 34.7835 | 1619867.1368 |  | -0.0025629
49 | 230.0021 | 0.50002 | 115.0055 | 115.0055 | 0.0084569 | 49.9999 | -0.0042132 | 0.017581 | 3.5664e-06 | -5.9759e-10 | 6.2705e-08 | -5.5102e-09 | 74.7547 | 100.2723 | 0.019721 | 3.0617e-06 | 2.1872e-09 | 6.0371e-08 | -7.0456e-10 | 74.8534 | 16.4469 | 3.4028e+38 | 3.4028e+38 | 115.0055 | 115.0055 | 0.0084568 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 0.50002 | 115.0055 | 115.0053 | 0.0083091 | 49.9999 | -0.0041396 | 0.015708 | 3.5827e-06 | 1.3795e-09 | 5.4972e-08 | -2.2202e-09 | 74.6944 | 93.6623 | 0.016943 | 3.1728e-06 | 2.4897e-09 | 5.2098e-08 | -4.9358e-11 | 74.7505 | 1.1118 | 3.4028e+38 | 3.4028e+38 | 115.0055 | 115.0053 | 0.0083091 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 0.50002 | 115.0053 | 115.0052 | 0.0085407 | 50 | -0.004255 | 0.014424 | 3.5686e-06 | -6.3518e-10 | 5.1458e-08 | -1.0186e-09 | 74.6778 | 68.6554 | 0.01556 | 3.1623e-06 | 1.7268e-09 | 4.903e-08 | -8.2192e-10 | 74.8649 | 60.2617 | 3.4028e+38 | 3.4028e+38 | 115.0053 | 115.0052 | 0.0085407 | 3.4028e+38 | 3.4028e+38 | 230.0024 | 0.50002 | 115.0056 | 115.0056 | 0.0083759 | 49.9999 | -0.0044891 | 0.015049 | 3.6189e-06 | -8.2658e-11 | 5.3725e-08 | -1.6343e-09 | 74.7176 | 45.1157 | 0.016218 | 3.0832e-06 | 2.393e-09 | 5.0962e-08 | -3.4785e-10 | 74.8364 | 22.8965 | 3.4028e+38 | 3.4028e+38 | 115.0056 | 115.0056 | 0.0083759 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 0.50002 | 115.0057 | 115.0062 | 0.0078212 | 50.0001 | -0.0042908 | 0.01767 | 3.5375e-06 | -9.4866e-10 | 6.2444e-08 | -4.273e-09 | 74.5923 | 95.493 | 0.01928 | 3.0585e-06 | 1.2428e-09 | 5.8915e-08 | -8.8573e-10 | 74.7528 | 22.0024 | 3.4028e+38 | 3.4028e+38 | 115.0057 | 115.0062 | 0.0078212 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 0.50002 | 115.006 | 115.006 | 0.0086868 | 49.9999 | -0.0043278 | 1 | 3.4028e+38 | 0 | 230.004 | 115.0066 | 115.0064 | 0.013663 | 2.3293e-06 | 0.017822 | 3.6206e-06 | -7.0791e-10 | 6.3039e-08 | -4.6373e-09 | 74.6202 | 90.4969 | -0.00058064 | 3.4028e+38 | 9.5144 | 0.020715 | 0 | 0 | 0 | -5.3807e-05 | 0.019203 | 3.0508e-06 | 3.6402e-09 | 5.8615e-08 | -1.7449e-10 | 74.7767 | 2.8839 | 0.05789 | 3.4028e+38 | 7.5555 | 229.6666 | 8.1477e-10 | 0.00070169 | 5.2096e-10 | 1.9846e-05 | 3.4028e+38 | 3.4028e+38 | 115.0058 | 115.0058 | 0.0083157 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 115.0051 | 115.0044 | 0.011893 | -0.00017815 | 230.0019 | 0.50002 | 115.0055 | 115.0055 | 0.0083008 | 50 | -0.0042775 | [17 empty] | 23.2267 | 13.3282 | 35.9717 | 36.0306 | 1076137.8608 |  | -0.0022856
50 | 230.0016 | 0.50002 | 115.0052 | 115.0051 | 0.0079389 | 50 | -0.0039552 | 0.017783 | 3.5716e-06 | -6.0145e-10 | 6.3864e-08 | -4.2156e-09 | 74.6838 | 98.1633 | 0.019626 | 3.0508e-06 | 2.2407e-09 | 6.2222e-08 | -1.0446e-09 | 74.7974 | -16.4072 | 3.4028e+38 | 3.4028e+38 | 115.0052 | 115.0051 | 0.0079389 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 0.50002 | 115.0055 | 115.0054 | 0.0089224 | 49.9999 | -0.0043218 | 0.014279 | 3.5445e-06 | 1.2232e-09 | 5.0578e-08 | -1.1432e-09 | 74.7879 | 113.8062 | 0.015442 | 3.0508e-06 | 1.1571e-09 | 4.7506e-08 | -1.499e-09 | 74.6987 | 9.4852 | 3.4028e+38 | 3.4028e+38 | 115.0055 | 115.0054 | 0.0089224 | 3.4028e+38 | 3.4028e+38 | 230.002 | 0.50002 | 115.0057 | 115.0057 | 0.0081357 | 49.9999 | -0.0041653 | 0.017616 | 3.5308e-06 | -5.0695e-10 | 6.3271e-08 | -4.3657e-09 | 74.7151 | 109.1043 | 0.019752 | 3.1956e-06 | 1.7323e-09 | 6.3025e-08 | -9.5499e-10 | 74.775 | 33.3963 | 3.4028e+38 | 3.4028e+38 | 115.0057 | 115.0057 | 0.0081357 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 0.50002 | 115.0059 | 115.0059 | 0.0084304 | 49.9999 | -0.0041227 | 0.016468 | 3.5589e-06 | -4.8015e-10 | 5.9006e-08 | -3.6247e-09 | 74.7229 | 110.8764 | 0.018558 | 3.1797e-06 | 7.5851e-10 | 5.7977e-08 | -1.4973e-09 | 74.763 | 63.3821 | 3.4028e+38 | 3.4028e+38 | 115.0059 | 115.0059 | 0.0084304 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 0.50002 | 115.0054 | 115.0051 | 0.0085668 | 50 | -0.0041658 | 0.016428 | 3.5667e-06 | -2.0635e-10 | 5.7521e-08 | -2.3641e-09 | 74.7794 | 86.5627 | 0.017877 | 3.0916e-06 | 2.616e-09 | 5.6616e-08 | 3.9126e-10 | 74.7615 | -0.14703 | 3.4028e+38 | 3.4028e+38 | 115.0054 | 115.0051 | 0.0085668 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 0.50002 | 115.0054 | 115.0053 | 0.0082173 | 50 | -0.0040939 | 1 | 3.4028e+38 | 0 | 230.0005 | 115.0045 | 115.0044 | 0.011195 | -0.00021013 | 0.016748 | 3.5926e-06 | 6.0139e-10 | 5.8367e-08 | -3.2772e-09 | 74.7578 | 79.6017 | 0.0020964 | 3.4028e+38 | -5.0145 | 0.020189 | 0 | 0 | 0 | -5.1361e-05 | 0.018637 | 3.2048e-06 | 2.448e-09 | 5.5734e-08 | 3.4679e-10 | 74.7472 | -7.3342 | 0.048343 | 3.4028e+38 | -35.0019 | 229.6804 | 1.185e-09 | 0.00070551 | 1.139e-11 | 2.0127e-05 | 3.4028e+38 | 3.4028e+38 | 115.0052 | 115.0051 | 0.0082886 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 115.005 | 115.005 | 0.0119 | -0.00021805 | 230.0018 | 0.50002 | 115.0055 | 115.0054 | 0.0083988 | 49.9999 | -0.0041462 | [17 empty] | 22.2004 | 15.113 | 36.7273 | 35.7767 | 1376742.5835 |  | -0.0023856
51 | 230.0019 | 0.50002 | 57.4946 | 115.0051 | 99.6022 | 49.9999 | -60.0047 | 0.017662 | 3.5575e-06 | 3.2456e-09 | 6.3254e-08 | -1.1112e-09 | 74.8133 | 46.3112 | 0.019383 | 3.1484e-06 | 4.629e-09 | 6.0726e-08 | 1.0629e-09 | 74.6632 | -13.3771 | 3.4028e+38 | 3.4028e+38 | 57.4946 | 115.0051 | 99.6022 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 0.50002 | 57.4954 | 115.0052 | 99.6019 | 50 | -60.0042 | 0.01781 | 3.5554e-06 | 2.5243e-09 | 6.3586e-08 | -1.1358e-09 | 74.7679 | 19.3909 | 0.01942 | 3.1101e-06 | 4.5381e-09 | 6.2489e-08 | 1.9656e-09 | 74.7192 | -22.7416 | 3.4028e+38 | 3.4028e+38 | 57.4954 | 115.0052 | 99.6019 | 3.4028e+38 | 3.4028e+38 | 230.0012 | 0.50002 | 57.4953 | 115.0054 | 99.602 | 50 | -60.0044 | 0.017007 | 3.4997e-06 | 2.8276e-09 | 5.9444e-08 | -1.3085e-09 | 74.7588 | 29.009 | 0.018824 | 3.0744e-06 | 2.7673e-09 | 5.7827e-08 | -6.588e-10 | 74.7565 | 12.8042 | 3.4028e+38 | 3.4028e+38 | 57.4953 | 115.0054 | 99.602 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 0.50002 | 57.4949 | 115.0055 | 99.6024 | 49.9999 | -60.0045 | 0.01573 | 3.5214e-06 | 3.1222e-09 | 5.4552e-08 | 5.7589e-10 | 74.6542 | -9.3902 | 0.016912 | 3.1272e-06 | 4.2643e-09 | 5.2561e-08 | 1.7044e-09 | 74.6687 | -21.7858 | 3.4028e+38 | 3.4028e+38 | 57.4949 | 115.0055 | 99.6024 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 0.50002 | 57.4956 | 115.0053 | 99.6018 | 49.9999 | -60.0044 | 0.017039 | 3.5617e-06 | 2.4496e-09 | 6.0306e-08 | -1.2413e-09 | 74.7254 | 37.0719 | 0.018901 | 3.1988e-06 | 3.9703e-09 | 6.0328e-08 | -1.0177e-10 | 74.8506 | 1.5658 | 3.4028e+38 | 3.4028e+38 | 57.4956 | 115.0053 | 99.6018 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0018 | 0.50002 | 57.4956 | 115.0055 | 99.6021 | 49.9999 | -60.0041 | 0.49994 | 3.4028e+38 | 0 | 230.0064 | 57.4948 | 115.0083 | 99.6049 | -9.1733e-05 | 0.014932 | 3.5575e-06 | 1.9158e-09 | 5.3138e-08 | -9.0966e-10 | 74.8322 | 14.1336 | 0.041074 | 3.4028e+38 | -13.9816 | 0.019598 | 0 | 0 | 0 | -9.5e-06 | 0.015235 | 3.0534e-06 | 2.2016e-09 | 4.6543e-08 | 3.5156e-10 | 74.6733 | -8.6958 | 0.048078 | 3.4028e+38 | 42.0251 | 229.5875 | -4.9333e-10 | 0.00070672 | 7.8595e-11 | 1.71e-05 | 3.4028e+38 | 3.4028e+38 | 57.4951 | 115.0055 | 99.6023 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 57.4923 | 57.4923 | 99.6017 | -0.00020018 | 230.0016 | 0.50002 | 57.4952 | 115.0053 | 99.6021 | 49.9999 | -60.0044 | [17 empty] | 21.5984 | 48.9607 | -74.3873 | 72.1115 | 123.2535 |  | -0.0049591
52 | 230.002 | 0.50002 | 57.495 | 115.0053 | 99.602 | 49.9999 | -60.0047 | 0.015282 | 3.5525e-06 | 2.5017e-09 | 5.4242e-08 | 5.5076e-10 | 74.8071 | -19.7252 | 0.016682 | 3.053e-06 | 2.6849e-09 | 5.0663e-08 | 1.1907e-09 | 74.7599 | -16.8926 | 3.4028e+38 | 3.4028e+38 | 57.495 | 115.0053 | 99.602 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 0.50002 | 57.4953 | 115.0053 | 99.602 | 49.9999 | -60.0042 | 0.017047 | 3.5804e-06 | 2.4273e-09 | 6.1035e-08 | -1.3667e-09 | 74.6589 | 36.938 | 0.018437 | 3.1892e-06 | 4.4133e-09 | 5.6409e-08 | -3.5509e-10 | 74.6404 | 7.768 | 3.4028e+38 | 3.4028e+38 | 57.4953 | 115.0053 | 99.602 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 0.50002 | 57.4955 | 115.0056 | 99.6029 | 49.9999 | -60.0046 | 0.014258 | 3.4956e-06 | 1.1292e-09 | 5.0423e-08 | -3.6251e-10 | 74.7705 | -14.5215 | 0.015113 | 3.1759e-06 | 8.4963e-10 | 4.9085e-08 | -6.7439e-10 | 74.7394 | 10.2502 | 3.4028e+38 | 3.4028e+38 | 57.4955 | 115.0056 | 99.6029 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 0.50002 | 57.4953 | 115.0055 | 99.6019 | 50 | -60.0044 | 0.015335 | 3.5788e-06 | 1.0018e-09 | 5.3781e-08 | -1.458e-09 | 74.8065 | 50.2363 | 0.016782 | 3.0984e-06 | 9.9734e-10 | 5.1709e-08 | -7.584e-10 | 74.7536 | 36.137 | 3.4028e+38 | 3.4028e+38 | 57.4953 | 115.0055 | 99.6019 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 0.50002 | 57.4951 | 115.0051 | 99.6019 | 49.9999 | -60.0044 | 0.016383 | 3.6005e-06 | 2.2997e-09 | 5.8935e-08 | 6.8231e-10 | 74.7459 | 1.8855 | 0.017855 | 3.1355e-06 | 3.8727e-09 | 5.5789e-08 | 1.623e-09 | 74.7199 | -34.236 | 3.4028e+38 | 3.4028e+38 | 57.4951 | 115.0051 | 99.6019 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0017 | 0.50002 | 57.4954 | 115.0053 | 99.6018 | 50 | -60.0045 | 0.49993 | 3.4028e+38 | 0 | 230.0004 | 57.4908 | 115.0043 | 99.601 | -0.00034377 | 0.016425 | 3.5708e-06 | 2.8492e-09 | 5.8604e-08 | 5.227e-10 | 74.7331 | -6.7953 | 0.074678 | 3.4028e+38 | -31.988 | 0.020925 | 0 | 0 | 0 | -1.9834e-05 | 0.018144 | 3.1035e-06 | 3.7986e-09 | 5.6873e-08 | 2.8162e-09 | 74.751 | -27.1977 | 0.09198 | 3.4028e+38 | -47.5952 | 229.2898 | -7.0664e-11 | 0.00073502 | 8.0788e-10 | 3.956e-06 | 3.4028e+38 | 3.4028e+38 | 57.495 | 115.0053 | 99.602 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 57.4904 | 57.4904 | 99.6021 | -0.00039271 | 230.0019 | 0.50002 | 57.4953 | 115.0053 | 99.6021 | 49.9999 | -60.0044 | [17 empty] | 21.9571 | 40.9019 | -91.9856 | 61.146 | 115.0591 |  | -0.0052299
53 | 230.002 | 0.50002 | 57.495 | 115.0059 | 99.6028 | 50 | -60.0046 | 0.01427 | 3.5511e-06 | 8.154e-10 | 5.1152e-08 | 6.28e-11 | 74.7586 | -2.6084 | 0.015275 | 3.2268e-06 | 5.9546e-10 | 4.9277e-08 | -3.0679e-11 | 74.7177 | 1.5349 | 3.4028e+38 | 3.4028e+38 | 57.495 | 115.0059 | 99.6028 | 3.4028e+38 | 3.4028e+38 | 230.002 | 0.50002 | 57.4949 | 115.0054 | 99.6021 | 49.9999 | -60.0044 | 0.016047 | 3.537e-06 | 1.6806e-09 | 5.7793e-08 | 9.3155e-10 | 74.8297 | -13.7339 | 0.017833 | 3.1377e-06 | 4.0109e-09 | 5.6595e-08 | 1.8793e-09 | 74.7692 | -27.727 | 3.4028e+38 | 3.4028e+38 | 57.4949 | 115.0054 | 99.6021 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 0.50002 | 57.4949 | 115.0057 | 99.6026 | 50 | -60.0044 | 0.014426 | 3.5363e-06 | 1.3229e-09 | 5.0993e-08 | 3.6781e-10 | 74.7048 | -8.1192 | 0.015246 | 3.0716e-06 | 3.1276e-10 | 4.8734e-08 | -5.7935e-10 | 74.6769 | 34.3152 | 3.4028e+38 | 3.4028e+38 | 57.4949 | 115.0057 | 99.6026 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 0.50002 | 57.4953 | 115.0057 | 99.6027 | 49.9999 | -60.0045 | 0.014362 | 3.5714e-06 | 6.1289e-10 | 5.0346e-08 | -3.4011e-10 | 74.5951 | 75.5723 | 0.014975 | 3.1919e-06 | 1.5377e-09 | 4.6668e-08 | -2.3848e-10 | 74.7709 | 7.8499 | 3.4028e+38 | 3.4028e+38 | 57.4953 | 115.0057 | 99.6026 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 0.50002 | 57.4952 | 115.0054 | 99.6023 | 50 | -60.0042 | 0.014486 | 3.6344e-06 | 4.6071e-10 | 5.2666e-08 | -3.9685e-10 | 74.7633 | 2.8379 | 0.015179 | 3.1919e-06 | 1.4422e-09 | 4.84e-08 | 2.3568e-11 | 74.6124 | -1.2832 | 3.4028e+38 | 3.4028e+38 | 57.4952 | 115.0054 | 99.6023 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0018 | 0.50002 | 57.4952 | 115.0059 | 99.6025 | 49.9999 | -60.0046 | 0.49993 | 3.4028e+38 | 0 | 230.0014 | 57.493 | 115.0055 | 99.6025 | -0.00014207 | 0.014598 | 3.5145e-06 | 1.0193e-09 | 5.1283e-08 | 1.0704e-09 | 74.7767 | -57.6326 | 0.013211 | 3.4028e+38 | -20.7985 | 0.020088 | 0 | 0 | 0 | -1.4064e-05 | 0.01658 | 3.1493e-06 | 2.8612e-09 | 5.222e-08 | 1.8458e-09 | 74.7505 | -24.4835 | 0.060259 | 3.4028e+38 | 22.2934 | 229.0077 | 4.2219e-10 | 0.00071986 | -4.3141e-10 | 9.0289e-06 | 3.4028e+38 | 3.4028e+38 | 57.4945 | 115.0048 | 99.6016 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 57.4904 | 57.4903 | 99.6018 | -0.00048144 | 230.0019 | 0.50002 | 57.4951 | 115.0056 | 99.6025 | 49.9999 | -60.0044 | [17 empty] | 22.0407 | 42.7584 | -89.4746 | 64.6268 | 118.0975 |  | -0.005065
54 | 230.0017 | 0.50002 | 91.9739 | 115.0053 | -69.0457 | 49.9999 | 36.8961 | 0.016981 | 3.6243e-06 | -2.694e-09 | 6.2326e-08 | -3.8911e-09 | 74.7913 | 108.2183 | 0.018707 | 3.1205e-06 | -3.1277e-10 | 5.8335e-08 | -9.0402e-10 | 74.7536 | 99.7222 | 3.4028e+38 | 3.4028e+38 | 91.9739 | 115.0053 | -69.0457 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 0.50002 | 91.9731 | 115.0051 | -69.0456 | 50 | 36.896 | 0.014092 | 3.5452e-06 | -4.8679e-11 | 5.0539e-08 | -1.1009e-09 | 74.7591 | 147.8477 | 0.015472 | 3.112e-06 | 2.4075e-10 | 4.75e-08 | -1.215e-09 | 74.8069 | 136.6975 | 3.4028e+38 | 3.4028e+38 | 91.9731 | 115.0051 | -69.0456 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 0.50002 | 91.9731 | 115.0052 | -69.0459 | 50 | 36.8963 | 0.014742 | 3.5388e-06 | -5.2333e-10 | 5.2823e-08 | -1.5562e-09 | 74.6813 | 97.9484 | 0.016035 | 3.0779e-06 | 9.8978e-10 | 5.0708e-08 | -5.8598e-10 | 74.7313 | 42.7634 | 3.4028e+38 | 3.4028e+38 | 91.9731 | 115.0052 | -69.046 | 3.4028e+38 | 3.4028e+38 | 230.002 | 0.50002 | 91.9731 | 115.0054 | -69.0459 | 49.9999 | 36.8962 | 0.014443 | 3.5372e-06 | -1.752e-09 | 5.229e-08 | -1.6267e-09 | 74.7387 | 119.6596 | 0.015772 | 3.0786e-06 | -7.8232e-11 | 4.8839e-08 | -4.5589e-10 | 74.7503 | 104.6586 | 3.4028e+38 | 3.4028e+38 | 91.9731 | 115.0054 | -69.0459 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 0.50002 | 91.9733 | 115.0052 | -69.0455 | 49.9999 | 36.896 | 0.016305 | 3.6023e-06 | -3.5808e-09 | 5.8774e-08 | -3.2166e-09 | 74.7706 | 129.3759 | 0.017676 | 3.1797e-06 | -1.3124e-09 | 5.4658e-08 | -1.4036e-09 | 74.676 | 111.2585 | 3.4028e+38 | 3.4028e+38 | 91.9733 | 115.0052 | -69.0456 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 0.50002 | 91.9738 | 0 | 0 | 49.9999 | 36.8961 | 0.79973 | 3.4028e+38 | 0 | 230.0068 | 91.9778 | 115.0081 | -69.0423 | 0.00034421 | 0.016725 | 3.6038e-06 | -3.5418e-09 | 6.1095e-08 | -3.6989e-09 | 74.7725 | 121.4963 | -0.037707 | 3.4028e+38 | 31.06 | 0.020263 | 0 | 0 | 0 | -4.9375e-05 | 0.018607 | 3.1589e-06 | -8.8449e-10 | 6.0491e-08 | -8.4114e-10 | 74.7787 | 56.1241 | 0.0051164 | 3.4028e+38 | 27.7216 | 229.3488 | 1.3067e-10 | 0.00072114 | -6.8318e-11 | 5.6595e-06 | 3.4028e+38 | 3.4028e+38 | 91.9729 | 115.0054 | -69.0464 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 91.9774 | 91.9774 | -69.0411 | 0.00022889 | 230.0019 | 0.50002 | 91.9733 | 115.0052 | -69.0457 | 49.9999 | 36.8961 | [17 empty] | 22.2905 | 51.6716 | 85.7767 | 72.3006 | 68.8755 |  | -0.00045496
55 | 230.0016 | 0.50002 | 91.9735 | 115.0057 | -69.0461 | 49.9999 | 36.8961 | 0.015693 | 3.5768e-06 | -1.9182e-09 | 5.4355e-08 | -1.5813e-09 | 74.8407 | 122.921 | 0.016861 | 3.1335e-06 | -3.7041e-10 | 5.0979e-08 | -1.0579e-09 | 74.79 | 80.4337 | 3.4028e+38 | 3.4028e+38 | 91.9735 | 115.0057 | -69.0461 | 3.4028e+38 | 3.4028e+38 | 230.002 | 0.50002 | 91.9732 | 115.0055 | -69.0461 | 50 | 36.8962 | 0.014727 | 3.5634e-06 | -2.9317e-09 | 5.2521e-08 | -2.5236e-09 | 74.6547 | 106.6715 | 0.016102 | 3.1753e-06 | 1.9669e-09 | 5.1038e-08 | -4.0329e-10 | 74.7978 | 11.5973 | 3.4028e+38 | 3.4028e+38 | 91.9732 | 115.0055 | -69.0461 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 0.50002 | 91.9734 | 115.0056 | -69.0459 | 49.9999 | 36.8965 | 0.017758 | 3.577e-06 | -4.6394e-09 | 6.2487e-08 | -5.0181e-09 | 74.7432 | 123.1543 | 0.019393 | 3.087e-06 | 1.4671e-09 | 6.2401e-08 | -8.933e-11 | 74.7407 | 2.5648 | 3.4028e+38 | 3.4028e+38 | 91.9734 | 115.0056 | -69.0459 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 0.50002 | 91.9734 | 115.0056 | -69.0456 | 49.9999 | 36.8959 | 0.016322 | 3.5785e-06 | -3.0349e-09 | 5.8379e-08 | -3.3535e-09 | 74.7634 | 131.1839 | 0.017683 | 3.0448e-06 | -9.7512e-10 | 5.444e-08 | -1.3583e-09 | 74.7149 | 50.1582 | 3.4028e+38 | 3.4028e+38 | 91.9734 | 115.0056 | -69.0456 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 0.50002 | 91.9736 | 115.0057 | -69.0461 | 49.9999 | 36.8961 | 0.01466 | 3.5984e-06 | -8.981e-10 | 5.3025e-08 | -1.2564e-09 | 74.8111 | 143.8197 | 0.016024 | 3.1402e-06 | -1.1213e-09 | 5.0365e-08 | -1.0012e-09 | 74.741 | 134.0152 | 3.4028e+38 | 3.4028e+38 | 91.9736 | 115.0057 | -69.0461 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0021 | 0.50002 | 91.9738 | 115.0052 | -69.0452 | 49.9999 | 36.8958 | 0.79973 | 3.4028e+38 | 0 | 230.0056 | 91.9762 | 115.0074 | -69.044 | 9.3845e-05 | 0.015124 | 3.567e-06 | -1.6952e-09 | 5.2789e-08 | -2.1316e-09 | 74.8723 | 136.1245 | -0.037674 | 3.4028e+38 | 51.929 | 0.020221 | 0 | 0 | 0 | -3.9629e-05 | 0.016329 | 3.1503e-06 | -6.6898e-10 | 5.1256e-08 | -1.601e-09 | 74.762 | 141.9648 | -0.021645 | 3.4028e+38 | -36.9003 | 229.5382 | 1.7236e-09 | 0.00072766 | -1.0248e-10 | 1.9783e-05 | 3.4028e+38 | 3.4028e+38 | 91.9735 | 115.0057 | -69.0457 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 91.9773 | 91.9773 | -69.0432 | 0.00025573 | 230.0018 | 0.50002 | 91.9734 | 115.0056 | -69.046 | 49.9999 | 36.8962 | [17 empty] | 22.0455 | 45.5375 | 81.192 | 68.3887 | 62.5707 |  | -0.00049275
56 | 230.0021 | 0.50002 | 91.9734 | 115.0054 | -69.0461 | 49.9999 | 36.8963 | 0.01661 | 3.5373e-06 | -2.8934e-09 | 5.9616e-08 | -4.4793e-09 | 74.7188 | 107.0405 | 0.018458 | 3.0996e-06 | 1.4317e-09 | 5.8462e-08 | 1.897e-10 | 74.756 | -5.9921 | 3.4028e+38 | 3.4028e+38 | 91.9734 | 115.0054 | -69.0461 | 3.4028e+38 | 3.4028e+38 | 230.002 | 0.50002 | 91.9735 | 115.0056 | -69.0459 | 49.9999 | 36.896 | 0.014387 | 3.5844e-06 | -5.8607e-10 | 5.0951e-08 | -1.0328e-09 | 74.8154 | 145.2599 | 0.015193 | 3.0844e-06 | 3.1142e-10 | 4.6876e-08 | -1.3061e-09 | 74.7658 | 70.4503 | 3.4028e+38 | 3.4028e+38 | 91.9735 | 115.0056 | -69.046 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 0.50002 | 91.973 | 115.0057 | -69.0462 | 50 | 36.8963 | 0.017536 | 3.5012e-06 | -2.7427e-09 | 6.3632e-08 | -5.3703e-09 | 74.7423 | 111.8181 | 0.019488 | 3.1378e-06 | 1.3592e-09 | 6.0737e-08 | -3.1449e-10 | 74.6159 | 14.7052 | 3.4028e+38 | 3.4028e+38 | 91.973 | 115.0057 | -69.0462 | 3.4028e+38 | 3.4028e+38 | 230.002 | 0.50002 | 91.9736 | 115.0059 | -69.0461 | 49.9998 | 36.8962 | 0.016596 | 3.5961e-06 | -1.7948e-09 | 6.0809e-08 | -3.5069e-09 | 74.7936 | 134.0003 | 0.018364 | 3.0681e-06 | -1.488e-09 | 5.8053e-08 | -1.1751e-09 | 74.7503 | 141.7009 | 3.4028e+38 | 3.4028e+38 | 91.9736 | 115.0059 | -69.0461 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 0.50002 | 91.9728 | 115.0057 | -69.0462 | 50 | 36.8963 | 0.016555 | 3.5522e-06 | -8.3652e-10 | 5.8803e-08 | -4.7661e-09 | 74.7325 | 107.9366 | 0.01794 | 3.0622e-06 | 1.8912e-09 | 5.4925e-08 | -5.7659e-10 | 74.6772 | 17.012 | 3.4028e+38 | 3.4028e+38 | 91.9728 | 115.0057 | -69.0462 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0017 | 0.50002 | 91.9721 | 115.0047 | -69.0462 | 50 | 36.8965 | 0.79973 | 3.4028e+38 | 0 | 230.0008 | 91.975 | 115.0049 | -69.0421 | 1.9389e-05 | 0.016569 | 3.5552e-06 | -1.5768e-09 | 5.8879e-08 | -3.7918e-09 | 74.7008 | 121.7993 | -0.039533 | 3.4028e+38 | 23.036 | 0.020975 | 0 | 0 | 0 | -4.691e-05 | 0.01828 | 3.1042e-06 | 1.6876e-09 | 5.8548e-08 | -5.3261e-10 | 74.7639 | 17.5156 | 0.027338 | 3.4028e+38 | -46.3941 | 229.1299 | -2.9692e-10 | 0.00072293 | -5.2801e-11 | 1.7033e-05 | 3.4028e+38 | 3.4028e+38 | 91.9725 | 115.0051 | -69.046 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 91.9756 | 91.9754 | -69.0415 | 0.00010129 | 230.0019 | 0.50002 | 91.9733 | 115.0057 | -69.0461 | 49.9999 | 36.8962 | [17 empty] | 21.0924 | 50.3874 | 86.5392 | 72.9136 | 63.8037 |  | -0.00067237
57 | 230.0015 | 1 | 230.0095 | 230.0096 | 0.014644 | 49.9999 | -0.003648 | 0.01732 | 3.6539e-06 | 2.7644e-11 | 6.4133e-08 | -8.2569e-09 | 74.7201 | 88.0302 | 0.019298 | 3.1229e-06 | 2.177e-09 | 6.0486e-08 | -2.7046e-09 | 74.6184 | 32.1493 | 3.4028e+38 | 3.4028e+38 | 230.0095 | 230.0096 | 0.014644 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 1 | 230.0103 | 230.0096 | 0.014954 | 49.9999 | -0.0037251 | 0.014352 | 3.6294e-06 | -4.3308e-10 | 5.2089e-08 | -2.8442e-09 | 74.7591 | 109.3776 | 0.015326 | 3.0755e-06 | 3.2556e-10 | 4.7156e-08 | -1.7883e-09 | 74.7657 | 49.1986 | 3.4028e+38 | 3.4028e+38 | 230.0103 | 230.0096 | 0.014954 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 1 | 230.0098 | 230.0096 | 0.014204 | 49.9999 | -0.0037708 | 0.017609 | 3.6877e-06 | 2.9183e-10 | 6.4993e-08 | -8.3455e-09 | 74.7725 | 97.5754 | 0.019761 | 3.1943e-06 | 1.9337e-09 | 6.0908e-08 | -2.3626e-09 | 74.783 | 42.5338 | 3.4028e+38 | 3.4028e+38 | 230.0098 | 230.0096 | 0.014204 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 1 | 230.0101 | 230.0094 | 0.01572 | 49.9999 | -0.0039158 | 0.014231 | 3.5933e-06 | 3.3705e-10 | 5.2331e-08 | -2.0739e-09 | 74.8332 | 117.5555 | 0.015293 | 3.139e-06 | 8.3616e-11 | 4.9416e-08 | -1.3651e-09 | 74.7718 | 54.1826 | 3.4028e+38 | 3.4028e+38 | 230.0101 | 230.0094 | 0.01572 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | 230.0105 | 230.0103 | 0.015532 | 50 | -0.0038459 | 0.014598 | 3.5876e-06 | -2.1669e-09 | 5.2383e-08 | -2.7734e-09 | 74.76 | 104.4179 | 0.015707 | 3.0936e-06 | 1.9878e-10 | 4.8549e-08 | -2.0911e-09 | 74.806 | 126.1139 | 3.4028e+38 | 3.4028e+38 | 230.0104 | 230.0103 | 0.015532 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0022 | 1 | 230.01 | 230.0098 | 0.015725 | 49.9999 | -0.0039063 | 1 | 3.4028e+38 | 0 | 230.0039 | 230.0122 | 230.012 | 0.019587 | -0.00016123 | 0.014112 | 3.6364e-06 | -7.1635e-10 | 5.1856e-08 | -2.5787e-09 | 74.7016 | 96.1174 | 0.011166 | 3.4028e+38 | -51.6411 | 0.020049 | 0 | 0 | 0 | -9.1305e-05 | 0.015368 | 3.1875e-06 | 2.2212e-09 | 4.8989e-08 | -1.3695e-09 | 74.661 | 101.4378 | -0.0064596 | 3.4028e+38 | -70.8932 | 229.2025 | -3.5126e-10 | 0.00072818 | 3.7703e-10 | 5.9757e-05 | 3.4028e+38 | 3.4028e+38 | 230.01 | 230.0099 | 0.014767 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 230.0101 | 230.0101 | 0.018437 | -0.0002491 | 230.0018 | 1 | 230.01 | 230.0097 | 0.015011 | 49.9999 | -0.0037811 | [17 empty] | 21.2413 | 12.2819 | 33.845 | 32.4172 | 1783513.6067 |  | -0.0024377
58 | 230.0019 | 1 | 230.01 | 230.0093 | 0.014492 | 50 | -0.0036099 | 0.017535 | 3.6208e-06 | 1.8504e-11 | 6.3469e-08 | -6.3632e-09 | 74.7573 | 78.5307 | 0.019306 | 3.1776e-06 | 3.3407e-09 | 6.1444e-08 | -1.7709e-09 | 74.699 | 21.4309 | 3.4028e+38 | 3.4028e+38 | 230.01 | 230.0093 | 0.014492 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 1 | 230.0094 | 230.0093 | 0.015096 | 49.9999 | -0.0037637 | 0.017235 | 3.6974e-06 | 1.4694e-09 | 6.1258e-08 | -5.3358e-09 | 74.7285 | 89.1885 | 0.018659 | 3.0936e-06 | 2.7565e-09 | 5.968e-08 | -1.4651e-09 | 74.7173 | 29.0584 | 3.4028e+38 | 3.4028e+38 | 230.0094 | 230.0093 | 0.015096 | 3.4028e+38 | 3.4028e+38 | 230.002 | 1 | 230.0098 | 230.0097 | 0.015359 | 49.9999 | -0.0034668 | 0.017523 | 3.6801e-06 | -3.9724e-10 | 6.455e-08 | -8.3866e-09 | 74.7191 | 90.9877 | 0.019463 | 3.1093e-06 | 7.9769e-10 | 6.0579e-08 | -3.4994e-09 | 74.6985 | 88.713 | 3.4028e+38 | 3.4028e+38 | 230.0098 | 230.0097 | 0.015359 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | 230.0099 | 230.0097 | 0.014724 | 50 | -0.0036678 | 0.01534 | 3.5608e-06 | 1.146e-09 | 5.4127e-08 | -2.7087e-09 | 74.7056 | 63.3229 | 0.016618 | 3.1312e-06 | 2.5315e-09 | 5.1189e-08 | -1.0884e-09 | 74.7914 | 8.4568 | 3.4028e+38 | 3.4028e+38 | 230.0099 | 230.0097 | 0.014724 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 1 | 230.0099 | 230.0098 | 0.01432 | 50 | -0.0036535 | 0.017637 | 3.6275e-06 | 9.651e-10 | 6.3987e-08 | -6.2889e-09 | 74.7291 | 82.8209 | 0.019255 | 3.1814e-06 | 3.0232e-09 | 6.1016e-08 | -2.7364e-09 | 74.7426 | 47.9265 | 3.4028e+38 | 3.4028e+38 | 230.0099 | 230.0098 | 0.01432 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0015 | 1 | 230.0094 | 230.0092 | 0.014684 | 50 | -0.0036578 | 1 | 3.4028e+38 | 0 | 230.0022 | 230.0102 | 230.0101 | 0.023445 | 8.8681e-05 | 0.017421 | 3.6452e-06 | -5.9635e-10 | 6.3888e-08 | -7.7412e-09 | 74.7723 | 100.0959 | 0.0030845 | 3.4028e+38 | -43.2421 | 0.020954 | 0 | 0 | 0 | -0.00010984 | 0.019527 | 3.066e-06 | 1.7057e-09 | 6.03e-08 | -3.0911e-09 | 74.6785 | 68.7905 | 0.017216 | 3.4028e+38 | 34.7344 | 229.7527 | -1.4895e-09 | 0.00071895 | -1.5398e-09 | 4.1016e-05 | 3.4028e+38 | 3.4028e+38 | 230.01 | 230.0094 | 0.014392 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 230.0126 | 230.0126 | 0.025046 | 0.00012509 | 230.0017 | 1 | 230.0098 | 230.0095 | 0.014798 | 49.9999 | -0.0036323 | [17 empty] | 21.6423 | 10.2602 | 32.2706 | 31.0553 | 2028857.8558 |  | -0.0024153
59 | 230.0016 | 1 | 230.0098 | 230.0097 | 0.014478 | 50 | -0.0036064 | 0.017685 | 3.6111e-06 | -1.0041e-09 | 6.3394e-08 | -7.1084e-09 | 74.6986 | 99.5425 | 0.019486 | 3.0524e-06 | 1.4234e-09 | 6.3313e-08 | -2.8038e-09 | 74.8085 | 54.8819 | 3.4028e+38 | 3.4028e+38 | 230.0098 | 230.0097 | 0.014478 | 3.4028e+38 | 3.4028e+38 | 230.0013 | 1 | 230.0094 | 230.0093 | 0.01499 | 49.9999 | -0.0036613 | 0.017651 | 3.6184e-06 | -1.6786e-09 | 6.3875e-08 | -8.5083e-09 | 74.7483 | 104.3626 | 0.019455 | 3.1043e-06 | 1.108e-09 | 6.0365e-08 | -3.3122e-09 | 74.788 | 61.3348 | 3.4028e+38 | 3.4028e+38 | 230.0094 | 230.0093 | 0.01499 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 1 | 230.0103 | 230.0099 | 0.014643 | 49.9999 | -0.0036477 | 0.016943 | 3.6975e-06 | -2.2722e-09 | 6.1874e-08 | -6.9263e-09 | 74.7657 | 108.4151 | 0.018664 | 3.1449e-06 | 1.8564e-10 | 5.8341e-08 | -3.228e-09 | 74.7936 | 86.7087 | 3.4028e+38 | 3.4028e+38 | 230.0103 | 230.0099 | 0.014643 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 1 | 230.009 | 230.0089 | 0.015451 | 49.9999 | -0.0034303 | 0.017258 | 3.6756e-06 | 1.3251e-09 | 6.3449e-08 | -6.385e-09 | 74.7402 | 84.5846 | 0.018698 | 3.183e-06 | 3.1178e-09 | 6.1251e-08 | -1.7216e-09 | 74.8348 | 35.9353 | 3.4028e+38 | 3.4028e+38 | 230.009 | 230.0089 | 0.015451 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 1 | 230.0102 | 230.0099 | 0.016278 | 49.9999 | -0.0040548 | 0.014242 | 3.5598e-06 | -2.4072e-09 | 5.0982e-08 | -3.8686e-09 | 74.7196 | 121.8909 | 0.015097 | 3.1352e-06 | -3.8503e-10 | 4.8564e-08 | -1.6198e-09 | 74.6643 | 110.5395 | 3.4028e+38 | 3.4028e+38 | 230.0102 | 230.0099 | 0.016278 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 1 | 230.0098 | 230.0097 | 0.015539 | 50 | -0.0035505 | 1 | 3.4028e+38 | 0 | 230.0049 | 230.0129 | 230.0127 | 0.020024 | -6.0007e-05 | 0.014516 | 3.6192e-06 | -2.6167e-09 | 5.3525e-08 | -4.5016e-09 | 74.7636 | 109.4188 | -0.033036 | 3.4028e+38 | 28.7785 | 0.020152 | 0 | 0 | 0 | -9.286e-05 | 0.016444 | 3.1091e-06 | -2.0223e-09 | 5.1116e-08 | -2.705e-09 | 74.8122 | 109.5554 | -0.018432 | 3.4028e+38 | 21.7087 | 229.7584 | -1.6528e-10 | 0.00071891 | 2.5097e-10 | 4.7775e-05 | 3.4028e+38 | 3.4028e+38 | 230.0101 | 230.0099 | 0.01551 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 230.0144 | 230.0144 | 0.023929 | 0.0002519 | 230.0016 | 1 | 230.0098 | 230.0095 | 0.015168 | 49.9999 | -0.0036801 | [17 empty] | 20.8464 | 12.9576 | 33.3672 | 32.3244 | 1914086.5093 |  | -0.0023835
60 | 230.0019 | 1 | 114.992 | 230.0099 | 199.2015 | 50 | -60.0039 | 0.015908 | 3.6273e-06 | 3.988e-09 | 5.771e-08 | -2.5452e-09 | 74.7626 | 39.0758 | 0.017365 | 3.2044e-06 | 3.0087e-09 | 5.565e-08 | -1.5982e-09 | 74.7843 | 26.5112 | 3.4028e+38 | 3.4028e+38 | 114.992 | 230.0099 | 199.2015 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | 114.9917 | 230.0103 | 199.2021 | 50 | -60.0039 | 0.015054 | 3.5768e-06 | 5.1556e-09 | 5.4461e-08 | 1.623e-09 | 74.7369 | -16.776 | 0.016227 | 3.1227e-06 | 4.624e-09 | 5.2511e-08 | 1.6933e-09 | 74.7233 | -22.472 | 3.4028e+38 | 3.4028e+38 | 114.9917 | 230.0103 | 199.2021 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | 114.9917 | 230.0097 | 199.2024 | 50 | -60.0038 | 0.01619 | 3.6096e-06 | 4.5175e-09 | 5.865e-08 | -3.864e-09 | 74.7587 | 49.2787 | 0.017373 | 3.1521e-06 | 2.6431e-09 | 5.477e-08 | -2.4553e-09 | 74.6994 | 46.0977 | 3.4028e+38 | 3.4028e+38 | 114.9917 | 230.0097 | 199.2024 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | 114.9924 | 230.0101 | 199.2026 | 49.9999 | -60.0038 | 0.015841 | 3.6308e-06 | 3.3837e-09 | 5.6343e-08 | -2.769e-09 | 74.7836 | 34.7654 | 0.016662 | 3.0948e-06 | 2.7848e-09 | 5.2217e-08 | -1.5635e-09 | 74.7253 | 59.801 | 3.4028e+38 | 3.4028e+38 | 114.9924 | 230.0101 | 199.2026 | 3.4028e+38 | 3.4028e+38 | 230.0011 | 1 | 114.9913 | 230.0096 | 199.2023 | 49.9999 | -60.0039 | 0.0143 | 3.6749e-06 | 2.6418e-09 | 5.192e-08 | 9.1256e-10 | 74.7485 | -15.2594 | 0.015355 | 3.1933e-06 | 1.8363e-09 | 4.7441e-08 | 3.7247e-10 | 74.7593 | -13.9453 | 3.4028e+38 | 3.4028e+38 | 114.9913 | 230.0096 | 199.2023 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0016 | 1 | 114.9913 | 230.0099 | 199.2024 | 50 | -60.0038 | 0.49994 | 3.4028e+38 | 0 | 230.0004 | 114.9842 | 230.0083 | 199.2011 | -0.00068247 | 0.015912 | 3.6059e-06 | 6.109e-09 | 5.861e-08 | 1.8781e-09 | 74.7466 | -17.089 | 0.091381 | 3.4028e+38 | -49.8083 | 0.020841 | 0 | 0 | 0 | -3.9311e-05 | 0.018919 | 3.1355e-06 | 6.9429e-09 | 5.9322e-08 | 2.8621e-09 | 74.7421 | -21.6342 | 0.12025 | 3.4028e+38 | -33.8157 | 229.1162 | 1.9674e-09 | 0.00074058 | -3.856e-10 | 1.4186e-05 | 3.4028e+38 | 3.4028e+38 | 114.9907 | 230.0087 | 199.2018 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 114.9827 | 114.9827 | 199.2025 | -0.00083069 | 230.0017 | 1 | 114.9918 | 230.0099 | 199.2022 | 50 | -60.0038 | [17 empty] | 20.9045 | 49.0455 | -90.4636 | 68.6351 | 122.5619 |  | -0.0053319
61 | 230.0016 | 1 | 114.9912 | 230.0089 | 199.2019 | 49.9999 | -60.0039 | 0.017672 | 3.582e-06 | 7.1865e-09 | 6.3312e-08 | -1.7155e-09 | 74.6335 | 17.5012 | 0.019518 | 3.1242e-06 | 7.6285e-09 | 6.1857e-08 | 1.8466e-09 | 74.8201 | -13.4449 | 3.4028e+38 | 3.4028e+38 | 114.9912 | 230.0089 | 199.2019 | 3.4028e+38 | 3.4028e+38 | 230.002 | 1 | 114.9919 | 230.0104 | 199.2022 | 49.9999 | -60.0041 | 0.014433 | 3.6698e-06 | 1.7733e-09 | 5.2353e-08 | -1.5089e-09 | 74.8155 | 25.3629 | 0.016001 | 3.1691e-06 | 1.8184e-09 | 4.8982e-08 | -1.0555e-09 | 74.7633 | 42.6202 | 3.4028e+38 | 3.4028e+38 | 114.9919 | 230.0104 | 199.2022 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 1 | 114.9926 | 230.0093 | 199.2018 | 49.9999 | -60.0038 | 0.017589 | 3.6322e-06 | 5.6114e-09 | 6.2825e-08 | -2.2232e-09 | 74.7212 | 21.6203 | 0.019124 | 3.214e-06 | 4.8716e-09 | 6.1539e-08 | 4.9753e-10 | 74.8042 | -4.9988 | 3.4028e+38 | 3.4028e+38 | 114.9926 | 230.0093 | 199.2018 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 1 | 114.9911 | 230.0097 | 199.2026 | 49.9999 | -60.0039 | 0.014733 | 3.5796e-06 | 3.3998e-09 | 5.3327e-08 | 9.7453e-10 | 74.7471 | -16.3924 | 0.015638 | 3.1391e-06 | 2.9777e-09 | 5.1143e-08 | 1.431e-09 | 74.7563 | -25.6678 | 3.4028e+38 | 3.4028e+38 | 114.9911 | 230.0097 | 199.2026 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 1 | 114.9914 | 230.0094 | 199.202 | 49.9999 | -60.0039 | 0.015881 | 3.6036e-06 | 4.1342e-09 | 5.6326e-08 | 1.1337e-09 | 74.6482 | -13.8648 | 0.017151 | 3.1502e-06 | 3.861e-09 | 5.381e-08 | 1.9253e-09 | 74.7009 | -29.7539 | 3.4028e+38 | 3.4028e+38 | 114.9914 | 230.0094 | 199.202 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0016 | 1 | 114.991 | 230.0096 | 199.2023 | 49.9999 | -60.0039 | 0.49994 | 3.4028e+38 | 0 | 229.9997 | 114.9836 | 230.0073 | 199.2003 | -0.00050706 | 0.015738 | 3.6343e-06 | 3.8353e-09 | 5.7486e-08 | 1.6748e-09 | 74.7333 | -20.593 | 0.077537 | 3.4028e+38 | 7.424 | 0.019763 | 0 | 0 | 0 | -3.4911e-05 | 0.016948 | 3.1325e-06 | 5.8306e-09 | 5.3099e-08 | 2.3216e-09 | 74.7689 | -21.5182 | 0.1089 | 3.4028e+38 | -61.8251 | 229.1172 | 4.3681e-10 | 0.00072127 | -6.4849e-10 | 2.5465e-05 | 3.4028e+38 | 3.4028e+38 | 114.991 | 230.0089 | 199.2023 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 114.9858 | 114.9854 | 199.2014 | -0.00046735 | 230.0016 | 1 | 114.9917 | 230.0095 | 199.2021 | 49.9999 | -60.0039 | [17 empty] | 20.2402 | 38.3914 | -82.4835 | 57.6296 | 106.4133 |  | -0.0047814
62 | 230.0013 | 1 | 114.9919 | 230.0091 | 199.2013 | 49.9999 | -60.0037 | 0.017747 | 3.6206e-06 | 5.6909e-09 | 6.2174e-08 | -1.8308e-09 | 74.7485 | 15.9792 | 0.019781 | 3.1884e-06 | 7.2165e-09 | 6.2221e-08 | 2.2125e-09 | 74.7423 | -11.6316 | 3.4028e+38 | 3.4028e+38 | 114.9919 | 230.0091 | 199.2013 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 1 | 114.9915 | 230.0092 | 199.2024 | 49.9999 | -60.0038 | 0.01479 | 3.5739e-06 | 3.8976e-09 | 5.2885e-08 | 1.6289e-09 | 74.7517 | -31.1214 | 0.016315 | 3.158e-06 | 3.9871e-09 | 5.1153e-08 | 7.2724e-10 | 74.7032 | -22.1935 | 3.4028e+38 | 3.4028e+38 | 114.9915 | 230.0092 | 199.2024 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 1 | 114.991 | 230.0099 | 199.2023 | 50 | -60.0038 | 0.014247 | 3.6093e-06 | 3.2167e-09 | 5.2291e-08 | 5.0181e-10 | 74.7263 | 13.0332 | 0.015275 | 3.0953e-06 | 1.2909e-09 | 4.7277e-08 | -9.9366e-10 | 74.7891 | 32.6252 | 3.4028e+38 | 3.4028e+38 | 114.991 | 230.0099 | 199.2023 | 3.4028e+38 | 3.4028e+38 | 230.0012 | 1 | 114.9915 | 230.0095 | 199.2023 | 50 | -60.0037 | 0.017832 | 3.5972e-06 | 6.1336e-09 | 6.4251e-08 | -1.3197e-09 | 74.6889 | 12.1424 | 0.019501 | 3.142e-06 | 4.5042e-09 | 6.1667e-08 | -6.9585e-10 | 74.7315 | -2.904 | 3.4028e+38 | 3.4028e+38 | 114.9915 | 230.0095 | 199.2023 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | 114.9917 | 230.0099 | 199.2022 | 50 | -60.0039 | 0.014965 | 3.5829e-06 | 4.1792e-09 | 5.4152e-08 | 2.0843e-09 | 74.7199 | -26.2525 | 0.016247 | 3.2304e-06 | 3.1264e-09 | 5.0397e-08 | 9.2792e-10 | 74.7156 | -21.1755 | 3.4028e+38 | 3.4028e+38 | 114.9917 | 230.0099 | 199.2022 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0011 | 1 | 114.9913 | 230.0102 | 199.2026 | 50 | -60.0039 | 0.49994 | 3.4028e+38 | 0 | 229.9998 | 114.9831 | 230.0077 | 199.2007 | -0.0007646 | 0.016343 | 3.6451e-06 | 6.324e-09 | 5.9535e-08 | 1.2231e-09 | 74.7199 | -10.8175 | 0.11017 | 3.4028e+38 | -4.7465 | 0.020287 | 0 | 0 | 0 | -4.2142e-05 | 0.019432 | 3.1918e-06 | 8.2077e-09 | 6.1622e-08 | 2.5262e-09 | 74.7396 | -17.4189 | 0.1093 | 3.4028e+38 | 9.597 | 229.608 | 1.5708e-13 | 0.00072337 | -3.9214e-10 | 9.6497e-06 | 3.4028e+38 | 3.4028e+38 | 114.9917 | 230.0093 | 199.2016 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 114.9869 | 114.9876 | 199.2069 | -0.00027454 | 230.0015 | 1 | 114.9915 | 230.0095 | 199.2021 | 50 | -60.0038 | [17 empty] | 20.0114 | 45.9596 | -84.4451 | 65.6058 | 117.6189 |  | -0.0049563
63 | 230.0015 | 1 | 183.9434 | 230.0085 | -138.0924 | 50 | 36.8967 | 0.015359 | 3.6585e-06 | -2.3317e-09 | 5.711e-08 | -5.3936e-09 | 74.6573 | 99.9282 | 0.016911 | 3.1416e-06 | 6.898e-10 | 5.3209e-08 | -2.3836e-09 | 74.8319 | 54.7947 | 3.4028e+38 | 3.4028e+38 | 183.9434 | 230.0085 | -138.0924 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 1 | 183.9447 | 230.0102 | -138.093 | 49.9999 | 36.8966 | 0.016215 | 3.6873e-06 | -5.4703e-09 | 5.9779e-08 | -5.4191e-09 | 74.6969 | 137.3464 | 0.017476 | 3.0379e-06 | -3.5034e-09 | 5.5114e-08 | -3.199e-09 | 74.7979 | 134.1677 | 3.4028e+38 | 3.4028e+38 | 183.9447 | 230.0102 | -138.093 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 1 | 183.9441 | 230.0096 | -138.0927 | 49.9998 | 36.8968 | 0.01665 | 3.6255e-06 | -7.2045e-09 | 6.4113e-08 | -8.0854e-09 | 74.7856 | 129.8398 | 0.0178 | 3.125e-06 | -1.7071e-09 | 5.8397e-08 | -3.0942e-09 | 74.7408 | 125.2023 | 3.4028e+38 | 3.4028e+38 | 183.9441 | 230.0096 | -138.0927 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 1 | 183.9446 | 230.0093 | -138.0922 | 50 | 36.8967 | 0.014364 | 3.5836e-06 | -2.4157e-09 | 5.1751e-08 | -2.7071e-09 | 74.7734 | 134.279 | 0.015298 | 3.1953e-06 | -2.1284e-10 | 4.8937e-08 | -1.7883e-09 | 74.761 | 109.1774 | 3.4028e+38 | 3.4028e+38 | 183.9446 | 230.0093 | -138.0922 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 1 | 183.9442 | 230.0097 | -138.0932 | 50 | 36.8968 | 0.014355 | 3.5944e-06 | -2.5446e-09 | 5.2247e-08 | -2.3864e-09 | 74.7607 | 142.7652 | 0.015426 | 3.1748e-06 | -1.3764e-09 | 4.894e-08 | -2.1444e-09 | 74.7749 | 118.0067 | 3.4028e+38 | 3.4028e+38 | 183.9442 | 230.0097 | -138.0932 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0021 | 1 | 183.9453 | 230.0102 | -138.0927 | 50 | 36.8966 | 0.79973 | 3.4028e+38 | 0 | 230.0063 | 183.9529 | 230.0141 | -138.0849 | 0.00077101 | 0.017435 | 3.6287e-06 | -6.3433e-09 | 6.407e-08 | -7.9821e-09 | 74.7628 | 122.9085 | -0.080624 | 3.4028e+38 | 1.9198 | 0.019928 | 0 | 0 | 0 | -9.8047e-05 | 0.019489 | 3.1622e-06 | -2.9659e-10 | 6.1668e-08 | -2.8991e-09 | 74.759 | 74.7607 | 0.013112 | 3.4028e+38 | -53.9156 | 229.023 | -9.7881e-10 | 0.00070978 | -7.8173e-10 | 3.5204e-05 | 3.4028e+38 | 3.4028e+38 | 183.9439 | 230.0095 | -138.0935 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 183.9478 | 183.9478 | -138.0836 | -6.8508e-05 | 230.0018 | 1 | 183.9442 | 230.0095 | -138.0927 | 49.9999 | 36.8967 | [17 empty] | 21.5057 | 45.6418 | 73.2282 | 66.4137 | 72.928 |  | 3.8574e-06
64 | 230.0016 | 1 | 183.9437 | 230.0092 | -138.0929 | 49.9999 | 36.8967 | 0.015169 | 3.6424e-06 | -1.3142e-09 | 5.3727e-08 | -4.3134e-09 | 74.7232 | 110.0787 | 0.0164 | 3.2229e-06 | 1.0538e-09 | 5.2893e-08 | -1.8174e-09 | 74.8098 | 90.5274 | 3.4028e+38 | 3.4028e+38 | 183.9436 | 230.0092 | -138.0929 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 1 | 183.9448 | 230.0102 | -138.0935 | 50 | 36.8968 | 0.017362 | 3.6608e-06 | -4.6086e-09 | 6.3591e-08 | -6.6877e-09 | 74.7842 | 132.997 | 0.019112 | 3.1046e-06 | -2.1896e-09 | 5.9601e-08 | -3.1439e-09 | 74.7502 | 125.1098 | 3.4028e+38 | 3.4028e+38 | 183.9448 | 230.0102 | -138.0935 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | 183.9438 | 230.0093 | -138.0929 | 49.9999 | 36.8968 | 0.016834 | 3.6126e-06 | -3.246e-09 | 5.963e-08 | -6.0954e-09 | 74.7155 | 119.1714 | 0.018373 | 3.088e-06 | 5.5637e-10 | 5.6692e-08 | -2.546e-09 | 74.7333 | 72.4177 | 3.4028e+38 | 3.4028e+38 | 183.9438 | 230.0093 | -138.0929 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | 183.9441 | 230.0094 | -138.0926 | 49.9999 | 36.8968 | 0.014601 | 3.6662e-06 | -1.0138e-09 | 5.4711e-08 | -3.8762e-09 | 74.7221 | 118.7974 | 0.015802 | 3.0948e-06 | -3.2556e-10 | 5.006e-08 | -1.8585e-09 | 74.7486 | 59.2777 | 3.4028e+38 | 3.4028e+38 | 183.9441 | 230.0094 | -138.0926 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 1 | 183.944 | 230.0103 | -138.0931 | 50 | 36.8967 | 0.017592 | 3.6803e-06 | -7.0865e-09 | 6.3764e-08 | -8.324e-09 | 74.7195 | 127.881 | 0.019618 | 3.1828e-06 | -2.2642e-09 | 6.0786e-08 | -3.0898e-09 | 74.7398 | 116.3915 | 3.4028e+38 | 3.4028e+38 | 183.944 | 230.0103 | -138.0932 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 1 | 183.9445 | 230.0104 | -138.0932 | 50 | 36.8968 | 0.79972 | 3.4028e+38 | 0 | 230.0057 | 183.9538 | 230.014 | -138.0836 | 0.00077684 | 0.017538 | 3.6298e-06 | -5.0134e-09 | 6.3695e-08 | -7.7073e-09 | 74.7789 | 129.5146 | -0.10042 | 3.4028e+38 | -22.9001 | 0.020378 | 0 | 0 | 0 | -9.5188e-05 | 0.018759 | 3.092e-06 | -2.3948e-09 | 5.8005e-08 | -1.8403e-09 | 74.6802 | 137.8928 | -0.034726 | 3.4028e+38 | -24.57 | 229.1389 | -7.5092e-11 | 0.00072344 | 2.4479e-10 | 3.7188e-05 | 3.4028e+38 | 3.4028e+38 | 183.9445 | 230.01 | -138.0933 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 183.9529 | 183.9529 | -138.087 | 0.00046091 | 230.0017 | 1 | 183.9441 | 230.0097 | -138.093 | 49.9999 | 36.8968 | [17 empty] | 20.9972 | 45.8206 | 78.1499 | 67.223 | 64.9251 |  | -0.00038829
65 | 230.0016 | 1 | 183.9443 | 230.0095 | -138.0932 | 49.9999 | 36.8968 | 0.017564 | 3.6722e-06 | -6.4273e-09 | 6.2599e-08 | -7.4594e-09 | 74.8075 | 134.6142 | 0.018983 | 3.2185e-06 | -2.0885e-09 | 5.8543e-08 | -3.728e-09 | 74.6605 | 126.011 | 3.4028e+38 | 3.4028e+38 | 183.9443 | 230.0095 | -138.0932 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | 183.9437 | 230.0099 | -138.0941 | 50 | 36.8968 | 0.01742 | 3.68e-06 | -4.5999e-09 | 6.4026e-08 | -9.2933e-09 | 74.6979 | 116.655 | 0.018944 | 3.1758e-06 | -8.4425e-11 | 6.0843e-08 | -2.6601e-09 | 74.7746 | 107.3879 | 3.4028e+38 | 3.4028e+38 | 183.9437 | 230.0099 | -138.0941 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 1 | 183.945 | 230.0093 | -138.0921 | 49.9999 | 36.8966 | 0.014956 | 3.6262e-06 | -2.6245e-09 | 5.4693e-08 | -3.7131e-09 | 74.7876 | 141.2711 | 0.016312 | 3.0968e-06 | -1.9856e-09 | 5.2726e-08 | -2.4017e-09 | 74.8034 | 132.9505 | 3.4028e+38 | 3.4028e+38 | 183.945 | 230.0093 | -138.0921 | 3.4028e+38 | 3.4028e+38 | 230.002 | 1 | 183.9447 | 230.0103 | -138.0934 | 50 | 36.8967 | 0.014718 | 3.5948e-06 | -2.329e-09 | 5.2987e-08 | -3.534e-09 | 74.7671 | 134.8203 | 0.015936 | 3.0883e-06 | -1.7113e-09 | 4.8587e-08 | -2.4352e-09 | 74.7429 | 128.6198 | 3.4028e+38 | 3.4028e+38 | 183.9447 | 230.0103 | -138.0934 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 1 | 183.9441 | 230.01 | -138.0935 | 50 | 36.8967 | 0.016812 | 3.6434e-06 | -3.8203e-09 | 6.0271e-08 | -8.2573e-09 | 74.7211 | 117.2895 | 0.018653 | 3.1552e-06 | 7.7608e-10 | 5.9987e-08 | -3.0665e-09 | 74.8183 | 67.6902 | 3.4028e+38 | 3.4028e+38 | 183.9441 | 230.01 | -138.0935 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0016 | 1 | 183.9436 | 230.0092 | -138.0929 | 49.9999 | 36.8968 | 0.79972 | 3.4028e+38 | 0 | 230.0012 | 183.9489 | 230.0092 | -138.0842 | 0.00025156 | 0.01701 | 3.6182e-06 | -4.6396e-09 | 6.2215e-08 | -8.3981e-09 | 74.7344 | 114.3064 | -0.062259 | 3.4028e+38 | 51.8976 | 0.020192 | 0 | 0 | 0 | -9.9897e-05 | 0.019302 | 3.1702e-06 | 8.7592e-10 | 5.9285e-08 | -2.9397e-09 | 74.7404 | 67.9397 | -0.0037191 | 3.4028e+38 | 12.356 | 229.6349 | -5.1523e-11 | 0.00069299 | 3.5258e-10 | 3.3342e-05 | 3.4028e+38 | 3.4028e+38 | 183.9436 | 230.0101 | -138.0945 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 183.9527 | 183.9536 | -138.0833 | 0.00071914 | 230.0019 | 1 | 183.9444 | 230.0098 | -138.0932 | 50 | 36.8967 | [17 empty] | 21.246 | 52.1514 | 77.3729 | 72.213 | 83.3442 |  | 8.1526e-05
66 | 230.0022 | 1 | 184.0179 | 230.0114 | -137.9976 | 49.9999 | 36.8669 | 0.017265 | 3.4545e-06 | -3.5165e-09 | 5.9565e-08 | -8.0513e-09 | 74.6655 | 113.594 | 0.019168 | 3.1121e-06 | 6.6831e-10 | 5.9715e-08 | -3.1433e-09 | 74.7427 | 90.1837 | 3.4028e+38 | 3.4028e+38 | 184.0178 | 230.0114 | -137.9976 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 1 | 184.0182 | 230.0117 | -137.9982 | 49.9999 | 36.8667 | 0.017587 | 3.4772e-06 | -6.1542e-09 | 6.1018e-08 | -8.0444e-09 | 74.7492 | 123.3593 | 0.019287 | 3.0348e-06 | -2.4876e-09 | 6.0247e-08 | -3.1613e-09 | 74.887 | 128.1988 | 3.4028e+38 | 3.4028e+38 | 184.0182 | 230.0117 | -137.9982 | 3.4028e+38 | 3.4028e+38 | 230.0029 | 1 | 184.0188 | 230.0111 | -137.9974 | 49.9999 | 36.8667 | 0.015362 | 3.467e-06 | -5.2699e-09 | 5.328e-08 | -4.6957e-09 | 74.7758 | 136.7987 | 0.016917 | 2.9901e-06 | -3.516e-09 | 5.0181e-08 | -3.1184e-09 | 74.5929 | 124.6035 | 3.4028e+38 | 3.4028e+38 | 184.0188 | 230.0111 | -137.9974 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 1 | 184.0175 | 230.011 | -137.9978 | 50 | 36.867 | 0.016314 | 3.4174e-06 | -1.8211e-09 | 5.7367e-08 | -7.0383e-09 | 74.7507 | 116.4931 | 0.01803 | 3.1523e-06 | 5.6949e-10 | 5.6917e-08 | -2.5651e-09 | 74.7939 | 107.1168 | 3.4028e+38 | 3.4028e+38 | 184.0175 | 230.011 | -137.9978 | 3.4028e+38 | 3.4028e+38 | 230.0024 | 1 | 184.018 | 230.0114 | -137.9981 | 49.9999 | 36.8667 | 0.015181 | 3.423e-06 | -4.4677e-09 | 5.2612e-08 | -3.4345e-09 | 74.7007 | 142.449 | 0.01679 | 3.1315e-06 | -8.7169e-10 | 5.2562e-08 | -2.2966e-09 | 74.7705 | 137.8211 | 3.4028e+38 | 3.4028e+38 | 184.0179 | 230.0114 | -137.9981 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0018 | 1 | 184.0172 | 230.0117 | -137.9975 | 50 | 36.8666 | 0.80004 | 3.4028e+38 | 0 | 230.0051 | 184.0223 | 230.0139 | -137.9934 | 2.378e-05 | 0.014294 | 3.4699e-06 | -1.3431e-09 | 4.9673e-08 | -3.0978e-09 | 74.5597 | 118.8469 | -0.039174 | 3.4028e+38 | -51.3824 | 0.019719 | 0 | 0 | 0 | -7.8225e-05 | 0.015301 | 3.0825e-06 | -3.2111e-10 | 4.7761e-08 | -7.9555e-10 | 74.8008 | 111.9802 | 0.00082576 | 3.4028e+38 | 19.617 | 229.8502 | 5.3027e-10 | 0.00070457 | 1.0426e-10 | 5.1655e-05 | 3.4028e+38 | 3.4028e+38 | 184.0171 | 230.0106 | -137.9971 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 184.0196 | 184.0196 | -137.9921 | -0.0001306 | 230.0024 | 1 | 184.0181 | 230.0113 | -137.9978 | 49.9999 | 36.8668 | [17 empty] | 25.472 | 38.5651 | 76.7217 | 62.5291 | 58.9855 |  | -0.00043548
67 | 230.0018 | 1 | 184.0164 | 230.0101 | -137.9982 | 50 | 36.867 | 0.017145 | 3.4283e-06 | -3.7659e-09 | 6.0283e-08 | -9.0852e-09 | 74.78 | 112.0458 | 0.019235 | 3.0231e-06 | 8.6658e-10 | 6.038e-08 | -2.0681e-09 | 74.734 | 107.7914 | 3.4028e+38 | 3.4028e+38 | 184.0164 | 230.0101 | -137.9982 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 1 | 184.018 | 230.0112 | -137.9975 | 50 | 36.8667 | 0.016697 | 3.4545e-06 | -5.9357e-09 | 5.7661e-08 | -7.0169e-09 | 74.7709 | 129.0526 | 0.018581 | 3.1208e-06 | -1.6119e-09 | 5.506e-08 | -3.4059e-09 | 74.7708 | 125.6321 | 3.4028e+38 | 3.4028e+38 | 184.018 | 230.0112 | -137.9975 | 3.4028e+38 | 3.4028e+38 | 230.0024 | 1 | 184.0165 | 230.01 | -137.9969 | 49.9999 | 36.8669 | 0.015443 | 3.4227e-06 | -2.1768e-09 | 5.2858e-08 | -6.1432e-09 | 74.7751 | 110.266 | 0.01675 | 3.0883e-06 | 3.493e-10 | 5.1665e-08 | -1.7032e-09 | 74.8571 | 85.2206 | 3.4028e+38 | 3.4028e+38 | 184.0165 | 230.01 | -137.9969 | 3.4028e+38 | 3.4028e+38 | 230.0025 | 1 | 184.0169 | 230.0102 | -137.9971 | 49.9999 | 36.8669 | 0.014977 | 3.4921e-06 | -1.1313e-09 | 5.0597e-08 | -4.2424e-09 | 74.7307 | 103.9464 | 0.016093 | 3.0739e-06 | -1.1682e-10 | 4.9469e-08 | -1.5894e-09 | 74.752 | 77.5679 | 3.4028e+38 | 3.4028e+38 | 184.0169 | 230.0102 | -137.9971 | 3.4028e+38 | 3.4028e+38 | 230.0024 | 1 | 184.0168 | 230.0102 | -137.9974 | 49.9999 | 36.8669 | 0.014717 | 3.4611e-06 | -2.0516e-09 | 5.1163e-08 | -4.4206e-09 | 74.7352 | 114.8959 | 0.016003 | 3.1241e-06 | 3.9532e-10 | 4.9895e-08 | -2.319e-09 | 74.7848 | 93.2213 | 3.4028e+38 | 3.4028e+38 | 184.0168 | 230.0102 | -137.9974 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0024 | 1 | 184.016 | 230.0096 | -137.997 | 49.9999 | 36.867 | 0.80004 | 3.4028e+38 | 0 | 230.0013 | 184.0188 | 230.0089 | -137.9917 | -0.00017222 | 0.014702 | 3.4635e-06 | -1.1595e-09 | 5.0905e-08 | -4.1697e-09 | 74.8097 | 122.1374 | -0.05167 | 3.4028e+38 | -28.7301 | 0.020377 | 0 | 0 | 0 | -8.4267e-05 | 0.015908 | 3.0417e-06 | 2.7673e-10 | 4.8723e-08 | -1.4159e-09 | 74.8109 | 78.9417 | 0.015192 | 3.4028e+38 | -2.6647 | 229.4394 | 2.9216e-10 | 0.00070726 | 1.7002e-10 | 5.4343e-05 | 3.4028e+38 | 3.4028e+38 | 184.0176 | 230.0108 | -137.9969 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 184.0191 | 184.0191 | -137.9926 | -0.00011373 | 230.0022 | 1 | 184.0169 | 230.0103 | -137.9974 | 49.9999 | 36.8669 | [17 empty] | 23.2618 | 43.1881 | 76.7887 | 65.7944 | 65.8371 |  | -0.00028671
68 | 230.0018 | 1 | 184.0168 | 230.011 | -137.9974 | 49.9999 | 36.8668 | 0.017717 | 3.431e-06 | -6.6713e-09 | 6.1372e-08 | -9.3479e-09 | 74.8601 | 124.9582 | 0.019531 | 3.1048e-06 | -3.8976e-10 | 6.0631e-08 | -2.9465e-09 | 74.8092 | 112.1714 | 3.4028e+38 | 3.4028e+38 | 184.0168 | 230.011 | -137.9975 | 3.4028e+38 | 3.4028e+38 | 230.002 | 1 | 184.0173 | 230.0102 | -137.9961 | 50 | 36.8665 | 0.014727 | 3.4311e-06 | -4.3294e-09 | 5.064e-08 | -3.4584e-09 | 74.6758 | 141.3813 | 0.015871 | 3.1284e-06 | -7.6282e-10 | 4.7982e-08 | -1.9317e-09 | 74.7799 | 126.1547 | 3.4028e+38 | 3.4028e+38 | 184.0173 | 230.0102 | -137.9961 | 3.4028e+38 | 3.4028e+38 | 230.002 | 1 | 184.0172 | 230.0105 | -137.9973 | 50 | 36.8667 | 0.017159 | 3.4006e-06 | -5.4323e-09 | 5.8351e-08 | -6.7709e-09 | 74.7008 | 128.676 | 0.018841 | 3.0687e-06 | -3.2618e-09 | 5.8565e-08 | -3.1888e-09 | 74.7271 | 135.6483 | 3.4028e+38 | 3.4028e+38 | 184.0172 | 230.0105 | -137.9973 | 3.4028e+38 | 3.4028e+38 | 230.002 | 1 | 184.0172 | 230.0104 | -137.9965 | 50 | 36.8666 | 0.015694 | 3.4228e-06 | -4.8696e-09 | 5.374e-08 | -4.1728e-09 | 74.7151 | 139.7312 | 0.017416 | 3.1184e-06 | -2.0935e-09 | 5.3272e-08 | -3.0631e-09 | 74.7777 | 135.6238 | 3.4028e+38 | 3.4028e+38 | 184.0172 | 230.0104 | -137.9965 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 1 | 184.0163 | 230.01 | -137.9977 | 50 | 36.8668 | 0.014837 | 3.4578e-06 | -1.7762e-09 | 4.9789e-08 | -3.2348e-09 | 74.6083 | 112.0291 | 0.015761 | 3.0259e-06 | -2.0348e-10 | 4.7703e-08 | -1.3507e-09 | 74.7775 | 59.6486 | 3.4028e+38 | 3.4028e+38 | 184.0163 | 230.01 | -137.9977 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0021 | 1 | 184.0163 | 230.0097 | -137.9966 | 50 | 36.8668 | 0.80004 | 3.4028e+38 | 0 | 230.0049 | 184.0215 | 230.0128 | -137.9928 | 0.00023973 | 0.014818 | 3.4147e-06 | -5.0313e-09 | 5.1118e-08 | -4.2465e-09 | 74.6833 | 142.2292 | -0.091818 | 3.4028e+38 | -56.2126 | 0.020052 | 0 | 0 | 0 | -9.0842e-05 | 0.016995 | 3.0878e-06 | -3.1938e-09 | 5.3605e-08 | -2.8362e-09 | 74.7609 | 138.394 | -0.050532 | 3.4028e+38 | 35.4599 | 229.1062 | -3.3027e-10 | 0.00069113 | -3.7832e-10 | 3.767e-05 | 3.4028e+38 | 3.4028e+38 | 184.0172 | 230.0105 | -137.9969 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 184.0263 | 184.0263 | -137.9882 | 0.00056299 | 230.002 | 1 | 184.017 | 230.0104 | -137.997 | 50 | 36.8667 | [17 empty] | 23.7246 | 38.827 | 74.8559 | 62.4089 | 56.5651 |  | -0.00058592
69 | 230.0018 | 0.50002 | 92.0094 | 115.0056 | -68.9981 | 49.9999 | 36.8665 | 0.017825 | 3.3946e-06 | -3.5446e-09 | 6.0579e-08 | -5.6779e-09 | 74.6477 | 120.5932 | 0.019761 | 3.0958e-06 | 4.2549e-10 | 6.11e-08 | -1.0582e-09 | 74.7297 | 22.3334 | 3.4028e+38 | 3.4028e+38 | 92.0094 | 115.0056 | -68.9981 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 0.50002 | 92.0094 | 115.0056 | -68.998 | 50 | 36.8663 | 0.014364 | 3.3337e-06 | -3.3312e-10 | 4.8066e-08 | -1.4436e-09 | 74.7793 | 102.9943 | 0.015164 | 3.0403e-06 | 4.2693e-10 | 4.7335e-08 | -9.3347e-10 | 74.709 | 80.6943 | 3.4028e+38 | 3.4028e+38 | 92.0094 | 115.0056 | -68.998 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 0.50002 | 92.0088 | 115.0053 | -68.9983 | 49.9999 | 36.8664 | 0.015499 | 3.3984e-06 | -7.2212e-11 | 5.2948e-08 | -3.298e-09 | 74.7058 | 123.27 | 0.017061 | 2.9738e-06 | 1.5009e-09 | 4.9883e-08 | -5.0175e-10 | 74.7668 | 11.7963 | 3.4028e+38 | 3.4028e+38 | 92.0087 | 115.0053 | -68.9983 | 3.4028e+38 | 3.4028e+38 | 230.002 | 0.50002 | 92.0097 | 115.0061 | -68.9983 | 49.9998 | 36.8665 | 0.017774 | 3.4168e-06 | -3.9968e-09 | 6.0651e-08 | -5.097e-09 | 74.5747 | 118.3459 | 0.019658 | 3.1132e-06 | 3.4472e-10 | 6.144e-08 | -1.5199e-09 | 74.8283 | 64.1463 | 3.4028e+38 | 3.4028e+38 | 92.0097 | 115.0061 | -68.9983 | 3.4028e+38 | 3.4028e+38 | 230.002 | 0.50002 | 92.0089 | 115.0055 | -68.9985 | 49.9999 | 36.8666 | 0.017394 | 3.35e-06 | -2.3317e-09 | 5.9961e-08 | -5.1017e-09 | 74.8282 | 114.5628 | 0.01941 | 3.0263e-06 | 5.2961e-10 | 5.9149e-08 | -2.8782e-10 | 74.6843 | 28.522 | 3.4028e+38 | 3.4028e+38 | 92.0089 | 115.0055 | -68.9986 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 0.50002 | 92.0089 | 115.0054 | -68.9982 | 49.9999 | 36.8664 | 0.80004 | 3.4028e+38 | 0 | 230.0026 | 92.0125 | 115.0058 | -68.9931 | 0.00028998 | 0.017435 | 3.384e-06 | -2.8944e-09 | 5.8929e-08 | -5.2541e-09 | 74.7616 | 116.617 | -0.042678 | 3.4028e+38 | -37.2069 | 0.02063 | 0 | 0 | 0 | -5.4877e-05 | 0.01936 | 3.0963e-06 | 6.2117e-10 | 6.0861e-08 | -5.4781e-10 | 74.8253 | 77.9431 | 0.0030087 | 3.4028e+38 | -7.3759 | 229.0849 | -1.0805e-09 | 0.0006972 | -5.5706e-11 | 9.4112e-06 | 3.4028e+38 | 3.4028e+38 | 92.009 | 115.0051 | -68.998 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 92.0133 | 92.0133 | -68.993 | 0.00028804 | 230.002 | 0.50002 | 92.0092 | 115.0056 | -68.9983 | 49.9999 | 36.8664 | [17 empty] | 23.2705 | 51.5717 | 85.5277 | 73.6492 | 70.1002 |  | -0.00042855
70 | 230.0018 | 0.50002 | 92.009 | 115.0053 | -68.998 | 49.9999 | 36.8664 | 0.015079 | 3.3464e-06 | -1.0165e-09 | 5.0363e-08 | -2.58e-09 | 74.747 | 119.0846 | 0.01616 | 3.0863e-06 | 6.304e-10 | 5.0436e-08 | -5.5196e-10 | 74.7644 | 5.6121 | 3.4028e+38 | 3.4028e+38 | 92.009 | 115.0053 | -68.998 | 3.4028e+38 | 3.4028e+38 | 230.002 | 0.50002 | 92.0091 | 115.0055 | -68.9983 | 49.9999 | 36.8665 | 0.017125 | 3.3917e-06 | -1.9301e-09 | 5.8022e-08 | -3.9739e-09 | 74.6908 | 112.155 | 0.019026 | 3.0337e-06 | 2.3726e-09 | 5.7739e-08 | -1.6411e-10 | 74.7367 | 5.5419 | 3.4028e+38 | 3.4028e+38 | 92.0091 | 115.0055 | -68.9983 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 0.50002 | 92.0095 | 115.0056 | -68.9981 | 49.9999 | 36.8662 | 0.014807 | 3.4003e-06 | -1.417e-09 | 4.8694e-08 | -1.9097e-09 | 74.7632 | 116.1161 | 0.015824 | 3.0473e-06 | -8.1886e-10 | 4.7966e-08 | -1.1447e-09 | 74.7003 | 98.1898 | 3.4028e+38 | 3.4028e+38 | 92.0095 | 115.0056 | -68.9981 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 0.50002 | 92.0095 | 115.0056 | -68.9979 | 49.9999 | 36.8662 | 0.01438 | 3.3377e-06 | -1.3428e-09 | 4.865e-08 | -1.2385e-09 | 74.7036 | 130.9546 | 0.015304 | 3.0003e-06 | 2.575e-10 | 4.7101e-08 | -1.1829e-09 | 74.691 | 77.7189 | 3.4028e+38 | 3.4028e+38 | 92.0095 | 115.0056 | -68.998 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 0.50002 | 92.0098 | 115.006 | -68.9983 | 49.9999 | 36.8664 | 0.016636 | 3.3626e-06 | -3.0428e-09 | 5.5943e-08 | -2.8571e-09 | 74.7278 | 121.5477 | 0.017824 | 3.0136e-06 | -1.1902e-09 | 5.3721e-08 | -1.197e-09 | 74.7175 | 99.4143 | 3.4028e+38 | 3.4028e+38 | 92.0098 | 115.006 | -68.9983 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0018 | 0.50002 | 92.0097 | 115.0055 | -68.9981 | 49.9999 | 36.8663 | 0.80004 | 3.4028e+38 | 0 | 230.007 | 92.0132 | 115.0084 | -68.9961 | 0.00013405 | 0.015188 | 3.3724e-06 | -1.6028e-09 | 5.2152e-08 | -2.5268e-09 | 74.7431 | 137.0751 | -0.052094 | 3.4028e+38 | -5.0546 | 0.020291 | 0 | 0 | 0 | -3.9064e-05 | 0.015821 | 3.0754e-06 | -1.8533e-10 | 4.7189e-08 | -6.5959e-10 | 74.8515 | 109.6408 | -0.0049884 | 3.4028e+38 | 15.3269 | 229.4561 | -2.429e-10 | 0.00070233 | -1.9224e-10 | 2.7455e-05 | 3.4028e+38 | 3.4028e+38 | 92.0092 | 115.0053 | -68.9976 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 92.0103 | 92.0103 | -68.9955 | -6.3093e-05 | 230.0021 | 0.50002 | 92.0094 | 115.0056 | -68.9981 | 49.9999 | 36.8663 | [17 empty] | 22.8884 | 46.1397 | 77.8775 | 68.1386 | 69.2693 |  | -0.00025383
71 | 230.0021 | 0.50002 | 92.0098 | 115.0061 | -68.998 | 50 | 36.8662 | 0.01527 | 3.3824e-06 | -2.6631e-09 | 5.1684e-08 | -1.808e-09 | 74.8371 | 142.1617 | 0.016602 | 3.0526e-06 | -1.3759e-10 | 5.0147e-08 | -1.4014e-09 | 74.7642 | 139.4084 | 3.4028e+38 | 3.4028e+38 | 92.0098 | 115.0061 | -68.998 | 3.4028e+38 | 3.4028e+38 | 230.002 | 0.50002 | 92.0091 | 115.0054 | -68.9981 | 49.9999 | 36.8664 | 0.015807 | 3.3807e-06 | -8.5548e-10 | 5.3455e-08 | -3.1944e-09 | 74.75 | 91.576 | 0.017226 | 3.0659e-06 | 1.9895e-09 | 5.262e-08 | 6.787e-10 | 74.7505 | 25.3026 | 3.4028e+38 | 3.4028e+38 | 92.0091 | 115.0054 | -68.9981 | 3.4028e+38 | 3.4028e+38 | 230.0024 | 0.50002 | 92.0089 | 115.0052 | -68.998 | 49.9999 | 36.8668 | 0.016962 | 3.428e-06 | -2.5716e-09 | 5.7266e-08 | -4.0089e-09 | 74.7558 | 127.4689 | 0.018854 | 3.0276e-06 | 1.3108e-09 | 5.7273e-08 | 3.3842e-11 | 74.7479 | -1.0389 | 3.4028e+38 | 3.4028e+38 | 92.0089 | 115.0052 | -68.998 | 3.4028e+38 | 3.4028e+38 | 230.002 | 0.50002 | 92.0097 | 115.0062 | -68.9983 | 49.9999 | 36.8663 | 0.01608 | 3.3426e-06 | -1.7424e-09 | 5.3657e-08 | -2.763e-09 | 74.7577 | 127.4497 | 0.01769 | 3.112e-06 | -3.3395e-10 | 5.3047e-08 | -1.4702e-09 | 74.7712 | 138.834 | 3.4028e+38 | 3.4028e+38 | 92.0097 | 115.0062 | -68.9983 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 0.50002 | 92.0089 | 115.0054 | -68.9984 | 50 | 36.8666 | 0.01663 | 3.3444e-06 | -3.0722e-09 | 5.5552e-08 | -4.5319e-09 | 74.7414 | 116.6951 | 0.018114 | 3.1286e-06 | 8.2675e-10 | 5.7154e-08 | -6.9608e-10 | 74.7971 | -1.5029 | 3.4028e+38 | 3.4028e+38 | 92.0089 | 115.0054 | -68.9984 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0021 | 0.50002 | 92.009 | 115.0053 | -68.9982 | 49.9999 | 36.8665 | 0.80004 | 3.4028e+38 | 0 | 230.0018 | 92.0116 | 115.0055 | -68.9941 | 0.00015676 | 0.017126 | 3.3759e-06 | -1.9911e-09 | 5.7824e-08 | -4.0696e-09 | 74.7728 | 109.8524 | -0.025411 | 3.4028e+38 | -17.8394 | 0.020427 | 0 | 0 | 0 | -5.3236e-05 | 0.019193 | 3.0417e-06 | 5.8883e-10 | 5.9832e-08 | -9.921e-10 | 74.7463 | 33.9029 | 0.024673 | 3.4028e+38 | -34.2567 | 229.7067 | 3.5891e-10 | 0.00070863 | -2.2415e-10 | 1.29e-05 | 3.4028e+38 | 3.4028e+38 | 92.0086 | 115.0052 | -68.9988 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 92.0123 | 92.0123 | -68.9935 | 0.00016523 | 230.0022 | 0.50002 | 92.0093 | 115.0056 | -68.9982 | 49.9999 | 36.8664 | [17 empty] | 22.4199 | 49.9439 | 82.0381 | 72.6133 | 72.6521 |  | -0.0002069
72 | 230.002 | 0.50002 | 115.0054 | 115.0053 | 0.0080211 | 49.9999 | -0.0043472 | 0.016169 | 3.3442e-06 | 4.0722e-10 | 5.2754e-08 | -1.4972e-09 | 74.7659 | 61.9974 | 0.017444 | 3.0479e-06 | 3.8287e-09 | 5.3172e-08 | 5.5259e-10 | 74.8141 | -5.9281 | 3.4028e+38 | 3.4028e+38 | 115.0054 | 115.0053 | 0.0080211 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 0.50002 | 115.0051 | 115.005 | 0.0082474 | 49.9999 | -0.0041088 | 0.014823 | 3.372e-06 | 5.1761e-10 | 4.9974e-08 | -1.4828e-09 | 74.6574 | 37.7778 | 0.016345 | 3.0816e-06 | 1.0867e-09 | 5.0465e-08 | -6.1151e-10 | 74.8376 | 22.1339 | 3.4028e+38 | 3.4028e+38 | 115.0051 | 115.005 | 0.0082474 | 3.4028e+38 | 3.4028e+38 | 230.002 | 0.50002 | 115.0057 | 115.0056 | 0.0081696 | 49.9999 | -0.0040701 | 0.016655 | 3.3008e-06 | -7.312e-10 | 5.5035e-08 | -3.4467e-09 | 74.6904 | 104.7252 | 0.0181 | 2.9868e-06 | 1.7375e-09 | 5.6201e-08 | -1.4475e-09 | 74.7552 | 75.3961 | 3.4028e+38 | 3.4028e+38 | 115.0057 | 115.0056 | 0.0081696 | 3.4028e+38 | 3.4028e+38 | 230.0025 | 0.50002 | 115.0058 | 115.0057 | 0.0087153 | 50 | -0.0043561 | 0.014303 | 3.3649e-06 | 8.294e-10 | 4.7652e-08 | -1.0649e-09 | 74.6984 | 132.3929 | 0.015417 | 3.1169e-06 | 4.3997e-10 | 4.5415e-08 | -9.7717e-10 | 74.755 | 144.5635 | 3.4028e+38 | 3.4028e+38 | 115.0058 | 115.0057 | 0.0087153 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 0.50002 | 115.0051 | 115.0051 | 0.0081363 | 50 | -0.0040535 | 0.016425 | 3.3452e-06 | -6.6274e-11 | 5.5781e-08 | -3.1463e-09 | 74.7419 | 84.2925 | 0.018087 | 3.1247e-06 | 3.4528e-09 | 5.7581e-08 | 1.4844e-10 | 74.7248 | -1.2847 | 3.4028e+38 | 3.4028e+38 | 115.0051 | 115.0051 | 0.0081363 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0022 | 0.50002 | 115.0056 | 115.0055 | 0.008186 | 50 | -0.0040782 | 1 | 3.4028e+38 | 0 | 230.0031 | 115.0062 | 115.006 | 0.012812 | 6.7988e-05 | 0.017865 | 3.3945e-06 | -1.4636e-10 | 6.0576e-08 | -4.705e-09 | 74.7348 | 96.3655 | -0.008523 | 3.4028e+38 | 9.2778 | 0.020861 | 0 | 0 | 0 | -6.1199e-05 | 0.019455 | 3.0209e-06 | 1.6759e-09 | 5.8712e-08 | -8.5041e-10 | 74.7425 | 21.5099 | 0.035479 | 3.4028e+38 | -39.9538 | 229.4163 | -1.1764e-09 | 0.00068502 | -2.8943e-10 | 1.4828e-05 | 3.4028e+38 | 3.4028e+38 | 115.0061 | 115.006 | 0.0078852 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 115.0079 | 115.008 | 0.01344 | 1.998e-05 | 230.0021 | 0.50002 | 115.0054 | 115.0053 | 0.0082579 | 49.9999 | -0.0041871 | [17 empty] | 22.8856 | 13.3738 | 36.0022 | 35.3573 | 1025469.6111 |  | -0.0021559
73 | 230.0022 | 0.50002 | 115.0056 | 115.0055 | 0.0087832 | 49.9999 | -0.0043758 | 0.014888 | 3.3804e-06 | -2.082e-10 | 5.0359e-08 | -1.7758e-09 | 74.6989 | 119.0584 | 0.0162 | 3.042e-06 | -1.0285e-09 | 4.9871e-08 | -1.9641e-09 | 74.6845 | 109.542 | 3.4028e+38 | 3.4028e+38 | 115.0056 | 115.0055 | 0.0087832 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 0.50002 | 115.0061 | 115.006 | 0.0086287 | 50 | -0.0044652 | 0.014847 | 3.3484e-06 | -6.5825e-10 | 4.9648e-08 | -1.8232e-09 | 74.6809 | 127.1585 | 0.015896 | 3.018e-06 | -3.6898e-10 | 4.8273e-08 | -1.9707e-09 | 74.8457 | 116.4127 | 3.4028e+38 | 3.4028e+38 | 115.0061 | 115.006 | 0.0086287 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 0.50002 | 115.0057 | 115.0057 | 0.0076745 | 49.9999 | -0.0038234 | 0.017165 | 3.3984e-06 | -4.5815e-10 | 5.8296e-08 | -3.7216e-09 | 74.7784 | 114.0269 | 0.018839 | 3.0896e-06 | 4.7434e-10 | 5.8048e-08 | -1.814e-09 | 74.8116 | 89.1627 | 3.4028e+38 | 3.4028e+38 | 115.0057 | 115.0057 | 0.0076745 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 0.50002 | 115.0053 | 115.0054 | 0.008574 | 50 | -0.0042716 | 0.01464 | 3.3474e-06 | 6.6129e-10 | 5.0139e-08 | -1.1112e-09 | 74.7004 | 50.7157 | 0.015941 | 3.0286e-06 | 3.9362e-09 | 4.9601e-08 | 2.6199e-10 | 74.7263 | -3.808 | 3.4028e+38 | 3.4028e+38 | 115.0053 | 115.0054 | 0.008574 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 0.50002 | 115.0058 | 115.0058 | 0.0084868 | 49.9999 | -0.0042379 | 0.014723 | 3.4259e-06 | -7.8597e-10 | 4.8136e-08 | -1.5444e-09 | 74.7199 | 105.4119 | 0.015886 | 3.05e-06 | 1.9034e-10 | 4.8471e-08 | -1.301e-09 | 74.7236 | 103.789 | 3.4028e+38 | 3.4028e+38 | 115.0058 | 115.0058 | 0.0084868 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 0.50002 | 115.0057 | 115.0056 | 0.0080381 | 50 | -0.0040046 | 1 | 3.4028e+38 | 0 | 230.0051 | 115.007 | 115.0069 | 0.011184 | -4.8426e-05 | 0.014245 | 3.378e-06 | -7.1766e-10 | 4.817e-08 | -1.4421e-09 | 74.649 | 98.7722 | -0.0046299 | 3.4028e+38 | -47.5618 | 0.019596 | 0 | 0 | 0 | -4.1257e-05 | 0.015077 | 3.0885e-06 | 1.0087e-09 | 4.6768e-08 | -9.1466e-10 | 74.7791 | 83.3009 | 0.0022971 | 3.4028e+38 | 13.031 | 229.7925 | 5.8785e-10 | 0.00070923 | -8.2216e-10 | 3.1882e-05 | 3.4028e+38 | 3.4028e+38 | 115.0059 | 115.0059 | 0.0087781 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 115.0066 | 115.0066 | 0.010946 | -3.4928e-05 | 230.0022 | 0.50002 | 115.0057 | 115.0057 | 0.0084294 | 49.9999 | -0.0042348 | [17 empty] | 23.3347 | 4.6372 | 27.9854 | 27.7124 | 1038739.7649 |  | -0.0021749
74 | 230.0016 | 1 | 0.014423 | 230.0116 | -230.016 | 50 | 89.9963 | 0.01555 | 3.6797e-06 | -5.5944e-09 | 5.6805e-08 | -4.6783e-09 | 74.7116 | 131.7026 | 0.017068 | 3.1259e-06 | -1.0413e-09 | 5.3893e-08 | -1.1667e-09 | 74.7952 | 102.5579 | 3.4028e+38 | 3.4028e+38 | 0.014423 | 230.0116 | -230.016 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 1 | 0.014048 | 230.0112 | -230.0158 | 49.9999 | 89.9964 | 0.01444 | 3.6274e-06 | -3.1984e-09 | 5.2952e-08 | -3.2132e-09 | 74.7696 | 139.8885 | 0.015718 | 3.1603e-06 | -8.7359e-10 | 4.8338e-08 | -9.8425e-10 | 74.7529 | 117.7601 | 3.4028e+38 | 3.4028e+38 | 0.014048 | 230.0112 | -230.0158 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 1 | 0.014575 | 230.0113 | -230.0156 | 50 | 89.9964 | 0.017254 | 3.6935e-06 | -8.0854e-09 | 6.3017e-08 | -5.7762e-09 | 74.7473 | 147.715 | 0.018888 | 3.181e-06 | -1.7174e-09 | 5.9995e-08 | -1.4688e-09 | 74.7693 | 146.7766 | 3.4028e+38 | 3.4028e+38 | 0.014575 | 230.0113 | -230.0156 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 1 | 0.014183 | 230.0109 | -230.0155 | 49.9999 | 89.9965 | 0.016246 | 3.6591e-06 | -5.6984e-09 | 5.894e-08 | -4.6072e-09 | 74.7548 | 144.361 | 0.017308 | 3.1256e-06 | -9.7849e-10 | 5.5407e-08 | -1.7261e-09 | 74.8004 | 132.4068 | 3.4028e+38 | 3.4028e+38 | 0.014183 | 230.0109 | -230.0155 | 3.4028e+38 | 3.4028e+38 | 230.0024 | 1 | 0.014449 | 230.0104 | -230.0148 | 50 | 89.9963 | 0.014301 | 3.6789e-06 | -3.446e-09 | 5.2585e-08 | -1.5233e-09 | 74.7449 | 161.8096 | 0.015442 | 3.1714e-06 | -2.4237e-10 | 4.8949e-08 | -1.3712e-09 | 74.7846 | 151.9738 | 3.4028e+38 | 3.4028e+38 | 0.014449 | 230.0104 | -230.0148 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0023 | 1 | 0.014465 | 230.0107 | -230.0151 | 49.9999 | 89.9963 | 6.5202e-05 | 3.4028e+38 | 0 | 230.0037 | 0.019281 | 230.0119 | -230.0145 | 9.8334e-05 | 0.014241 | 3.6605e-06 | -3.472e-09 | 5.2628e-08 | -1.3257e-09 | 74.7069 | 161.0171 | -0.073228 | 3.4028e+38 | 26.1578 | 0.020375 | 0 | 0 | 0 | -8.6111e-06 | 0.01534 | 3.064e-06 | -2.043e-09 | 4.7473e-08 | -8.87e-10 | 74.691 | 124.1022 | -0.022147 | 3.4028e+38 | -8.2454 | 229.0725 | -5.283e-10 | 0.0007005 | -2.1006e-11 | 9.4927e-06 | 3.4028e+38 | 3.4028e+38 | 0.014593 | 230.0104 | -230.0155 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 0.019621 | 0.019814 | -230.0126 | 0.00012888 | 230.002 | 1 | 0.014336 | 230.0111 | -230.0155 | 50 | 89.9964 | [17 empty] | 24.0082 | 61.8553 | -16733885.5893 | 85.9254 | 105.2135 |  | -0.0038601
75 | 230.0017 | 1 | -0.015457 | 230.01 | 230.01 | 50 | -90.0038 | 0.015743 | 3.6777e-06 | 6.7733e-09 | 5.7898e-08 | 2.2465e-10 | 74.738 | -2.0721 | 0.017121 | 3.1038e-06 | 3.6878e-09 | 5.3127e-08 | -7.5889e-10 | 74.7579 | 10.1579 | 3.4028e+38 | 3.4028e+38 | -0.015457 | 230.01 | 230.01 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 1 | -0.01564 | 230.0099 | 230.0099 | 49.9999 | -90.0041 | 0.01732 | 3.6179e-06 | 6.8553e-09 | 6.2731e-08 | 4.9688e-09 | 74.7158 | -30.9448 | 0.019169 | 3.2126e-06 | 6.9332e-09 | 5.9844e-08 | 4.8394e-09 | 74.7646 | -31.6998 | 3.4028e+38 | 3.4028e+38 | -0.01564 | 230.0099 | 230.0099 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 1 | -0.01647 | 230.0105 | 230.0105 | 50 | -90.0041 | 0.01519 | 3.6154e-06 | 3.8988e-09 | 5.4847e-08 | 3.659e-09 | 74.6802 | -44.0706 | 0.016595 | 3.2198e-06 | 4.6639e-09 | 5.3532e-08 | 3.271e-09 | 74.7111 | -49.8816 | 3.4028e+38 | 3.4028e+38 | -0.01647 | 230.0105 | 230.0105 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | -0.015111 | 230.0103 | 230.0104 | 49.9999 | -90.0038 | 0.016107 | 3.609e-06 | 6.6605e-09 | 5.7907e-08 | -1.0156e-10 | 74.756 | 0.87363 | 0.017222 | 3.0677e-06 | 4.0381e-09 | 5.4923e-08 | 2.0457e-10 | 74.706 | -2.5108 | 3.4028e+38 | 3.4028e+38 | -0.015111 | 230.0103 | 230.0104 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 1 | -0.016289 | 230.0092 | 230.0091 | 50 | -90.004 | 0.016966 | 3.6308e-06 | 6.5442e-09 | 6.1011e-08 | 4.2743e-09 | 74.6947 | -26.4588 | 0.018649 | 3.0782e-06 | 7.1584e-09 | 5.7781e-08 | 4.6227e-09 | 74.7261 | -33.7457 | 3.4028e+38 | 3.4028e+38 | -0.016289 | 230.0092 | 230.0091 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0013 | 1 | -0.016275 | 230.0098 | 230.0097 | 50 | -90.0039 | -6.749e-05 | 3.4028e+38 | 0 | 230.0016 | -0.026054 | 230.0097 | 230.0049 | -0.00084256 | 0.017229 | 3.6479e-06 | 9.0622e-09 | 6.286e-08 | 4.8963e-09 | 74.7929 | -24.3966 | 0.13685 | 3.4028e+38 | 30.9576 | 0.019925 | 0 | 0 | 0 | 1.1245e-05 | 0.019227 | 3.1759e-06 | 6.2267e-09 | 6.1347e-08 | 5.0129e-09 | 74.7774 | -35.1005 | 0.10524 | 3.4028e+38 | 41.0021 | 229.0354 | -2.071e-10 | 0.00074388 | 1.0488e-10 | -1.3859e-05 | 3.4028e+38 | 3.4028e+38 | -0.015958 | 230.0095 | 230.0098 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -0.025465 | -0.025465 | 230.0046 | -0.00080015 | 230.0019 | 1 | -0.015793 | 230.01 | 230.01 | 50 | -90.004 | [17 empty] | 23.1312 | 65.0399 | 416557147.7957 | 87.2823 | 87.2099 |  | -0.0039441
76 | 230.002 | 1 | -0.015282 | 230.0095 | 230.0094 | 50 | -90.0037 | 0.016734 | 3.6319e-06 | 5.9287e-09 | 6.1849e-08 | 6.6718e-10 | 74.7959 | -5.5903 | 0.018537 | 3.2207e-06 | 5.7689e-09 | 5.7688e-08 | 2.2988e-10 | 74.7085 | -2.6948 | 3.4028e+38 | 3.4028e+38 | -0.015282 | 230.0095 | 230.0094 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | -0.015869 | 230.0106 | 230.0104 | 50 | -90.004 | 0.014676 | 3.6559e-06 | 2.8769e-09 | 5.4141e-08 | 3.1014e-09 | 74.7525 | -37.4957 | 0.016104 | 3.2076e-06 | 3.7113e-09 | 5.0187e-08 | 2.31e-09 | 74.7769 | -34.0262 | 3.4028e+38 | 3.4028e+38 | -0.015869 | 230.0106 | 230.0104 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 1 | -0.015103 | 230.0099 | 230.0102 | 50 | -90.0038 | 0.016378 | 3.6647e-06 | 7.0307e-09 | 5.9978e-08 | 3.9096e-10 | 74.7567 | -3.5663 | 0.018044 | 3.0971e-06 | 5.3157e-09 | 5.5798e-08 | 7.8041e-10 | 74.749 | -10.3789 | 3.4028e+38 | 3.4028e+38 | -0.015103 | 230.0099 | 230.0102 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 1 | -0.015169 | 230.01 | 230.0099 | 49.9999 | -90.0038 | 0.017027 | 3.7057e-06 | 8.4094e-09 | 6.31e-08 | 5.9464e-10 | 74.7659 | -4.6941 | 0.01844 | 3.2111e-06 | 4.7207e-09 | 5.9158e-08 | -2.7847e-10 | 74.723 | 2.9447 | 3.4028e+38 | 3.4028e+38 | -0.015169 | 230.01 | 230.0099 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 1 | -0.016127 | 230.01 | 230.0103 | 49.9999 | -90.0039 | 0.016479 | 3.5848e-06 | 6.1287e-09 | 5.9615e-08 | 4.0486e-09 | 74.7271 | -37.8833 | 0.017965 | 3.1846e-06 | 6.1852e-09 | 5.6459e-08 | 5.5334e-09 | 74.6645 | -36.8919 | 3.4028e+38 | 3.4028e+38 | -0.016127 | 230.01 | 230.0103 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0016 | 1 | -0.015717 | 230.0095 | 230.0101 | 49.9999 | -90.004 | -7.0265e-05 | 3.4028e+38 | 0 | 230.0014 | -0.024573 | 230.009 | 230.0047 | -0.00074302 | 0.016791 | 3.5638e-06 | 6.7014e-09 | 6.1171e-08 | 5.1549e-09 | 74.7449 | -37.5689 | 0.13044 | 3.4028e+38 | 30.7419 | 0.020781 | 0 | 0 | 0 | 9.9597e-06 | 0.019112 | 3.2066e-06 | 7.2068e-09 | 6.1349e-08 | 4.1038e-09 | 74.7028 | -32.6232 | 0.10779 | 3.4028e+38 | -12.8437 | 229.6409 | -6.7879e-10 | 0.00072467 | -1.3522e-10 | -1.2267e-05 | 3.4028e+38 | 3.4028e+38 | -0.015555 | 230.0091 | 230.0091 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -0.02617 | -0.02617 | 230.004 | -0.00073075 | 230.002 | 1 | -0.01551 | 230.01 | 230.01 | 49.9999 | -90.0038 | [17 empty] | 22.2508 | 74.305 | 112945563.3935 | 95.7279 | 95.8678 |  | -0.0038067
77 | 230.0018 | 1 | -0.015763 | 230.0096 | 230.0098 | 49.9999 | -90.0038 | 0.016225 | 3.5954e-06 | 4.8092e-09 | 5.8281e-08 | -8.7577e-10 | 74.7499 | 9.8138 | 0.017836 | 3.1407e-06 | 4.8332e-09 | 5.6042e-08 | -4.1786e-10 | 74.6936 | 5.3716 | 3.4028e+38 | 3.4028e+38 | -0.015763 | 230.0096 | 230.0098 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 1 | -0.015242 | 230.0094 | 230.0096 | 49.9999 | -90.0037 | 0.017767 | 3.6319e-06 | 8.9608e-09 | 6.3841e-08 | 2.7391e-09 | 74.7637 | -18.5065 | 0.019301 | 3.1339e-06 | 7.6969e-09 | 6.1534e-08 | 2.7129e-09 | 74.6112 | -21.0282 | 3.4028e+38 | 3.4028e+38 | -0.015242 | 230.0094 | 230.0096 | 3.4028e+38 | 3.4028e+38 | 230.002 | 1 | -0.016242 | 230.0096 | 230.0095 | 49.9999 | -90.004 | 0.016638 | 3.6751e-06 | 5.7367e-09 | 6.1078e-08 | 4.7011e-09 | 74.7508 | -28.7086 | 0.018394 | 3.2453e-06 | 7.0125e-09 | 5.7919e-08 | 4.2767e-09 | 74.8141 | -31.3777 | 3.4028e+38 | 3.4028e+38 | -0.016242 | 230.0096 | 230.0095 | 3.4028e+38 | 3.4028e+38 | 230.002 | 1 | -0.015288 | 230.0102 | 230.0102 | 50 | -90.0039 | 0.014364 | 3.6358e-06 | 3.4657e-09 | 5.2212e-08 | 1.9374e-10 | 74.7756 | -3.8851 | 0.015279 | 3.2068e-06 | 1.9067e-09 | 4.8999e-08 | 1.4995e-10 | 74.7416 | -4.0815 | 3.4028e+38 | 3.4028e+38 | -0.015288 | 230.0102 | 230.0102 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | -0.015562 | 230.0095 | 230.0092 | 50 | -90.0039 | 0.017487 | 3.6199e-06 | 7.4832e-09 | 6.1088e-08 | 4.9507e-09 | 74.7077 | -36.1733 | 0.0192 | 3.1032e-06 | 7.4686e-09 | 6.1162e-08 | 5.2932e-09 | 74.7356 | -38.7364 | 3.4028e+38 | 3.4028e+38 | -0.015562 | 230.0095 | 230.0092 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0013 | 1 | -0.016161 | 230.0094 | 230.0097 | 49.9999 | -90.004 | -6.878e-05 | 3.4028e+38 | 0 | 230.0007 | -0.024517 | 230.0087 | 230.0042 | -0.00053581 | 0.016376 | 3.592e-06 | 4.7389e-09 | 5.827e-08 | 4.3057e-09 | 74.6965 | -42.5669 | 0.076444 | 3.4028e+38 | 3.5217 | 0.020763 | 0 | 0 | 0 | 2.1142e-05 | 0.01731 | 3.167e-06 | 5.7151e-09 | 5.5116e-08 | 3.7095e-09 | 74.8379 | -34.6498 | 0.086631 | 3.4028e+38 | -24.9793 | 229.7326 | -2.6347e-10 | 0.00074276 | 5.2067e-11 | -5.7298e-06 | 3.4028e+38 | 3.4028e+38 | -0.015758 | 230.0099 | 230.0096 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -0.022976 | -0.022976 | 230.0046 | -0.00034986 | 230.0018 | 1 | -0.015619 | 230.0097 | 230.0097 | 49.9999 | -90.0039 | [17 empty] | 21.7596 | 82.1499 | -57634731.0734 | 103.5544 | 103.6012 |  | -0.0039399
78 | 230.0018 | 1 | 199.1874 | 230.01 | 115.0186 | 49.9999 | -30.0038 | 0.017706 | 3.6883e-06 | 3.0832e-09 | 6.5301e-08 | -6.2466e-09 | 74.7144 | 69.8035 | 0.01939 | 3.2177e-06 | 2.6553e-09 | 6.2321e-08 | -2.2107e-09 | 74.7378 | 33.1609 | 3.4028e+38 | 3.4028e+38 | 199.1874 | 230.01 | 115.0186 | 3.4028e+38 | 3.4028e+38 | 230.0013 | 1 | 199.1862 | 230.009 | 115.0175 | 50 | -30.0036 | 0.016954 | 3.6024e-06 | 5.8915e-09 | 6.0831e-08 | -1.3953e-09 | 74.7527 | 24.541 | 0.018566 | 3.1466e-06 | 5.221e-09 | 5.9617e-08 | -5.6414e-11 | 74.8139 | -0.65049 | 3.4028e+38 | 3.4028e+38 | 199.1862 | 230.009 | 115.0175 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 1 | 199.1869 | 230.0099 | 115.0182 | 49.9999 | -30.0039 | 0.016161 | 3.6185e-06 | 4.4541e-09 | 5.8443e-08 | -1.7911e-09 | 74.7716 | 23.8047 | 0.01788 | 3.1254e-06 | 5.0771e-09 | 5.592e-08 | 4.1527e-10 | 74.5011 | -5.8026 | 3.4028e+38 | 3.4028e+38 | 199.1869 | 230.0099 | 115.0182 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | 199.1869 | 230.0097 | 115.0186 | 49.9999 | -30.0039 | 0.014367 | 3.5943e-06 | 1.3663e-09 | 5.1635e-08 | -1.031e-09 | 74.6813 | 17.7238 | 0.015708 | 3.1108e-06 | 1.6624e-09 | 4.8666e-08 | -1.0101e-09 | 74.8178 | 27.5391 | 3.4028e+38 | 3.4028e+38 | 199.1869 | 230.0097 | 115.0186 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 1 | 199.1865 | 230.0091 | 115.0181 | 49.9999 | -30.0037 | 0.01747 | 3.638e-06 | 4.2589e-09 | 6.1329e-08 | -3.282e-09 | 74.6607 | 32.6362 | 0.019259 | 3.1538e-06 | 5.3093e-09 | 5.8335e-08 | -2.8382e-10 | 74.7207 | -1.2806 | 3.4028e+38 | 3.4028e+38 | 199.1865 | 230.0091 | 115.0181 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0023 | 1 | 199.1869 | 230.01 | 115.0189 | 49.9999 | -30.0039 | 0.86599 | 3.4028e+38 | 0 | 230.0018 | 199.1844 | 230.0099 | 115.0212 | -0.0001037 | 0.014519 | 3.6257e-06 | 4.8996e-11 | 5.2138e-08 | -2.2887e-09 | 74.7611 | 68.2352 | 0.022711 | 3.4028e+38 | -80.5851 | 0.02 | 0 | 0 | 0 | -7.6607e-05 | 0.01681 | 3.095e-06 | 1.202e-09 | 5.595e-08 | -2.9316e-09 | 74.7346 | 97.0246 | 0.038987 | 3.4028e+38 | -42.1218 | 229.7151 | -7.8797e-10 | 0.00073491 | 5.9796e-10 | 3.7556e-05 | 3.4028e+38 | 3.4028e+38 | 199.1869 | 230.0098 | 115.0183 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 199.1839 | 199.1839 | 115.024 | -0.00061208 | 230.0016 | 1 | 199.1868 | 230.0095 | 115.0182 | 49.9999 | -30.0038 | [17 empty] | 20.8695 | 10.3249 | -6.184 | 30.3669 | 146.2361 |  | -0.0037684
79 | 230.0017 | 1 | 199.1866 | 230.0094 | 115.018 | 50 | -30.0038 | 0.017028 | 3.6346e-06 | 7.8927e-10 | 6.3674e-08 | -7.7237e-09 | 74.7115 | 89.1093 | 0.018901 | 3.1398e-06 | 2.1972e-09 | 6.0684e-08 | -3.1018e-09 | 74.8585 | 34.8458 | 3.4028e+38 | 3.4028e+38 | 199.1866 | 230.0094 | 115.0179 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 1 | 199.1866 | 230.0102 | 115.019 | 49.9999 | -30.004 | 0.014232 | 3.5615e-06 | 3.2248e-09 | 5.0693e-08 | -1.5259e-09 | 74.7454 | 36.8087 | 0.015412 | 3.1593e-06 | 2.4849e-09 | 5.0367e-08 | -6.1318e-10 | 74.8371 | 12.3959 | 3.4028e+38 | 3.4028e+38 | 199.1866 | 230.0102 | 115.019 | 3.4028e+38 | 3.4028e+38 | 230.002 | 1 | 199.1867 | 230.0099 | 115.0184 | 49.9999 | -30.0038 | 0.0144 | 3.6279e-06 | 2.5374e-09 | 5.3521e-08 | -1.3496e-09 | 74.7456 | 28.3607 | 0.016018 | 3.1382e-06 | 2.4146e-09 | 4.9703e-08 | 4.6214e-10 | 74.7778 | -4.7934 | 3.4028e+38 | 3.4028e+38 | 199.1867 | 230.0099 | 115.0184 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | 199.1868 | 230.0103 | 115.0186 | 50 | -30.0038 | 0.014316 | 3.6419e-06 | 2.0998e-09 | 5.2109e-08 | -1.4452e-09 | 74.6468 | 46.3981 | 0.01533 | 3.1615e-06 | 1.3274e-09 | 4.8378e-08 | -8.6674e-10 | 74.7027 | 58.2484 | 3.4028e+38 | 3.4028e+38 | 199.1868 | 230.0103 | 115.0186 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 1 | 199.1867 | 230.0095 | 115.018 | 50 | -30.0038 | 0.016912 | 3.6051e-06 | 5.3357e-09 | 5.9069e-08 | -1.7995e-09 | 74.7254 | 21.7193 | 0.018525 | 3.2369e-06 | 5.4952e-09 | 5.9222e-08 | 1.7437e-09 | 74.7851 | -12.0298 | 3.4028e+38 | 3.4028e+38 | 199.1867 | 230.0095 | 115.018 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0017 | 1 | 199.1869 | 230.01 | 115.0185 | 50 | -30.0039 | 0.86599 | 3.4028e+38 | 0 | 230.0005 | 199.1821 | 230.0077 | 115.0207 | -0.00045695 | 0.01509 | 3.6521e-06 | 3.8226e-09 | 5.4617e-08 | -5.8083e-10 | 74.7737 | 10.8699 | 0.041829 | 3.4028e+38 | -53.833 | 0.019765 | 0 | 0 | 0 | -6.9709e-05 | 0.016104 | 3.2248e-06 | 1.7672e-09 | 4.9802e-08 | -2.0162e-10 | 74.7183 | 4.5179 | 0.051235 | 3.4028e+38 | -14.8574 | 229.0789 | 7.4653e-10 | 0.00071335 | 1.0146e-09 | 4.8666e-05 | 3.4028e+38 | 3.4028e+38 | 199.187 | 230.0104 | 115.0183 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 199.1846 | 199.1846 | 115.0214 | -0.00023824 | 230.0018 | 1 | 199.1867 | 230.0098 | 115.0184 | 49.9999 | -30.0038 | [17 empty] | 21.5715 | 9.1186 | -4.4226 | 29.7364 | 133.5569 |  | -0.0034072
80 | 230.0019 | 1 | 199.1872 | 230.0104 | 115.0193 | 49.9999 | -30.004 | 0.015712 | 3.6235e-06 | 3.1959e-10 | 5.6962e-08 | -5.102e-09 | 74.7493 | 82.4938 | 0.016807 | 3.1043e-06 | 6.8157e-10 | 5.3518e-08 | -3.054e-09 | 74.746 | 111.3704 | 3.4028e+38 | 3.4028e+38 | 199.1872 | 230.0104 | 115.0193 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 1 | 199.1864 | 230.0095 | 115.0181 | 49.9999 | -30.004 | 0.016935 | 3.6171e-06 | 4.9153e-09 | 6.16e-08 | -2.1614e-09 | 74.6941 | 25.7555 | 0.018673 | 3.2122e-06 | 6.5593e-09 | 6.0463e-08 | 1.4579e-09 | 74.6302 | -3.6088 | 3.4028e+38 | 3.4028e+38 | 199.1864 | 230.0095 | 115.0181 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 1 | 199.1864 | 230.0098 | 115.019 | 49.9999 | -30.004 | 0.015196 | 3.6011e-06 | 3.5631e-09 | 5.4706e-08 | -1.3098e-10 | 74.7185 | 2.4354 | 0.016296 | 3.1204e-06 | 3.145e-09 | 5.3004e-08 | 6.0827e-10 | 74.8093 | -11.4692 | 3.4028e+38 | 3.4028e+38 | 199.1864 | 230.0098 | 115.019 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 1 | 199.1875 | 230.0105 | 115.0185 | 49.9999 | -30.0038 | 0.015886 | 3.6788e-06 | 1.8182e-09 | 5.792e-08 | -5.3434e-09 | 74.7214 | 77.9671 | 0.017221 | 3.1029e-06 | 1.3121e-09 | 5.3377e-08 | -3.2011e-09 | 74.6989 | 77.2754 | 3.4028e+38 | 3.4028e+38 | 199.1875 | 230.0105 | 115.0185 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | 199.1868 | 230.0096 | 115.018 | 49.9999 | -30.0038 | 0.016673 | 3.7188e-06 | 5.622e-10 | 6.1959e-08 | -7.1008e-09 | 74.7513 | 86.8563 | 0.018388 | 3.17e-06 | 9.082e-10 | 5.6323e-08 | -2.9865e-09 | 74.7226 | 55.6511 | 3.4028e+38 | 3.4028e+38 | 199.1868 | 230.0096 | 115.018 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0012 | 1 | 199.1865 | 230.0093 | 115.0182 | 49.9999 | -30.0038 | 0.86599 | 3.4028e+38 | 0 | 230.0068 | 199.1877 | 230.0146 | 115.0275 | -6.9793e-05 | 0.017148 | 3.6142e-06 | 1.1339e-09 | 6.2302e-08 | -5.5542e-09 | 74.7688 | 69.6911 | 0.029676 | 3.4028e+38 | 36.6663 | 0.020635 | 0 | 0 | 0 | -8.6688e-05 | 0.019066 | 3.1319e-06 | 2.279e-09 | 5.9691e-08 | -3.0135e-09 | 74.7397 | 64.0991 | 0.023721 | 3.4028e+38 | 3.6894 | 229.7173 | -5.5487e-10 | 0.0007091 | 9.7547e-10 | 3.2592e-05 | 3.4028e+38 | 3.4028e+38 | 199.1874 | 230.0104 | 115.0184 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 199.1862 | 199.1862 | 115.027 | -0.00019608 | 230.0018 | 1 | 199.1869 | 230.0099 | 115.0186 | 49.9999 | -30.0039 | [17 empty] | 21.7646 | 6.8471 | -9.0312 | 27.3114 | 140.2568 |  | -0.0037405
81 | 230.0019 | 1 | -138.0212 | 230.0093 | 183.998 | 49.9999 | -126.8744 | 0.017819 | 3.6137e-06 | 7.4392e-09 | 6.404e-08 | 7.3556e-09 | 74.7644 | -45.7541 | 0.019586 | 3.2123e-06 | 5.8668e-09 | 6.2895e-08 | 5.3774e-09 | 74.8879 | -43.6638 | 3.4028e+38 | 3.4028e+38 | -138.0212 | 230.0093 | 183.998 | 3.4028e+38 | 3.4028e+38 | 230.002 | 1 | -138.0213 | 230.0099 | 183.9986 | 49.9999 | -126.8743 | 0.017003 | 3.649e-06 | 6.9949e-09 | 6.2205e-08 | 3.8792e-09 | 74.7811 | -29.0846 | 0.018857 | 3.206e-06 | 5.3901e-09 | 6.0048e-08 | 2.7776e-09 | 74.7665 | -32.317 | 3.4028e+38 | 3.4028e+38 | -138.0213 | 230.0099 | 183.9986 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 1 | -138.0204 | 230.0092 | 183.9986 | 49.9999 | -126.8742 | 0.017324 | 3.655e-06 | 7.8097e-09 | 6.3383e-08 | 5.6978e-09 | 74.7712 | -31.5912 | 0.018938 | 3.1167e-06 | 7.1543e-09 | 6.1527e-08 | 4.1284e-09 | 74.8026 | -29.2904 | 3.4028e+38 | 3.4028e+38 | -138.0204 | 230.0092 | 183.9986 | 3.4028e+38 | 3.4028e+38 | 230.002 | 1 | -138.0211 | 230.0103 | 183.9992 | 50 | -126.8743 | 0.014538 | 3.6318e-06 | 3.5317e-09 | 5.2333e-08 | 4.0034e-09 | 74.7298 | -55.3977 | 0.015634 | 3.1191e-06 | 2.0243e-09 | 4.8772e-08 | 2.1386e-09 | 74.7865 | -42.6744 | 3.4028e+38 | 3.4028e+38 | -138.0211 | 230.0103 | 183.9992 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 1 | -138.0211 | 230.0099 | 183.9992 | 49.9999 | -126.8744 | 0.015692 | 3.5828e-06 | 4.4031e-09 | 5.6408e-08 | 5.6465e-09 | 74.7795 | -62.8757 | 0.017355 | 3.2117e-06 | 3.7709e-09 | 5.3702e-08 | 4.9545e-09 | 74.6519 | -49.719 | 3.4028e+38 | 3.4028e+38 | -138.0211 | 230.0099 | 183.9992 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0017 | 1 | -138.021 | 230.01 | 183.999 | 50 | -126.8743 | -0.60007 | 3.4028e+38 | 0 | 230.0007 | -138.0252 | 230.0091 | 183.9934 | 4.3966e-05 | 0.015021 | 3.6115e-06 | 2.4566e-09 | 5.4221e-08 | 3.9752e-09 | 74.7868 | -66.7281 | 0.035129 | 3.4028e+38 | 40.6675 | 0.019942 | 0 | 0 | 0 | 8.3906e-05 | 0.015491 | 3.1323e-06 | 2.2173e-09 | 5.1295e-08 | 2.1297e-09 | 74.707 | -43.6784 | 0.052534 | 3.4028e+38 | -1.67 | 229.2055 | 8.4979e-10 | 0.00073327 | 6.8758e-10 | -3.2612e-05 | 3.4028e+38 | 3.4028e+38 | -138.0215 | 230.0104 | 183.9992 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -138.0253 | -138.0249 | 183.995 | 8.4741e-05 | 230.002 | 1 | -138.021 | 230.0097 | 183.9987 | 49.9999 | -126.8743 | [17 empty] | 22.4273 | 63.5712 | 143.9166 | 85.6631 | 65.3496 |  | -0.0021964
82 | 230.0016 | 1 | -138.0209 | 230.0096 | 183.9987 | 50 | -126.8743 | 0.01505 | 3.6036e-06 | 6.5156e-09 | 5.5532e-08 | 3.0678e-09 | 74.7212 | -30.7203 | 0.016183 | 3.1987e-06 | 3.5611e-09 | 5.2818e-08 | 1.7804e-09 | 74.7404 | -11.5908 | 3.4028e+38 | 3.4028e+38 | -138.0209 | 230.0096 | 183.9987 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 1 | -138.0211 | 230.0101 | 183.9988 | 50 | -126.8744 | 0.014237 | 3.6301e-06 | 2.7817e-09 | 5.1731e-08 | 3.3802e-09 | 74.8303 | -57.5787 | 0.015666 | 3.0644e-06 | 1.7156e-09 | 5.0317e-08 | 2.3991e-09 | 74.7222 | -26.7634 | 3.4028e+38 | 3.4028e+38 | -138.0211 | 230.0101 | 183.9988 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | -138.0215 | 230.01 | 183.9985 | 50 | -126.8744 | 0.014657 | 3.6959e-06 | 2.1086e-09 | 5.4586e-08 | 5.0561e-09 | 74.7807 | -67.3616 | 0.0163 | 3.1436e-06 | 2.3397e-09 | 5.0155e-08 | 1.6769e-09 | 74.7491 | -35.8714 | 3.4028e+38 | 3.4028e+38 | -138.0215 | 230.01 | 183.9985 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 1 | -138.0216 | 230.0101 | 183.9985 | 49.9999 | -126.8744 | 0.014928 | 3.5968e-06 | 1.8673e-09 | 5.3919e-08 | 5.4432e-09 | 74.6728 | -69.6794 | 0.016303 | 3.2165e-06 | 2.5417e-09 | 5.0085e-08 | 3.3299e-09 | 74.7732 | -49.3176 | 3.4028e+38 | 3.4028e+38 | -138.0216 | 230.0101 | 183.9985 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 1 | -138.0217 | 230.0096 | 183.9979 | 49.9999 | -126.8743 | 0.017152 | 3.649e-06 | 4.7988e-09 | 6.0134e-08 | 7.3519e-09 | 74.6574 | -58.3323 | 0.018675 | 3.2272e-06 | 5.0979e-09 | 6.0189e-08 | 5.8366e-09 | 74.7626 | -50.5247 | 3.4028e+38 | 3.4028e+38 | -138.0217 | 230.0096 | 183.9979 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0026 | 1 | -138.0211 | 230.0095 | 183.9981 | 49.9999 | -126.8743 | -0.60007 | 3.4028e+38 | 0 | 230.0003 | -138.0271 | 230.0084 | 183.9904 | -0.00017339 | 0.016172 | 3.6106e-06 | 2.0525e-09 | 5.8277e-08 | 5.1879e-09 | 74.6475 | -71.2525 | 0.031966 | 3.4028e+38 | 5.4169 | 0.019564 | 0 | 0 | 0 | 8.0025e-05 | 0.016743 | 3.1655e-06 | 5.1526e-09 | 5.2968e-08 | 4.8554e-09 | 74.7983 | -55.3633 | 0.069114 | 3.4028e+38 | -35.0064 | 229.0192 | -4.2036e-10 | 0.00072499 | 5.4073e-10 | -3.6617e-05 | 3.4028e+38 | 3.4028e+38 | -138.021 | 230.0099 | 183.9989 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -138.0249 | -138.0249 | 183.9929 | 6.938e-05 | 230.0017 | 1 | -138.0214 | 230.0099 | 183.9985 | 50 | -126.8743 | [17 empty] | 20.8902 | 70.5267 | 142.8534 | 90.7302 | 72.9661 |  | -0.0018895
83 | 230.0019 | 1 | -138.0212 | 230.0101 | 183.9986 | 50 | -126.8743 | 0.014559 | 3.6331e-06 | 2.6913e-09 | 5.269e-08 | 4.1505e-09 | 74.7627 | -57.0396 | 0.015554 | 3.0841e-06 | 2.6017e-09 | 4.8002e-08 | 2.5029e-09 | 74.6794 | -33.8631 | 3.4028e+38 | 3.4028e+38 | -138.0212 | 230.0101 | 183.9986 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 1 | -138.0218 | 230.0102 | 183.9986 | 49.9999 | -126.8744 | 0.017659 | 3.5978e-06 | 4.5432e-09 | 6.2242e-08 | 6.5217e-09 | 74.76 | -58.9111 | 0.019235 | 3.235e-06 | 6.9184e-09 | 5.8353e-08 | 7.7054e-09 | 74.8096 | -53.248 | 3.4028e+38 | 3.4028e+38 | -138.0218 | 230.0102 | 183.9986 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 1 | -138.0215 | 230.0094 | 183.9981 | 49.9999 | -126.8744 | 0.015428 | 3.5794e-06 | 1.4265e-09 | 5.7168e-08 | 5.456e-09 | 74.7557 | -66.1407 | 0.016802 | 3.159e-06 | 3.3268e-09 | 5.3045e-08 | 3.8408e-09 | 74.7533 | -43.4855 | 3.4028e+38 | 3.4028e+38 | -138.0215 | 230.0094 | 183.9981 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 1 | -138.0206 | 230.0092 | 183.9981 | 49.9999 | -126.8744 | 0.015612 | 3.6845e-06 | 7.1666e-09 | 5.7559e-08 | 3.3848e-09 | 74.7116 | -29.6772 | 0.017079 | 3.1975e-06 | 3.7555e-09 | 5.4494e-08 | 1.5849e-09 | 74.7429 | -25.9568 | 3.4028e+38 | 3.4028e+38 | -138.0206 | 230.0092 | 183.9981 | 3.4028e+38 | 3.4028e+38 | 230.002 | 1 | -138.0208 | 230.0094 | 183.9987 | 49.9999 | -126.8743 | 0.015184 | 3.6639e-06 | 6.601e-09 | 5.5385e-08 | 3.366e-09 | 74.77 | -28.9867 | 0.016454 | 3.1491e-06 | 4.2731e-09 | 5.1835e-08 | 1.5128e-09 | 74.6987 | -21.1317 | 3.4028e+38 | 3.4028e+38 | -138.0208 | 230.0094 | 183.9987 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0014 | 1 | -138.021 | 230.0097 | 183.9987 | 49.9999 | -126.8744 | -0.60007 | 3.4028e+38 | 0 | 230.0068 | -138.029 | 230.0145 | 183.9944 | -0.00059394 | 0.016249 | 3.6005e-06 | 7.6248e-09 | 5.8398e-08 | 3.4566e-09 | 74.7267 | -27.2608 | 0.11628 | 3.4028e+38 | 28.8719 | 0.020146 | 0 | 0 | 0 | 7.091e-05 | 0.017796 | 3.1818e-06 | 5.9093e-09 | 5.6904e-08 | 2.4417e-09 | 74.7561 | -33.4347 | 0.06521 | 3.4028e+38 | 3.0021 | 229.0556 | -7.5261e-10 | 0.00072858 | -4.8839e-10 | -4.29e-05 | 3.4028e+38 | 3.4028e+38 | -138.0208 | 230.0096 | 183.9985 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -138.0304 | -138.0304 | 183.9938 | -0.00063265 | 230.0016 | 1 | -138.0212 | 230.0097 | 183.9984 | 49.9999 | -126.8744 | [17 empty] | 21.6609 | 71.0973 | 143.4105 | 92.6718 | 76.0226 |  | -0.0018721
84 | 230.0015 | 6.0003 | 1380.0875 | 1380.088 | 0.064145 | 50 | -0.0031723 | 0.016115 | 3.9123e-06 | 1.1601e-08 | 6.364e-08 | -3.5851e-10 | 74.6184 | 1.8608 | 0.017487 | 3.2392e-06 | 7.0851e-09 | 5.6623e-08 | 7.8809e-09 | 74.6111 | -51.773 | 3.4028e+38 | 3.4028e+38 | 1380.0875 | 1380.088 | 0.064145 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 6.0003 | 1380.0928 | 1380.0902 | 0.07127 | 49.9999 | -0.0029588 | 0.015304 | 3.921e-06 | 6.0736e-09 | 5.7257e-08 | -4.3088e-10 | 74.6466 | 1.7727 | 0.01614 | 3.3042e-06 | 4.4707e-09 | 5.0581e-08 | 4.0537e-09 | 74.8232 | -42.1999 | 3.4028e+38 | 3.4028e+38 | 1380.0928 | 1380.0902 | 0.07127 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 6.0003 | 1380.0868 | 1380.0873 | 0.075221 | 49.9999 | -0.0028576 | 0.017493 | 3.8972e-06 | 1.3698e-08 | 6.8159e-08 | -2.2943e-09 | 74.6282 | 8.7979 | 0.019457 | 3.3277e-06 | 9.278e-09 | 6.359e-08 | 8.6916e-09 | 74.7569 | -41.7906 | 3.4028e+38 | 3.4028e+38 | 1380.0868 | 1380.0873 | 0.075221 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 6.0003 | 1380.0851 | 1380.0879 | 0.066907 | 49.9999 | -0.0027777 | 0.01675 | 3.9224e-06 | 1.2037e-08 | 6.4242e-08 | -7.4172e-10 | 74.6653 | -2.1164 | 0.018301 | 3.2424e-06 | 8.9728e-09 | 5.8726e-08 | 8.0996e-09 | 74.7605 | -41.5876 | 3.4028e+38 | 3.4028e+38 | 1380.0851 | 1380.0879 | 0.066907 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 6.0003 | 1380.0876 | 1380.0883 | 0.072633 | 50 | -0.0029025 | 0.016676 | 4.02e-06 | 5.5862e-09 | 6.7168e-08 | -1.1976e-08 | 74.6375 | 69.6745 | 0.018278 | 3.229e-06 | 8.6976e-09 | 5.9009e-08 | 4.1761e-09 | 74.7783 | -22.4606 | 3.4028e+38 | 3.4028e+38 | 1380.0876 | 1380.0883 | 0.072633 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 6.0003 | 1380.0898 | 1380.0901 | 0.067933 | 50 | -0.0030413 | 1 | 3.4028e+38 | 0 | 230.007 | 1380.1191 | 1380.1199 | 0.12547 | 0.0019174 | 0.016844 | 4.0116e-06 | 5.6585e-09 | 6.9301e-08 | -1.1182e-08 | 74.7426 | 65.949 | 0.068926 | 3.4028e+38 | -44.3706 | 0.02133 | 0 | 0 | 0 | -0.00014875 | 0.019136 | 3.3049e-06 | 1.0647e-08 | 6.3274e-08 | 4.5794e-09 | 74.7656 | -21.5409 | 0.17463 | 3.4028e+38 | 54.451 | 229.1362 | -1.0098e-09 | 0.00075626 | -1.2976e-09 | -5.2323e-05 | 3.4028e+38 | 3.4028e+38 | 1380.0923 | 1380.0885 | 0.071334 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1380.1097 | 1380.1097 | 0.13659 | 0.0021346 | 230.0019 | 6.0003 | 1380.088 | 1380.0883 | 0.070035 | 49.9999 | -0.0029338 | [17 empty] | 23.0768 | 1.0934 | 23.5192 | 23.7899 | 384743.6592 |  | -0.00083402
85 | 230.0027 | 6.0003 | 1380.0895 | 1380.0898 | 0.072941 | 50 | -0.0030282 | 0.01427 | 3.9826e-06 | 7.011e-09 | 5.7362e-08 | -9.0922e-10 | 74.6028 | 2.9733 | 0.015438 | 3.2389e-06 | 4.6732e-09 | 5.1723e-08 | 3.9803e-09 | 74.7265 | -36.1783 | 3.4028e+38 | 3.4028e+38 | 1380.0895 | 1380.0898 | 0.072941 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 6.0003 | 1380.0803 | 1380.0811 | 0.065658 | 49.9999 | -0.0029078 | 0.017712 | 4.0001e-06 | 1.0797e-08 | 6.9905e-08 | -1.0408e-08 | 74.6402 | 44.6526 | 0.019838 | 3.1977e-06 | 1.3276e-08 | 6.3432e-08 | 6.987e-09 | 74.7981 | -33.0171 | 3.4028e+38 | 3.4028e+38 | 1380.0803 | 1380.0811 | 0.065658 | 3.4028e+38 | 3.4028e+38 | 230.0026 | 6.0003 | 1380.0852 | 1380.0854 | 0.074147 | 50 | -0.0030783 | 0.014454 | 3.9344e-06 | 5.2208e-09 | 5.5716e-08 | -2.5872e-09 | 74.5765 | 29.3731 | 0.015333 | 3.3339e-06 | 2.7244e-09 | 4.9921e-08 | 1.3742e-09 | 74.665 | -30.177 | 3.4028e+38 | 3.4028e+38 | 1380.0852 | 1380.0854 | 0.074147 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 6.0003 | 1380.0891 | 1380.0896 | 0.071382 | 49.9999 | -0.0029189 | 0.017498 | 4.0415e-06 | 7.3523e-09 | 6.998e-08 | -1.0181e-08 | 74.7253 | 45.0972 | 0.019222 | 3.2899e-06 | 9.9656e-09 | 6.3406e-08 | 6.3117e-09 | 74.8789 | -27.8622 | 3.4028e+38 | 3.4028e+38 | 1380.0891 | 1380.0896 | 0.071382 | 3.4028e+38 | 3.4028e+38 | 230.002 | 6.0003 | 1380.0845 | 1380.0836 | 0.069022 | 50 | -0.0028655 | 0.014302 | 3.9848e-06 | 3.2947e-09 | 5.6611e-08 | -3.8748e-09 | 74.7275 | 40.8532 | 0.015169 | 3.2097e-06 | 3.6478e-09 | 5.0574e-08 | 9.8436e-10 | 74.751 | -25.5972 | 3.4028e+38 | 3.4028e+38 | 1380.0845 | 1380.0836 | 0.069022 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0018 | 6.0003 | 1380.0768 | 1380.0771 | 0.075607 | 50 | -0.0029111 | 1 | 3.4028e+38 | 0 | 230.0018 | 1380.0808 | 1380.0812 | 0.094807 | -0.0022099 | 0.014623 | 3.9034e-06 | 7.886e-09 | 5.7045e-08 | 1.1673e-09 | 74.7488 | -7.7164 | 0.14683 | 3.4028e+38 | -15.9705 | 0.019811 | 0 | 0 | 0 | -0.00012714 | 0.01681 | 3.3815e-06 | 6.9533e-09 | 5.6858e-08 | 8.242e-09 | 74.7402 | -51.0985 | 0.12141 | 3.4028e+38 | 10.9376 | 229.6341 | 3.4888e-10 | 0.00074116 | 1.2652e-09 | -4.9085e-05 | 3.4028e+38 | 3.4028e+38 | 1380.0806 | 1380.0811 | 0.069434 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1380.0701 | 1380.0701 | 0.10984 | -0.0015882 | 230.0023 | 6.0003 | 1380.0857 | 1380.0859 | 0.07063 | 50 | -0.0029598 | [17 empty] | 24.4315 | 5.4588 | 29.1322 | 29.2789 | 462615.6869 |  | -0.00095488
86 | 230.0016 | 6.0003 | 1380.0791 | 1380.0765 | 0.069411 | 49.9999 | -0.002764 | 0.015218 | 3.9582e-06 | 8.977e-09 | 6.0578e-08 | 1.2836e-09 | 74.8048 | -6.9317 | 0.016599 | 3.2749e-06 | 6.5503e-09 | 5.3916e-08 | 7.7721e-09 | 74.7677 | -54.7153 | 3.4028e+38 | 3.4028e+38 | 1380.0791 | 1380.0765 | 0.069411 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 6.0003 | 1380.084 | 1380.0878 | 0.069939 | 49.9999 | -0.0029036 | 0.014587 | 4.0004e-06 | 4.2751e-09 | 5.8634e-08 | -5.2906e-09 | 74.7433 | 53.5832 | 0.015242 | 3.2175e-06 | 4.3135e-09 | 4.9015e-08 | 1.5929e-09 | 74.806 | -15.9002 | 3.4028e+38 | 3.4028e+38 | 1380.084 | 1380.0878 | 0.069939 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 6.0003 | 1380.0801 | 1380.0807 | 0.072746 | 49.9999 | -0.0027538 | 0.017509 | 4.0215e-06 | 7.1711e-09 | 7.027e-08 | -1.0289e-08 | 74.736 | 52.9618 | 0.019153 | 3.2441e-06 | 1.3844e-08 | 6.2849e-08 | 6.879e-09 | 74.6169 | -27.3853 | 3.4028e+38 | 3.4028e+38 | 1380.0801 | 1380.0807 | 0.072746 | 3.4028e+38 | 3.4028e+38 | 230.002 | 6.0003 | 1380.0868 | 1380.0873 | 0.069631 | 50 | -0.003065 | 0.014293 | 4.0009e-06 | 3.6794e-09 | 5.7492e-08 | -3.9872e-09 | 74.6222 | 55.5194 | 0.015345 | 3.3113e-06 | 3.8059e-09 | 4.9056e-08 | 1.8835e-09 | 74.739 | -17.9122 | 3.4028e+38 | 3.4028e+38 | 1380.0868 | 1380.0873 | 0.069631 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 6.0003 | 1380.085 | 1380.083 | 0.0739 | 49.9999 | -0.003068 | 0.01603 | 4.0028e-06 | 3.7602e-09 | 6.3921e-08 | -1.1388e-08 | 74.5568 | 68.7265 | 0.017653 | 3.278e-06 | 9.7937e-09 | 5.7931e-08 | 2.8029e-09 | 74.7186 | -18.7835 | 3.4028e+38 | 3.4028e+38 | 1380.085 | 1380.083 | 0.0739 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0022 | 6.0003 | 1380.0881 | 1380.0887 | 0.070361 | 50 | -0.0028516 | 1 | 3.4028e+38 | 0 | 230.0069 | 1380.1132 | 1380.1136 | 0.10681 | 0.0013408 | 0.015487 | 3.9746e-06 | 1.7127e-09 | 6.0564e-08 | -1.0129e-08 | 74.7059 | 80.9865 | 0.024609 | 3.4028e+38 | 15.1049 | 0.019743 | 0 | 0 | 0 | -0.00013488 | 0.01681 | 3.2431e-06 | 6.5363e-09 | 5.4638e-08 | 1.5377e-09 | 74.7397 | -12.3228 | 0.12943 | 3.4028e+38 | -67.2747 | 229.6471 | 5.7169e-10 | 0.00075163 | 4.0966e-10 | -4.5208e-05 | 3.4028e+38 | 3.4028e+38 | 1380.0829 | 1380.0835 | 0.072097 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1380.1168 | 1380.1141 | 0.12584 | 0.0018921 | 230.0018 | 6.0003 | 1380.083 | 1380.0831 | 0.071125 | 49.9999 | -0.0029109 | [17 empty] | 22.5197 | 5.0182 | 27.724 | 27.7818 | 513010.7541 |  | -0.0009592
87 | 230.0017 | 6.0003 | 689.9816 | 1380.0798 | 1195.2196 | 49.9999 | -60.0028 | 0.017481 | 3.934e-06 | 1.5275e-08 | 6.8756e-08 | 1.1744e-08 | 74.6996 | -37.7428 | 0.019495 | 3.2647e-06 | 4.9337e-09 | 6.3643e-08 | 1.0429e-08 | 74.8242 | -67.9026 | 3.4028e+38 | 3.4028e+38 | 689.9816 | 1380.0798 | 1195.2196 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 6.0003 | 689.9863 | 1380.0837 | 1195.2203 | 50 | -60.0027 | 0.016951 | 3.9922e-06 | 1.8217e-08 | 6.7162e-08 | 6.5454e-09 | 74.4923 | -17.2878 | 0.018413 | 3.1609e-06 | 5.9202e-09 | 5.8821e-08 | 8.2362e-09 | 74.8519 | -56.5295 | 3.4028e+38 | 3.4028e+38 | 689.9863 | 1380.0837 | 1195.2203 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 6.0003 | 689.9838 | 1380.0837 | 1195.2212 | 50 | -60.0029 | 0.015388 | 3.967e-06 | 1.0361e-08 | 6.1044e-08 | 1.1698e-08 | 74.6716 | -53.6246 | 0.016598 | 3.3075e-06 | 1.8367e-09 | 5.4888e-08 | 6.0832e-09 | 74.7329 | -75.1716 | 3.4028e+38 | 3.4028e+38 | 689.9838 | 1380.0837 | 1195.2212 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 6.0003 | 689.9807 | 1380.0826 | 1195.2216 | 49.9999 | -60.0027 | 0.017639 | 3.946e-06 | 1.8585e-08 | 6.9852e-08 | 1.1948e-08 | 74.7571 | -43.7443 | 0.019448 | 3.2077e-06 | 2.3099e-09 | 6.4443e-08 | 1.1598e-08 | 74.6594 | -77.0631 | 3.4028e+38 | 3.4028e+38 | 689.9807 | 1380.0826 | 1195.2216 | 3.4028e+38 | 3.4028e+38 | 230.0025 | 6.0003 | 689.9842 | 1380.084 | 1195.2245 | 50 | -60.0028 | 0.015506 | 3.8918e-06 | 9.5539e-09 | 6.0386e-08 | 1.1332e-08 | 74.7485 | -50.3585 | 0.016923 | 3.2113e-06 | 1.0519e-10 | 5.433e-08 | 5.5843e-09 | 74.763 | -78.0192 | 3.4028e+38 | 3.4028e+38 | 689.9842 | 1380.084 | 1195.2245 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0023 | 6.0003 | 689.9813 | 1380.0806 | 1195.2207 | 50 | -60.0029 | 0.49996 | 3.4028e+38 | 0 | 230.0011 | 689.9497 | 1380.0781 | 1195.22 | -0.0020715 | 0.014829 | 3.9578e-06 | 6.0578e-09 | 5.7515e-08 | 7.1881e-09 | 74.7099 | -47.0228 | 0.10964 | 3.4028e+38 | 9.9426 | 0.019941 | 0 | 0 | 0 | 0.00010443 | 0.015861 | 3.2017e-06 | 1.1698e-09 | 5.1524e-08 | 4.2022e-09 | 74.7813 | -69.7731 | 0.029411 | 3.4028e+38 | -13.4124 | 229.2258 | -1.0797e-09 | 0.00074757 | -2.0884e-10 | -0.00010008 | 3.4028e+38 | 3.4028e+38 | 689.9797 | 1380.0825 | 1195.2252 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 689.9671 | 689.9657 | 1195.2365 | -0.0006291 | 230.0022 | 6.0003 | 689.9833 | 1380.0828 | 1195.2214 | 50 | -60.0028 | [17 empty] | 23.5827 | 29.2517 | -42.2294 | 51.3242 | 83.4623 |  | -0.0030826
88 | 230.0021 | 6.0003 | 689.9848 | 1380.0834 | 1195.2233 | 49.9999 | -60.003 | 0.017866 | 3.9273e-06 | 1.9052e-08 | 6.9103e-08 | 1.2303e-08 | 74.6376 | -32.4686 | 0.019507 | 3.3125e-06 | 6.5692e-09 | 6.5302e-08 | 1.1282e-08 | 74.7347 | -62.464 | 3.4028e+38 | 3.4028e+38 | 689.9848 | 1380.0834 | 1195.2233 | 3.4028e+38 | 3.4028e+38 | 230.002 | 6.0003 | 689.9792 | 1380.0802 | 1195.2224 | 49.9999 | -60.0028 | 0.017049 | 3.9223e-06 | 1.303e-08 | 6.6853e-08 | 1.1737e-08 | 74.6652 | -45.0262 | 0.018683 | 3.2318e-06 | 2.1475e-09 | 6.1527e-08 | 9.2461e-09 | 74.7508 | -77.6601 | 3.4028e+38 | 3.4028e+38 | 689.9792 | 1380.0802 | 1195.2224 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 6.0003 | 689.9879 | 1380.0836 | 1195.224 | 49.9999 | -60.0028 | 0.016431 | 4.0153e-06 | 1.7389e-08 | 6.6035e-08 | 5.0447e-09 | 74.5894 | -14.9334 | 0.018104 | 3.1857e-06 | 4.9769e-09 | 5.7482e-08 | 8.3475e-09 | 74.7521 | -56.0802 | 3.4028e+38 | 3.4028e+38 | 689.9879 | 1380.0836 | 1195.224 | 3.4028e+38 | 3.4028e+38 | 230.0026 | 6.0003 | 689.9837 | 1380.0891 | 1195.2297 | 49.9999 | -60.0029 | 0.014473 | 4.0386e-06 | 6.8913e-09 | 5.7134e-08 | 4.1385e-09 | 74.703 | -33.3557 | 0.015302 | 3.2461e-06 | 1.7426e-09 | 4.8072e-08 | 2.9577e-09 | 74.7882 | -70.3382 | 3.4028e+38 | 3.4028e+38 | 689.9837 | 1380.0891 | 1195.2297 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 6.0003 | 689.9805 | 1380.0809 | 1195.2205 | 49.9999 | -60.0028 | 0.017792 | 3.9364e-06 | 1.7604e-08 | 6.9009e-08 | 8.4968e-09 | 74.6383 | -20.7008 | 0.019441 | 3.3212e-06 | 5.7119e-09 | 6.4486e-08 | 9.7831e-09 | 74.8056 | -58.7045 | 3.4028e+38 | 3.4028e+38 | 689.9805 | 1380.0809 | 1195.2205 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0021 | 6.0003 | 689.9855 | 1380.0837 | 1195.2251 | 49.9999 | -60.0028 | 0.49996 | 3.4028e+38 | 0 | 230.0074 | 689.9543 | 1380.1151 | 1195.2528 | -0.0017811 | 0.016692 | 4.0193e-06 | 1.5001e-08 | 6.7043e-08 | 3.843e-09 | 74.6414 | -15.751 | 0.2047 | 3.4028e+38 | -5.6306 | 0.020636 | 0 | 0 | 0 | 8.7828e-05 | 0.017355 | 3.1539e-06 | 5.5745e-09 | 5.463e-08 | 6.8794e-09 | 74.7605 | -50.9643 | 0.079951 | 3.4028e+38 | -54.3882 | 229.161 | -9.5332e-10 | 0.00072153 | 5.8513e-10 | -0.00010059 | 3.4028e+38 | 3.4028e+38 | 689.9902 | 1380.0878 | 1195.2295 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 689.9724 | 689.9724 | 1195.2417 | -0.0015077 | 230.0021 | 6.0003 | 689.9832 | 1380.0834 | 1195.224 | 49.9999 | -60.0029 | [17 empty] | 23.0914 | 19.0189 | -49.3483 | 41.5364 | 74.2247 |  | -0.0030784
89 | 230.0024 | 6.0003 | 689.9787 | 1380.0857 | 1195.2266 | 49.9999 | -60.003 | 0.014418 | 3.9815e-06 | 7.6938e-09 | 5.7088e-08 | 4.8664e-09 | 74.5698 | -32.3136 | 0.015244 | 3.2885e-06 | 1.4635e-09 | 5.0253e-08 | 2.7648e-09 | 74.7018 | -81.0101 | 3.4028e+38 | 3.4028e+38 | 689.9787 | 1380.0857 | 1195.2266 | 3.4028e+38 | 3.4028e+38 | 230.0024 | 6.0003 | 689.9834 | 1380.0867 | 1195.2294 | 49.9999 | -60.003 | 0.014446 | 3.9909e-06 | 7.3774e-09 | 5.6794e-08 | 4.4436e-09 | 74.6429 | -32.5965 | 0.015306 | 3.2137e-06 | 1.3518e-09 | 4.9201e-08 | 3.193e-09 | 74.7479 | -66.7719 | 3.4028e+38 | 3.4028e+38 | 689.9834 | 1380.0867 | 1195.2294 | 3.4028e+38 | 3.4028e+38 | 230.0025 | 6.0003 | 689.9844 | 1380.0876 | 1195.2327 | 49.9999 | -60.0032 | 0.015175 | 4.0136e-06 | 1.3039e-08 | 6.2344e-08 | 3.1499e-09 | 74.7473 | -13.3093 | 0.016451 | 3.2036e-06 | 3.7072e-09 | 5.3709e-08 | 5.56e-09 | 74.7537 | -49.1921 | 3.4028e+38 | 3.4028e+38 | 689.9844 | 1380.0876 | 1195.2327 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 6.0003 | 689.9834 | 1380.0813 | 1195.2213 | 50 | -60.0028 | 0.016176 | 3.9726e-06 | 1.5288e-08 | 6.369e-08 | 3.3518e-09 | 74.546 | -12.733 | 0.017718 | 3.2475e-06 | 6.485e-09 | 5.7394e-08 | 8.2972e-09 | 74.746 | -55.1052 | 3.4028e+38 | 3.4028e+38 | 689.9834 | 1380.0813 | 1195.2213 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 6.0003 | 689.9791 | 1380.0812 | 1195.2212 | 50 | -60.0029 | 0.01677 | 3.9207e-06 | 1.126e-08 | 6.5822e-08 | 1.2436e-08 | 74.7655 | -45.0515 | 0.018366 | 3.25e-06 | 1.2084e-09 | 6.0943e-08 | 8.6148e-09 | 74.6475 | -76.3496 | 3.4028e+38 | 3.4028e+38 | 689.9791 | 1380.0812 | 1195.2212 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0021 | 6.0003 | 689.98 | 1380.0809 | 1195.222 | 50 | -60.0031 | 0.49996 | 3.4028e+38 | 0 | 230.001 | 689.93 | 1380.0798 | 1195.2161 | -0.0054408 | 0.016861 | 3.9023e-06 | 1.3888e-08 | 6.529e-08 | 1.3952e-08 | 74.763 | -45.5693 | 0.22632 | 3.4028e+38 | -49.5942 | 0.020583 | 0 | 0 | 0 | 8.9585e-05 | 0.018664 | 3.2928e-06 | 1.8309e-09 | 6.0339e-08 | 9.7256e-09 | 74.7479 | -71.201 | 0.024044 | 3.4028e+38 | 28.4114 | 229.7362 | -1.1361e-09 | 0.00074504 | 1.8959e-10 | -0.00011678 | 3.4028e+38 | 3.4028e+38 | 689.9799 | 1380.0818 | 1195.2223 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 689.931 | 689.9355 | 1195.2087 | -0.0038155 | 230.0024 | 6.0003 | 689.9818 | 1380.0845 | 1195.2262 | 49.9999 | -60.003 | [17 empty] | 24.5884 | 25.414 | -41.1549 | 48.8548 | 81.4328 |  | -0.0030563
90 | 230.002 | 6.0003 | 1104.1053 | 1380.0781 | -828.0129 | 50 | 36.8677 | 0.01483 | 3.9223e-06 | 4.2232e-09 | 5.814e-08 | -6.0111e-09 | 74.657 | 61.86 | 0.016313 | 3.3623e-06 | 6.8228e-09 | 5.4542e-08 | 3.5129e-09 | 74.7222 | -26.1612 | 3.4028e+38 | 3.4028e+38 | 1104.1053 | 1380.0781 | -828.013 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 6.0003 | 1104.1071 | 1380.0779 | -828.0064 | 49.9999 | 36.8675 | 0.014897 | 3.9519e-06 | -4.604e-09 | 5.805e-08 | -9.6784e-09 | 74.7554 | 110.9984 | 0.016218 | 3.2157e-06 | 3.7468e-09 | 5.328e-08 | -2.6272e-09 | 74.8588 | 39.8386 | 3.4028e+38 | 3.4028e+38 | 1104.1069 | 1380.0779 | -828.0065 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 6.0003 | 1104.1023 | 1380.0751 | -828.0089 | 49.9999 | 36.8675 | 0.014135 | 3.9276e-06 | 2.2918e-09 | 5.543e-08 | -4.5789e-09 | 74.8544 | 67.6671 | 0.015415 | 3.3341e-06 | 5.732e-09 | 5.0156e-08 | 1.7672e-09 | 74.7579 | -20.544 | 3.4028e+38 | 3.4028e+38 | 1104.1023 | 1380.0751 | -828.009 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 6.0003 | 1104.1108 | 1380.0791 | -828.0082 | 50 | 36.8675 | 0.016229 | 4.0118e-06 | -5.9012e-09 | 6.6665e-08 | -1.7323e-08 | 74.7191 | 106.783 | 0.018071 | 3.2905e-06 | 5.7252e-09 | 5.943e-08 | -1.9055e-09 | 74.6967 | 18.8451 | 3.4028e+38 | 3.4028e+38 | 1104.1107 | 1380.0791 | -828.0082 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 6.0003 | 1104.1047 | 1380.0767 | -828.0074 | 50 | 36.8674 | 0.015214 | 3.9654e-06 | -2.8714e-09 | 5.9592e-08 | -9.4983e-09 | 74.6643 | 112.8937 | 0.016325 | 3.2035e-06 | 3.0369e-09 | 5.342e-08 | -2.9305e-09 | 74.7946 | 43.9781 | 3.4028e+38 | 3.4028e+38 | 1104.1046 | 1380.0767 | -828.0075 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 6.0003 | 1104.105 | 1380.0765 | -828.0082 | 49.9999 | 36.8675 | 0.80003 | 3.4028e+38 | 0 | 230.0054 | 1104.1399 | 1380.1012 | -827.9871 | 0.0017236 | 0.014599 | 3.9517e-06 | -1.918e-09 | 5.7861e-08 | -7.7946e-09 | 74.7311 | 106.0704 | -0.036278 | 3.4028e+38 | 36.2295 | 0.019828 | 0 | 0 | 0 | -0.00021188 | 0.015976 | 3.3025e-06 | 4.9551e-09 | 5.2791e-08 | -1.0737e-09 | 74.8076 | 21.3029 | 0.067466 | 3.4028e+38 | -30.8235 | 229.5789 | -2.3491e-10 | 0.00073737 | -7.825e-10 | 2.4944e-05 | 3.4028e+38 | 3.4028e+38 | 1104.1064 | 1380.0774 | -828.0073 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1104.1407 | 1104.1407 | -827.9818 | 0.00073799 | 230.0018 | 6.0003 | 1104.1061 | 1380.0774 | -828.0088 | 49.9999 | 36.8675 | [17 empty] | 22.3175 | 43.9725 | 64.9678 | 65.0036 | 90.7046 |  | 0.00067592
91 | 230.002 | 6.0003 | 1104.1108 | 1380.0825 | -828.0074 | 49.9999 | 36.8674 | 0.016192 | 3.9987e-06 | -4.2485e-09 | 6.4782e-08 | -1.4519e-08 | 74.7231 | 104.424 | 0.017791 | 3.2543e-06 | 5.9318e-09 | 5.7661e-08 | -2.9352e-09 | 74.7577 | 31.7391 | 3.4028e+38 | 3.4028e+38 | 1104.1107 | 1380.0825 | -828.0075 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 6.0003 | 1104.1049 | 1380.0789 | -828.0102 | 50 | 36.8679 | 0.017078 | 3.9366e-06 | 6.6951e-09 | 6.82e-08 | -1.2088e-08 | 74.6589 | 75.9529 | 0.018877 | 3.198e-06 | 1.0898e-08 | 6.2651e-08 | 4.635e-09 | 74.7591 | -23.0873 | 3.4028e+38 | 3.4028e+38 | 1104.1049 | 1380.0789 | -828.0103 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 6.0003 | 1104.1079 | 1380.0776 | -828.0074 | 49.9999 | 36.8675 | 0.016814 | 4.008e-06 | -4.9269e-09 | 6.7291e-08 | -1.5589e-08 | 74.7018 | 108.4094 | 0.018206 | 3.2235e-06 | 5.2853e-09 | 5.8551e-08 | -2.5344e-09 | 74.7784 | 20.537 | 3.4028e+38 | 3.4028e+38 | 1104.1078 | 1380.0776 | -828.0074 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 6.0003 | 1104.1038 | 1380.0776 | -828.0118 | 49.9999 | 36.8676 | 0.017674 | 4.0444e-06 | 2.795e-09 | 7.0829e-08 | -1.6623e-08 | 74.6107 | 82.8089 | 0.019519 | 3.2864e-06 | 1.308e-08 | 6.2935e-08 | 3.4741e-09 | 74.7047 | -17.7936 | 3.4028e+38 | 3.4028e+38 | 1104.1038 | 1380.0776 | -828.0118 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 6.0003 | 1104.1003 | 1380.0741 | -828.0107 | 50 | 36.8677 | 0.016482 | 3.9839e-06 | 6.3405e-09 | 6.5408e-08 | -9.1739e-09 | 74.6253 | 58.3233 | 0.017719 | 3.2602e-06 | 1.0028e-08 | 6.0678e-08 | 5.8354e-09 | 74.7527 | -30.3544 | 3.4028e+38 | 3.4028e+38 | 1104.1003 | 1380.0741 | -828.0108 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0014 | 6.0003 | 1104.1003 | 1380.0747 | -828.0098 | 50 | 36.8678 | 0.80003 | 3.4028e+38 | 0 | 230.0025 | 1104.1475 | 1380.0853 | -827.9528 | 0.0028599 | 0.017328 | 3.9308e-06 | 3.1418e-09 | 6.8108e-08 | -1.5078e-08 | 74.6351 | 83.0486 | 0.029422 | 3.4028e+38 | -38.7776 | 0.021039 | 0 | 0 | 0 | -0.00023106 | 0.019526 | 3.2848e-06 | 1.1803e-08 | 6.3761e-08 | 2.2652e-09 | 74.6857 | -12.4284 | 0.15897 | 3.4028e+38 | -34.2805 | 229.0365 | 3.8966e-11 | 0.0007422 | -1.2798e-09 | 8.3062e-06 | 3.4028e+38 | 3.4028e+38 | 1104.1095 | 1380.0804 | -828.0079 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1104.1611 | 1104.1611 | -827.9515 | 0.0045347 | 230.0017 | 6.0003 | 1104.1055 | 1380.0781 | -828.0095 | 50 | 36.8676 | [17 empty] | 22.0878 | 43.6844 | 64.5768 | 65.4371 | 91.0906 |  | 0.00077388
92 | 230.0021 | 6.0003 | 1104.1075 | 1380.08 | -828.0102 | 49.9999 | 36.8677 | 0.017215 | 4.0539e-06 | -3.1007e-09 | 6.934e-08 | -1.7625e-08 | 74.6794 | 101.9025 | 0.019266 | 3.1928e-06 | 1.1344e-08 | 6.3831e-08 | 5.3472e-10 | 74.7854 | -3.1741 | 3.4028e+38 | 3.4028e+38 | 1104.1074 | 1380.08 | -828.0103 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 6.0003 | 1104.1011 | 1380.0773 | -828.005 | 49.9999 | 36.8674 | 0.015962 | 3.9089e-06 | 5.5344e-09 | 6.0658e-08 | -7.431e-09 | 74.7781 | 57.6183 | 0.017141 | 3.3482e-06 | 7.7895e-09 | 5.6562e-08 | 3.688e-09 | 74.7207 | -25.3357 | 3.4028e+38 | 3.4028e+38 | 1104.1011 | 1380.0773 | -828.0051 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 6.0003 | 1104.1121 | 1380.0808 | -828.004 | 49.9999 | 36.8676 | 0.017079 | 3.9931e-06 | -1.7041e-09 | 6.8319e-08 | -1.6847e-08 | 74.6487 | 98.6639 | 0.01862 | 3.2892e-06 | 6.5698e-09 | 6.1327e-08 | -1.9596e-09 | 74.7153 | 6.4427 | 3.4028e+38 | 3.4028e+38 | 1104.1119 | 1380.0808 | -828.0041 | 3.4028e+38 | 3.4028e+38 | 230.002 | 6.0003 | 1104.1088 | 1380.0739 | -828.0095 | 49.9999 | 36.8677 | 0.014295 | 3.9642e-06 | -1.9527e-09 | 5.7706e-08 | -7.4098e-09 | 74.579 | 105.0136 | 0.014974 | 3.1929e-06 | 3.3174e-09 | 5.1317e-08 | -9.7894e-10 | 74.6926 | 18.4941 | 3.4028e+38 | 3.4028e+38 | 1104.1086 | 1380.0739 | -828.0095 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 6.0003 | 1104.1018 | 1380.0769 | -828.0109 | 49.9999 | 36.8676 | 0.017762 | 4.0014e-06 | 1.4722e-09 | 7.0796e-08 | -1.6115e-08 | 74.6945 | 81.0648 | 0.019642 | 3.2095e-06 | 1.2166e-08 | 6.2668e-08 | 2.7763e-09 | 74.6325 | -13.7322 | 3.4028e+38 | 3.4028e+38 | 1104.1018 | 1380.0769 | -828.0109 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 6.0003 | 1104.1085 | 1380.0817 | -828.0093 | 50 | 36.8675 | 0.80003 | 3.4028e+38 | 0 | 230.0025 | 1104.1422 | 1380.0819 | -827.9556 | 0.0019555 | 0.017306 | 4.0076e-06 | 3.4306e-09 | 6.9248e-08 | -1.3434e-08 | 74.575 | 70.4155 | 0.071817 | 3.4028e+38 | -55.9451 | 0.019824 | 0 | 0 | 0 | -0.00022414 | 0.018056 | 3.2501e-06 | 1.0863e-08 | 5.8497e-08 | 6.1102e-09 | 74.8056 | -30.8989 | 0.18831 | 3.4028e+38 | -46.4202 | 229.1378 | 5.0409e-10 | 0.00076583 | 1.2781e-09 | 1.9759e-05 | 3.4028e+38 | 3.4028e+38 | 1104.1005 | 1380.075 | -828.0138 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1104.12 | 1104.1185 | -827.9803 | -0.00089144 | 230.0018 | 6.0003 | 1104.1063 | 1380.0778 | -828.0079 | 49.9999 | 36.8676 | [17 empty] | 21.884 | 45.2513 | 62.6464 | 65.3479 | 94.2706 |  | 0.00094701
93 | 230.0022 | 10.0004 | 2300.1189 | 2300.1135 | 0.10838 | 50 | -0.0027893 | 0.015983 | 4.7618e-06 | 8.8933e-09 | 7.6125e-08 | -1.4709e-08 | 74.6659 | 53.1271 | 0.016963 | 3.4708e-06 | 1.2612e-08 | 5.9983e-08 | 4.2691e-09 | 74.6974 | -16.1302 | 3.4028e+38 | 3.4028e+38 | 2300.1189 | 2300.1135 | 0.10838 | 3.4028e+38 | 3.4028e+38 | 230.002 | 10.0004 | 2300.113 | 2300.1082 | 0.1062 | 49.9999 | -0.0026113 | 0.016103 | 4.7503e-06 | 4.44e-09 | 7.5069e-08 | -1.9103e-08 | 74.5141 | 79.3761 | 0.017139 | 3.4805e-06 | 1.1549e-08 | 5.9345e-08 | 3.4298e-09 | 74.6465 | -14.2626 | 3.4028e+38 | 3.4028e+38 | 2300.113 | 2300.1082 | 0.1062 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 10.0004 | 2300.1184 | 2300.1135 | 0.1071 | 50 | -0.0026679 | 0.016894 | 4.6328e-06 | 1.9999e-08 | 7.827e-08 | -3.69e-09 | 74.6236 | 9.9713 | 0.018678 | 3.5717e-06 | 1.5283e-08 | 6.6629e-08 | 1.5147e-08 | 74.5441 | -48.8224 | 3.4028e+38 | 3.4028e+38 | 2300.1184 | 2300.1135 | 0.1071 | 3.4028e+38 | 3.4028e+38 | 230.002 | 10.0004 | 2300.116 | 2300.1111 | 0.099746 | 49.9999 | -0.0025748 | 0.017193 | 4.6349e-06 | 2.0417e-08 | 7.874e-08 | -5.8455e-09 | 74.4882 | 15.9068 | 0.019068 | 3.4903e-06 | 1.4702e-08 | 6.6345e-08 | 1.2887e-08 | 74.7886 | -43.2026 | 3.4028e+38 | 3.4028e+38 | 2300.116 | 2300.1111 | 0.099746 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 10.0004 | 2300.1177 | 2300.1133 | 0.097304 | 50 | -0.0024238 | 0.017735 | 4.7177e-06 | 1.9893e-08 | 8.4367e-08 | -1.1592e-08 | 74.5548 | 24.627 | 0.019766 | 3.4706e-06 | 1.7154e-08 | 6.8548e-08 | 1.1479e-08 | 74.7272 | -37.272 | 3.4028e+38 | 3.4028e+38 | 2300.1177 | 2300.1133 | 0.097304 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0017 | 10.0004 | 2300.1167 | 2300.1113 | 0.10708 | 49.9999 | -0.0026065 | 1 | 3.4028e+38 | 0 | 230.0033 | 2300.1265 | 2300.1216 | 0.20155 | 0.00021114 | 0.0176 | 4.6446e-06 | 2.5179e-08 | 8.2377e-08 | -6.0282e-09 | 74.5266 | 10.6707 | 0.28667 | 3.4028e+38 | 24.8025 | 0.020239 | 0 | 0 | 0 | -0.00023947 | 0.019495 | 3.4126e-06 | 1.8138e-08 | 6.707e-08 | 1.5466e-08 | 74.6917 | -38.837 | 0.27866 | 3.4028e+38 | 46.9134 | 229.0936 | 2.7105e-09 | 0.00079994 | 1.5932e-09 | -8.3252e-05 | 3.4028e+38 | 3.4028e+38 | 2300.116 | 2300.1106 | 0.099853 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 2300.1162 | 2300.116 | 0.19202 | -0.0030645 | 230.0021 | 10.0004 | 2300.1168 | 2300.1119 | 0.10375 | 49.9999 | -0.0026134 | [17 empty] | 22.6309 | -0.47797 | 23.1202 | 21.0033 | 2543904.6085 |  | -0.0018842
94 | 230.0023 | 10.0004 | 2300.1223 | 2300.1172 | 0.1051 | 50 | -0.0027782 | 0.014555 | 4.6546e-06 | 5.575e-09 | 6.8137e-08 | -8.1858e-09 | 74.5716 | 46.8243 | 0.015582 | 3.5363e-06 | 8.005e-09 | 5.6017e-08 | 2.5803e-09 | 74.6568 | -20.0412 | 3.4028e+38 | 3.4028e+38 | 2300.1223 | 2300.1172 | 0.1051 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 10.0004 | 2300.1204 | 2300.115 | 0.11249 | 49.9999 | -0.002802 | 0.014354 | 4.7086e-06 | 7.3525e-09 | 6.7684e-08 | -5.9555e-09 | 74.6053 | 39.9705 | 0.015139 | 3.4278e-06 | 7.5454e-09 | 5.2576e-08 | 1.5396e-09 | 74.627 | -12.6217 | 3.4028e+38 | 3.4028e+38 | 2300.1204 | 2300.115 | 0.11249 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 10.0004 | 2300.1108 | 2300.1062 | 0.10579 | 50 | -0.0024708 | 0.015224 | 4.5286e-06 | 1.4665e-08 | 6.8675e-08 | 1.3527e-10 | 74.5795 | -2.6622 | 0.016405 | 3.4544e-06 | 7.8381e-09 | 5.6618e-08 | 9.3381e-09 | 74.719 | -50.2637 | 3.4028e+38 | 3.4028e+38 | 2300.1108 | 2300.1062 | 0.10579 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 10.0004 | 2300.1223 | 2300.1172 | 0.1032 | 50 | -0.0027494 | 0.014399 | 4.6821e-06 | 8.9725e-09 | 6.673e-08 | -1.6713e-09 | 74.7192 | 9.0734 | 0.015382 | 3.5421e-06 | 7.5049e-09 | 5.3761e-08 | 5.0261e-09 | 74.7137 | -40.5402 | 3.4028e+38 | 3.4028e+38 | 2300.1223 | 2300.1172 | 0.1032 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 10.0004 | 2300.1169 | 2300.1121 | 0.10509 | 50 | -0.0026177 | 0.016098 | 4.5913e-06 | 2.1615e-08 | 7.4894e-08 | 2.23e-09 | 74.7669 | -2.9119 | 0.017496 | 3.5472e-06 | 1.345e-08 | 6.1698e-08 | 1.4944e-08 | 74.6463 | -44.9151 | 3.4028e+38 | 3.4028e+38 | 2300.1169 | 2300.1121 | 0.10509 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 10.0004 | 2300.1096 | 2300.1047 | 0.098246 | 50 | -0.0026647 | 1 | 3.4028e+38 | 0 | 230.0026 | 2300.116 | 2300.1111 | 0.19469 | -7.5941e-05 | 0.01759 | 4.6793e-06 | 1.7203e-08 | 8.174e-08 | -1.2001e-08 | 74.516 | 33.2139 | 0.22686 | 3.4028e+38 | 21.375 | 0.021068 | 0 | 0 | 0 | -0.00024521 | 0.019452 | 3.4629e-06 | 1.6268e-08 | 6.7329e-08 | 9.3339e-09 | 74.7014 | -30.3228 | 0.23726 | 3.4028e+38 | 41.5411 | 229.7388 | -9.2964e-10 | 0.00078476 | 6.5557e-10 | -7.978e-05 | 3.4028e+38 | 3.4028e+38 | 2300.123 | 2300.1182 | 0.10292 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 2300.1699 | 2300.1699 | 0.18559 | 0.0025554 | 230.0021 | 10.0004 | 2300.1186 | 2300.1135 | 0.10633 | 50 | -0.0026836 | [17 empty] | 22.7069 | 0.90508 | 26.5405 | 24.3596 | 1761303.2754 |  | -0.0017244
95 | 230.0019 | 10.0004 | 2300.1143 | 2300.1094 | 0.10555 | 50 | -0.0025767 | 0.014917 | 4.6052e-06 | 1.2592e-08 | 6.856e-08 | -3.3552e-10 | 74.7178 | 1.6231 | 0.015952 | 3.4277e-06 | 7.2838e-09 | 5.6519e-08 | 7.441e-09 | 74.6638 | -44.6787 | 3.4028e+38 | 3.4028e+38 | 2300.1143 | 2300.1094 | 0.10555 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 10.0004 | 2300.1223 | 2300.1172 | 0.10502 | 49.9999 | -0.0024578 | 0.017552 | 4.6653e-06 | 2.4258e-08 | 8.265e-08 | -7.2761e-09 | 74.4513 | 14.908 | 0.019713 | 3.4681e-06 | 1.9428e-08 | 6.7744e-08 | 1.2118e-08 | 74.6922 | -34.326 | 3.4028e+38 | 3.4028e+38 | 2300.1223 | 2300.1172 | 0.10502 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 10.0004 | 2300.116 | 2300.1113 | 0.10008 | 49.9999 | -0.0024929 | 0.014602 | 4.6983e-06 | 5.2752e-09 | 6.9406e-08 | -9.1373e-09 | 74.612 | 60.001 | 0.015861 | 3.4669e-06 | 8.164e-09 | 5.3946e-08 | 1.3908e-09 | 74.4553 | -6.9587 | 3.4028e+38 | 3.4028e+38 | 2300.116 | 2300.1113 | 0.10008 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 10.0004 | 2300.1218 | 2300.1169 | 0.1014 | 50 | -0.002524 | 0.014645 | 4.6828e-06 | 5.7975e-09 | 6.8112e-08 | -1.0228e-08 | 74.3608 | 59.127 | 0.01591 | 3.4926e-06 | 7.0614e-09 | 5.4272e-08 | 8.6174e-10 | 74.7835 | -6.4619 | 3.4028e+38 | 3.4028e+38 | 2300.1218 | 2300.1169 | 0.1014 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 10.0004 | 2300.1184 | 2300.1133 | 0.1013 | 49.9999 | -0.0025234 | 0.01555 | 4.5663e-06 | 1.6992e-08 | 7.1155e-08 | 3.0143e-10 | 74.649 | -0.97551 | 0.016719 | 3.4596e-06 | 8.0182e-09 | 5.9939e-08 | 9.5747e-09 | 74.729 | -48.2595 | 3.4028e+38 | 3.4028e+38 | 2300.1184 | 2300.1133 | 0.1013 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0021 | 10.0004 | 2300.1123 | 2300.1074 | 0.10833 | 49.9999 | -0.0026435 | 1 | 3.4028e+38 | 0 | 230.0005 | 2300.0977 | 2300.0928 | 0.1631 | -0.0039563 | 0.016072 | 4.593e-06 | 1.9642e-08 | 7.4391e-08 | -1.923e-10 | 74.5982 | 0.58016 | 0.26217 | 3.4028e+38 | 1.8332 | 0.020629 | 0 | 0 | 0 | -0.00022503 | 0.018071 | 3.5626e-06 | 1.2348e-08 | 6.415e-08 | 1.5646e-08 | 74.7642 | -48.8788 | 0.21292 | 3.4028e+38 | -66.3593 | 229.5809 | 1.4247e-09 | 0.00079755 | 3.329e-09 | -8.2304e-05 | 3.4028e+38 | 3.4028e+38 | 2300.1047 | 2300.0996 | 0.10671 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 2300.0962 | 2300.0962 | 0.17895 | -0.0050534 | 230.0021 | 10.0004 | 2300.1186 | 2300.1136 | 0.10267 | 50 | -0.002515 | [17 empty] | 22.8995 | -0.25093 | 24.7614 | 22.6227 | 1789586.8709 |  | -0.0015982
96 | 230.0022 | 10.0004 | 1149.9764 | 2300.1226 | 1992.0131 | 50 | -60.0025 | 0.017283 | 4.6982e-06 | 3.0721e-08 | 8.1396e-08 | 1.2654e-08 | 74.7049 | -21.4987 | 0.018946 | 3.393e-06 | 1.0454e-08 | 6.49e-08 | 1.6801e-08 | 74.8435 | -56.3139 | 3.4028e+38 | 3.4028e+38 | 1149.9764 | 2300.1226 | 1992.0131 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 10.0004 | 1149.9639 | 2300.1162 | 1992.0176 | 50 | -60.0025 | 0.014784 | 4.6057e-06 | 1.3491e-08 | 6.9184e-08 | 1.6283e-08 | 74.593 | -50.4665 | 0.016167 | 3.38e-06 | 3.9204e-10 | 5.4701e-08 | 6.7201e-09 | 74.5588 | -102.8433 | 3.4028e+38 | 3.4028e+38 | 1149.9639 | 2300.1162 | 1992.0176 | 3.4028e+38 | 3.4028e+38 | 230.002 | 10.0004 | 1149.9725 | 2300.1162 | 1992.0115 | 49.9999 | -60.0026 | 0.017606 | 4.5794e-06 | 2.6944e-08 | 8.0377e-08 | 2.0225e-08 | 74.427 | -36.5024 | 0.019227 | 3.406e-06 | 3.4092e-09 | 6.5792e-08 | 1.6106e-08 | 74.7802 | -75.7638 | 3.4028e+38 | 3.4028e+38 | 1149.9725 | 2300.1162 | 1992.0115 | 3.4028e+38 | 3.4028e+38 | 230.0024 | 10.0004 | 1149.9719 | 2300.1167 | 1992.0145 | 49.9999 | -60.0026 | 0.014924 | 4.7418e-06 | 1.7981e-08 | 7.1838e-08 | 4.4981e-09 | 74.6081 | -13.8806 | 0.015925 | 3.3725e-06 | 6.3055e-09 | 5.5955e-08 | 7.3081e-09 | 74.6904 | -45.563 | 3.4028e+38 | 3.4028e+38 | 1149.9719 | 2300.1167 | 1992.0145 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 10.0004 | 1149.9684 | 2300.116 | 1992.0168 | 49.9999 | -60.0026 | 0.017883 | 4.6161e-06 | 2.8773e-08 | 8.267e-08 | 1.5728e-08 | 74.4702 | -31.0009 | 0.019618 | 3.4384e-06 | 1.022e-08 | 6.7396e-08 | 1.7357e-08 | 74.588 | -62.986 | 3.4028e+38 | 3.4028e+38 | 1149.9684 | 2300.116 | 1992.0168 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0021 | 10.0004 | 1149.9727 | 2300.1152 | 1992.0118 | 49.9999 | -60.0025 | 0.49996 | 3.4028e+38 | 0 | 230.005 | 1149.8898 | 2300.1406 | 1992.0381 | -0.0065451 | 0.017691 | 4.6481e-06 | 3.1138e-08 | 8.2637e-08 | 2.1429e-08 | 74.5173 | -32.2862 | 0.39648 | 3.4028e+38 | -17.9373 | 0.021148 | 0 | 0 | 0 | 0.00013933 | 0.019479 | 3.5169e-06 | 6.5611e-09 | 6.9851e-08 | 1.8966e-08 | 74.5941 | -68.642 | 0.11184 | 3.4028e+38 | -28.588 | 229.3006 | 4.2471e-09 | 0.0008008 | 5.8068e-09 | -0.00018602 | 3.4028e+38 | 3.4028e+38 | 1149.9664 | 2300.1106 | 1992.0077 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1149.8918 | 1149.8918 | 1992.0415 | -0.0074233 | 230.0022 | 10.0004 | 1149.9706 | 2300.1175 | 1992.0147 | 49.9999 | -60.0026 | [17 empty] | 23.485 | 36.3622 | -59.465 | 59.677 | 101.1939 |  | -0.0039671
97 | 230.0022 | 10.0004 | 1149.9741 | 2300.114 | 1992.0054 | 50 | -60.0026 | 0.016151 | 4.6273e-06 | 1.4405e-08 | 7.2471e-08 | 1.9417e-08 | 74.5877 | -51.2236 | 0.01713 | 3.42e-06 | 9.4668e-10 | 5.8813e-08 | 9.8795e-09 | 74.5936 | -91.5481 | 3.4028e+38 | 3.4028e+38 | 1149.9741 | 2300.114 | 1992.0054 | 3.4028e+38 | 3.4028e+38 | 230.0026 | 10.0004 | 1149.9757 | 2300.1252 | 1992.0153 | 49.9999 | -60.0026 | 0.014782 | 4.6493e-06 | 1.398e-08 | 6.8837e-08 | 1.3838e-08 | 74.7502 | -50.0584 | 0.015811 | 3.4045e-06 | -5.1775e-10 | 5.4091e-08 | 6.9161e-09 | 74.734 | -82.6877 | 3.4028e+38 | 3.4028e+38 | 1149.9756 | 2300.1252 | 1992.0153 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 10.0004 | 1149.9673 | 2300.116 | 1992.0188 | 49.9999 | -60.0026 | 0.01584 | 4.6433e-06 | 1.5596e-08 | 7.3418e-08 | 1.7379e-08 | 74.6653 | -50.056 | 0.0171 | 3.4215e-06 | 1.1784e-09 | 5.8153e-08 | 9.5704e-09 | 74.8438 | -88.3841 | 3.4028e+38 | 3.4028e+38 | 1149.9673 | 2300.116 | 1992.0188 | 3.4028e+38 | 3.4028e+38 | 230.0026 | 10.0004 | 1149.9675 | 2300.1125 | 1992.01 | 49.9999 | -60.0027 | 0.01696 | 4.572e-06 | 2.212e-08 | 7.929e-08 | 2.2953e-08 | 74.6367 | -46.3091 | 0.018795 | 3.5367e-06 | 2.5147e-09 | 6.646e-08 | 1.3193e-08 | 74.7203 | -83.5768 | 3.4028e+38 | 3.4028e+38 | 1149.9675 | 2300.1125 | 1992.01 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 10.0004 | 1149.9641 | 2300.1125 | 1992.015 | 49.9999 | -60.0027 | 0.016261 | 4.5918e-06 | 2.0599e-08 | 7.4497e-08 | 2.0219e-08 | 74.6084 | -46.0204 | 0.017679 | 3.4762e-06 | -6.265e-10 | 6.262e-08 | 1.1203e-08 | 74.6704 | -88.9123 | 3.4028e+38 | 3.4028e+38 | 1149.964 | 2300.1125 | 1992.015 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0024 | 10.0004 | 1149.9669 | 2300.1177 | 1992.0155 | 50 | -60.0026 | 0.49996 | 3.4028e+38 | 0 | 230.001 | 1149.8865 | 2300.1023 | 1992.0013 | -0.0070477 | 0.016712 | 4.5731e-06 | 2.2369e-08 | 7.6923e-08 | 2.3865e-08 | 74.722 | -45.3652 | 0.30632 | 3.4028e+38 | -55.7917 | 0.020897 | 0 | 0 | 0 | 0.0001536 | 0.018448 | 3.4932e-06 | 3.4768e-10 | 6.2047e-08 | 1.2741e-08 | 74.6392 | -88.4369 | -0.018098 | 3.4028e+38 | -28.2188 | 229.1926 | -2.6292e-09 | 0.00079167 | 6.2013e-10 | -0.00018728 | 3.4028e+38 | 3.4028e+38 | 1149.9713 | 2300.1191 | 1992.0153 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1149.8821 | 1149.8821 | 1992.0027 | -0.0065932 | 230.0023 | 10.0004 | 1149.9698 | 2300.1161 | 1992.0129 | 49.9999 | -60.0027 | [17 empty] | 23.6903 | 41.2361 | -61.3822 | 63.8922 | 107.1406 |  | -0.004255
98 | 230.0019 | 10.0004 | 1149.9641 | 2300.1155 | 1992.0104 | 49.9999 | -60.0026 | 0.01763 | 4.6399e-06 | 2.6993e-08 | 8.18e-08 | 1.1153e-08 | 74.5935 | -22.3338 | 0.019387 | 3.3935e-06 | 9.1216e-09 | 6.5977e-08 | 1.7513e-08 | 74.7357 | -61.3292 | 3.4028e+38 | 3.4028e+38 | 1149.9641 | 2300.1155 | 1992.0104 | 3.4028e+38 | 3.4028e+38 | 230.0027 | 10.0004 | 1149.9707 | 2300.1287 | 1992.0201 | 49.9999 | -60.0026 | 0.015703 | 4.7427e-06 | 1.923e-08 | 7.2875e-08 | 5.0876e-09 | 74.5326 | -15.9622 | 0.016561 | 3.4459e-06 | 6.2456e-09 | 5.492e-08 | 8.1533e-09 | 74.7191 | -49.5066 | 3.4028e+38 | 3.4028e+38 | 1149.9707 | 2300.1287 | 1992.0201 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 10.0004 | 1149.9709 | 2300.1194 | 1992.0144 | 50 | -60.0026 | 0.017331 | 4.6172e-06 | 2.2898e-08 | 7.9714e-08 | 2.0572e-08 | 74.4601 | -38.152 | 0.019016 | 3.5174e-06 | 2.1966e-09 | 6.7518e-08 | 1.5469e-08 | 74.6727 | -75.758 | 3.4028e+38 | 3.4028e+38 | 1149.9709 | 2300.1194 | 1992.0144 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 10.0004 | 1149.9656 | 2300.1113 | 1992.005 | 50 | -60.0025 | 0.016558 | 4.5907e-06 | 2.4841e-08 | 7.709e-08 | 2.4391e-08 | 74.5315 | -45.6844 | 0.018126 | 3.4319e-06 | 2.1224e-09 | 6.3302e-08 | 1.372e-08 | 74.7824 | -85.2321 | 3.4028e+38 | 3.4028e+38 | 1149.9656 | 2300.1113 | 1992.005 | 3.4028e+38 | 3.4028e+38 | 230.0026 | 10.0004 | 1149.9702 | 2300.1147 | 1992.0081 | 49.9999 | -60.0026 | 0.017735 | 4.5958e-06 | 3.2135e-08 | 8.1754e-08 | 1.9166e-08 | 74.5904 | -32.5181 | 0.019556 | 3.4027e-06 | 7.8418e-09 | 6.6549e-08 | 1.9132e-08 | 74.6196 | -63.4456 | 3.4028e+38 | 3.4028e+38 | 1149.9702 | 2300.1147 | 1992.0081 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 10.0004 | 1149.9733 | 2300.1121 | 1992.0092 | 49.9999 | -60.0026 | 0.49996 | 3.4028e+38 | 0 | 230.0059 | 1149.9011 | 2300.1482 | 1992.0472 | -0.0053 | 0.017717 | 4.6347e-06 | 3.4655e-08 | 8.2365e-08 | 1.784e-08 | 74.6052 | -27.8994 | 0.4112 | 3.4028e+38 | 24.6557 | 0.021056 | 0 | 0 | 0 | 0.00013849 | 0.019423 | 3.4281e-06 | 8.7064e-09 | 6.8512e-08 | 1.5814e-08 | 74.7777 | -62.4748 | 0.12028 | 3.4028e+38 | 13.7125 | 229.1359 | -1.0489e-09 | 0.00081296 | 1.3969e-09 | -0.00018632 | 3.4028e+38 | 3.4028e+38 | 1149.9712 | 2300.1128 | 1992.0066 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1149.9 | 1149.9 | 1992.0466 | -0.0073075 | 230.0023 | 10.0004 | 1149.9683 | 2300.1179 | 1992.0116 | 50 | -60.0026 | [17 empty] | 23.2793 | 32.3933 | -61.68 | 55.2832 | 93.618 |  | -0.0038608
99 | 230.0014 | 10.0004 | 1840.1504 | 2300.1099 | -1380.0049 | 49.9999 | 36.8677 | 0.01764 | 4.7044e-06 | -4.6717e-09 | 8.3076e-08 | -3.169e-08 | 74.4976 | 98.386 | 0.019527 | 3.4442e-06 | 1.5711e-08 | 6.6306e-08 | 1.4456e-09 | 74.5559 | -4.9478 | 3.4028e+38 | 3.4028e+38 | 1840.1503 | 2300.1099 | -1380.005 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 10.0004 | 1840.1473 | 2300.1079 | -1380.0135 | 49.9999 | 36.8679 | 0.016177 | 4.6087e-06 | 1.0365e-08 | 7.374e-08 | -1.2349e-08 | 74.5842 | 56.2964 | 0.017741 | 3.475e-06 | 1.541e-08 | 6.1775e-08 | 9.0386e-09 | 74.7526 | -29.2235 | 3.4028e+38 | 3.4028e+38 | 1840.1473 | 2300.1079 | -1380.0137 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 10.0004 | 1840.1541 | 2300.1086 | -1380.0037 | 49.9999 | 36.8677 | 0.01416 | 4.6482e-06 | -3.465e-10 | 6.58e-08 | -1.0407e-08 | 74.5847 | 93.4138 | 0.015298 | 3.4529e-06 | 5.3175e-09 | 5.2792e-08 | -9.3123e-10 | 74.6612 | 13.6581 | 3.4028e+38 | 3.4028e+38 | 1840.1539 | 2300.1086 | -1380.0038 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 10.0004 | 1840.1486 | 2300.1064 | -1380.0148 | 50 | 36.8677 | 0.017324 | 4.7645e-06 | -7.5842e-09 | 8.2614e-08 | -3.1113e-08 | 74.3265 | 102.4405 | 0.019039 | 3.4906e-06 | 1.4235e-08 | 6.4124e-08 | -1.2212e-09 | 74.7826 | 4.3999 | 3.4028e+38 | 3.4028e+38 | 1840.1484 | 2300.1064 | -1380.0149 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 10.0004 | 1840.1459 | 2300.105 | -1380.0142 | 49.9999 | 36.8678 | 0.016865 | 4.6455e-06 | 6.2464e-09 | 8.0007e-08 | -2.2139e-08 | 74.4239 | 72.2456 | 0.018613 | 3.5617e-06 | 1.6175e-08 | 6.567e-08 | 6.139e-09 | 74.7633 | -21.0769 | 3.4028e+38 | 3.4028e+38 | 1840.1459 | 2300.105 | -1380.0143 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 10.0004 | 1840.1456 | 2300.1021 | -1380.0056 | 49.9999 | 36.8679 | 0.80003 | 3.4028e+38 | 0 | 230.0029 | 1840.2189 | 2300.116 | -1379.907 | 0.0060855 | 0.017706 | 4.691e-06 | 7.7789e-09 | 8.3276e-08 | -2.2845e-08 | 74.6798 | 74.9566 | 0.075049 | 3.4028e+38 | -38.4628 | 0.020158 | 0 | 0 | 0 | -0.00038389 | 0.01942 | 3.4266e-06 | 1.8101e-08 | 6.6666e-08 | 4.2138e-09 | 74.6718 | -12.6352 | 0.27269 | 3.4028e+38 | -27.3913 | 229.4373 | 2.4531e-09 | 0.00078085 | 1.5267e-10 | 2.6163e-05 | 3.4028e+38 | 3.4028e+38 | 1840.1499 | 2300.1077 | -1380.0145 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1840.2406 | 1840.2407 | -1379.9186 | 0.0061816 | 230.0019 | 10.0004 | 1840.1492 | 2300.1076 | -1380.0102 | 49.9999 | 36.8677 | [17 empty] | 22.0166 | 30.4612 | 58.5287 | 51.9713 | 61.9774 |  | -3.6495e-05
100 | 230.0021 | 10.0004 | 1840.1483 | 2300.1047 | -1380.006 | 49.9999 | 36.8677 | 0.015383 | 4.6865e-06 | -9.0012e-09 | 7.173e-08 | -2.0533e-08 | 74.6787 | 112.0735 | 0.016619 | 3.4473e-06 | 5.418e-09 | 5.7752e-08 | -5.6986e-09 | 74.7043 | 39.9706 | 3.4028e+38 | 3.4028e+38 | 1840.1482 | 2300.1047 | -1380.0061 | 3.4028e+38 | 3.4028e+38 | 230.002 | 10.0004 | 1840.1541 | 2300.1096 | -1380.0061 | 50 | 36.8677 | 0.016555 | 4.7077e-06 | -7.4274e-09 | 7.8184e-08 | -2.8329e-08 | 74.5424 | 105.2806 | 0.018101 | 3.4228e-06 | 9.7355e-09 | 6.212e-08 | -3.4954e-09 | 74.8223 | 11.924 | 3.4028e+38 | 3.4028e+38 | 1840.1539 | 2300.1096 | -1380.0062 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 10.0004 | 1840.1448 | 2300.1042 | -1380.0144 | 50 | 36.8679 | 0.014824 | 4.6474e-06 | 7.8324e-09 | 6.7656e-08 | -7.0577e-09 | 74.6457 | 45.8724 | 0.015927 | 3.5353e-06 | 7.8955e-09 | 5.4753e-08 | 4.1457e-09 | 74.6274 | -25.8659 | 3.4028e+38 | 3.4028e+38 | 1840.1448 | 2300.1042 | -1380.0145 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 10.0004 | 1840.1555 | 2300.1094 | -1380.0055 | 49.9999 | 36.8677 | 0.01466 | 4.657e-06 | -2.338e-09 | 6.8248e-08 | -1.1545e-08 | 74.3292 | 105.6538 | 0.015236 | 3.4583e-06 | 5.0307e-09 | 5.4316e-08 | -2.0364e-09 | 74.6252 | 23.8367 | 3.4028e+38 | 3.4028e+38 | 1840.1554 | 2300.1094 | -1380.0056 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 10.0004 | 1840.1458 | 2300.1069 | -1380.0125 | 49.9999 | 36.8679 | 0.017702 | 4.7744e-06 | -1.9394e-09 | 8.393e-08 | -2.7537e-08 | 74.4365 | 99.3688 | 0.019368 | 3.3838e-06 | 1.6125e-08 | 6.5422e-08 | 1.4765e-09 | 74.6668 | -5.3398 | 3.4028e+38 | 3.4028e+38 | 1840.1456 | 2300.1069 | -1380.0126 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 10.0004 | 1840.1525 | 2300.1074 | -1380.0038 | 50 | 36.8677 | 0.80003 | 3.4028e+38 | 0 | 230.0066 | 1840.2448 | 2300.1592 | -1379.9327 | 0.0072338 | 0.017074 | 4.7202e-06 | -8.7909e-09 | 8.083e-08 | -2.889e-08 | 74.4088 | 103.5182 | -0.083414 | 3.4028e+38 | -32.6109 | 0.019874 | 0 | 0 | 0 | -0.00037359 | 0.017392 | 3.4829e-06 | 8.2535e-09 | 6.05e-08 | -4.3295e-09 | 74.7723 | 35.5649 | 0.11571 | 3.4028e+38 | -36.9191 | 229.7657 | 1.612e-09 | 0.00080775 | -1.2511e-09 | 4.3073e-05 | 3.4028e+38 | 3.4028e+38 | 1840.1564 | 2300.1111 | -1380.0083 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1840.1968 | 1840.1968 | -1379.9686 | -0.00041716 | 230.002 | 10.0004 | 1840.1497 | 2300.107 | -1380.0089 | 50 | 36.8678 | [17 empty] | 22.9397 | 35.7047 | 59.6202 | 57.2439 | 74.8614 |  | 0.00036565
101 | 230.002 | 10.0004 | 1840.1558 | 2300.1145 | -1380.0104 | 50 | 36.8677 | 0.014463 | 4.6945e-06 | -3.0721e-09 | 6.7791e-08 | -1.151e-08 | 74.5133 | 101.3039 | 0.015632 | 3.4207e-06 | 5.5676e-09 | 5.3114e-08 | -1.8853e-09 | 74.7433 | 17.1527 | 3.4028e+38 | 3.4028e+38 | 1840.1556 | 2300.1145 | -1380.0105 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 10.0004 | 1840.1533 | 2300.1072 | -1380.0039 | 50 | 36.8678 | 0.015534 | 4.6819e-06 | -6.4418e-09 | 7.2738e-08 | -1.7913e-08 | 74.6171 | 109.7787 | 0.016446 | 3.4031e-06 | 6.5384e-09 | 5.5866e-08 | -5.763e-09 | 74.7032 | 41.7332 | 3.4028e+38 | 3.4028e+38 | 1840.1532 | 2300.1072 | -1380.004 | 3.4028e+38 | 3.4028e+38 | 230.002 | 10.0004 | 1840.1526 | 2300.1067 | -1380.0049 | 50 | 36.8677 | 0.016313 | 4.73e-06 | -9.529e-09 | 7.7045e-08 | -2.4835e-08 | 74.6318 | 111.2841 | 0.017422 | 3.4741e-06 | 1.026e-08 | 5.9794e-08 | -4.5313e-09 | 74.7305 | 22.0258 | 3.4028e+38 | 3.4028e+38 | 1840.1525 | 2300.1067 | -1380.005 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 10.0004 | 1840.1445 | 2300.1038 | -1380.0093 | 49.9999 | 36.8679 | 0.014727 | 4.5979e-06 | 4.1313e-09 | 6.6476e-08 | -9.3652e-09 | 74.6562 | 62.743 | 0.015875 | 3.601e-06 | 7.9219e-09 | 5.6227e-08 | 4.1148e-09 | 74.7585 | -24.42 | 3.4028e+38 | 3.4028e+38 | 1840.1445 | 2300.1038 | -1380.0094 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 10.0004 | 1840.1537 | 2300.1096 | -1380.0094 | 50 | 36.8676 | 0.01595 | 4.7148e-06 | -8.1344e-09 | 7.4104e-08 | -2.1286e-08 | 74.723 | 111.8656 | 0.017011 | 3.4214e-06 | 5.9525e-09 | 6.0129e-08 | -5.923e-09 | 74.6761 | 44.8578 | 3.4028e+38 | 3.4028e+38 | 1840.1536 | 2300.1096 | -1380.0095 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 10.0004 | 1840.1555 | 2300.1072 | -1380.0087 | 49.9999 | 36.8678 | 0.80003 | 3.4028e+38 | 0 | 230.0071 | 1840.2322 | 2300.1624 | -1379.9495 | 0.0049919 | 0.01597 | 4.7123e-06 | -9.0616e-09 | 7.5754e-08 | -2.6693e-08 | 74.5817 | 111.1991 | -0.12348 | 3.4028e+38 | 46.6676 | 0.020309 | 0 | 0 | 0 | -0.00037651 | 0.017559 | 3.4908e-06 | 7.9248e-09 | 6.0294e-08 | -5.4841e-09 | 74.7222 | 27.5443 | 0.14028 | 3.4028e+38 | -13.4797 | 229.1594 | -8.7736e-11 | 0.00078335 | -1.5987e-09 | 3.7313e-05 | 3.4028e+38 | 3.4028e+38 | 1840.1606 | 2300.113 | -1380.0029 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1840.2367 | 1840.2367 | -1379.9452 | 0.0060797 | 230.002 | 10.0004 | 1840.152 | 2300.1083 | -1380.0076 | 50 | 36.8677 | [17 empty] | 23.2652 | 35.513 | 60.5348 | 57.2614 | 72.8943 |  | 0.00028441
102 | 230.0018 | 6.0003 | -0.072236 | 1380.0815 | 1380.0818 | 50 | -90.003 | 0.016957 | 3.9743e-06 | 1.4662e-08 | 6.8624e-08 | 1.1954e-08 | 74.6652 | -37.5327 | 0.018689 | 3.3007e-06 | 6.5998e-10 | 6.0551e-08 | 7.8875e-09 | 74.7639 | -85.217 | 3.4028e+38 | 3.4028e+38 | -0.072236 | 1380.0815 | 1380.0818 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 6.0003 | -0.073101 | 1380.076 | 1380.0771 | 49.9999 | -90.0029 | 0.017703 | 3.9154e-06 | 1.3934e-08 | 6.9987e-08 | 1.8214e-08 | 74.7535 | -49.5918 | 0.019708 | 3.271e-06 | -4.243e-10 | 6.4248e-08 | 8.9877e-09 | 74.7602 | -83.7109 | 3.4028e+38 | 3.4028e+38 | -0.073101 | 1380.076 | 1380.0771 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 6.0003 | -0.069748 | 1380.0837 | 1380.0851 | 50 | -90.003 | 0.014462 | 4.0176e-06 | 3.8264e-09 | 5.8196e-08 | 7.6973e-09 | 74.6152 | -61.2958 | 0.015653 | 3.1512e-06 | -5.8558e-10 | 5.0979e-08 | 2.1419e-09 | 74.7646 | -134.8059 | 3.4028e+38 | 3.4028e+38 | -0.069748 | 1380.0837 | 1380.0851 | 3.4028e+38 | 3.4028e+38 | 230.002 | 6.0003 | -0.069636 | 1380.0875 | 1380.0881 | 50 | -90.003 | 0.015334 | 4.0503e-06 | 1.1199e-08 | 6.2245e-08 | 8.1767e-09 | 74.6878 | -34.3993 | 0.016452 | 3.1852e-06 | 2.6592e-09 | 5.4064e-08 | 5.6147e-09 | 74.7919 | -63.8624 | 3.4028e+38 | 3.4028e+38 | -0.069636 | 1380.0875 | 1380.0881 | 3.4028e+38 | 3.4028e+38 | 230.0025 | 6.0003 | -0.067056 | 1380.0853 | 1380.0878 | 50 | -90.0028 | 0.015894 | 4.0202e-06 | 1.3314e-08 | 6.5271e-08 | 9.7847e-09 | 74.6073 | -35.952 | 0.017261 | 3.2602e-06 | 4.0804e-09 | 5.4905e-08 | 7.6459e-09 | 74.7073 | -67.4718 | 3.4028e+38 | 3.4028e+38 | -0.067056 | 1380.0853 | 1380.0878 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0026 | 6.0003 | -0.070688 | 1380.0825 | 1380.0837 | 50 | -90.0029 | -4.9079e-05 | 3.4028e+38 | 0 | 230.0071 | -0.1078 | 1380.111 | 1380.0831 | -0.0039607 | 0.016841 | 3.9783e-06 | 1.7087e-08 | 6.7625e-08 | 1.4032e-08 | 74.7334 | -39.5682 | 0.25381 | 3.4028e+38 | 27.8891 | 0.021026 | 0 | 0 | 0 | 0.00017198 | 0.019026 | 3.1949e-06 | 1.1219e-09 | 6.0813e-08 | 8.261e-09 | 74.8007 | -84.0617 | 0.013545 | 3.4028e+38 | -10.6097 | 229.1828 | -2.4091e-09 | 0.00074151 | 6.5551e-10 | -0.00010764 | 3.4028e+38 | 3.4028e+38 | -0.067843 | 1380.0763 | 1380.0771 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -0.13254 | -0.13254 | 1380.0604 | -0.005072 | 230.0021 | 6.0003 | -0.070356 | 1380.0828 | 1380.084 | 50 | -90.0029 | [17 empty] | 23.6492 | 67.3898 | 17324420.6371 | 91.3242 | 92.1503 |  | -0.0027764
103 | 230.0022 | 6.0003 | -0.069118 | 1380.0797 | 1380.0834 | 49.9999 | -90.0028 | 0.01705 | 4.0152e-06 | 8.1327e-09 | 6.7445e-08 | 1.7095e-08 | 74.8156 | -63.0747 | 0.018568 | 3.1803e-06 | -1.3257e-09 | 6.0914e-08 | 6.1418e-09 | 74.6722 | -102.2739 | 3.4028e+38 | 3.4028e+38 | -0.069118 | 1380.0797 | 1380.0834 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 6.0003 | -0.065921 | 1380.0771 | 1380.0803 | 49.9999 | -90.003 | 0.017739 | 3.9231e-06 | 1.4659e-08 | 7.0292e-08 | 1.8002e-08 | 74.7247 | -50.427 | 0.019477 | 3.2133e-06 | 5.9091e-10 | 6.2267e-08 | 9.7975e-09 | 74.5978 | -83.8412 | 3.4028e+38 | 3.4028e+38 | -0.065921 | 1380.0771 | 1380.0803 | 3.4028e+38 | 3.4028e+38 | 230.0028 | 6.0003 | -0.067131 | 1380.0811 | 1380.0835 | 49.9999 | -90.0029 | 0.017081 | 4.0165e-06 | 7.6418e-09 | 6.6469e-08 | 1.4197e-08 | 74.6515 | -63.2909 | 0.018819 | 3.244e-06 | -1.995e-09 | 5.9603e-08 | 5.7126e-09 | 74.7178 | -115.8244 | 3.4028e+38 | 3.4028e+38 | -0.067131 | 1380.0811 | 1380.0835 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 6.0003 | -0.065883 | 1380.0754 | 1380.0867 | 49.9999 | -90.0029 | 0.01651 | 3.9651e-06 | 6.966e-09 | 6.5465e-08 | 1.3e-08 | 74.6455 | -64.5683 | 0.018482 | 3.2641e-06 | -2.7877e-09 | 5.872e-08 | 4.2993e-09 | 74.7933 | -121.3352 | 3.4028e+38 | 3.4028e+38 | -0.065883 | 1380.0754 | 1380.0867 | 3.4028e+38 | 3.4028e+38 | 230.002 | 6.0003 | -0.071716 | 1380.0792 | 1380.0793 | 50 | -90.0029 | 0.016802 | 3.9475e-06 | 8.0101e-09 | 6.6891e-08 | 1.7024e-08 | 74.7524 | -64.2232 | 0.018743 | 3.1847e-06 | -3.0892e-09 | 6.0898e-08 | 6.7131e-09 | 74.7709 | -97.9079 | 3.4028e+38 | 3.4028e+38 | -0.071716 | 1380.0792 | 1380.0793 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0017 | 6.0003 | -0.070284 | 1380.0757 | 1380.0813 | 49.9999 | -90.0029 | -5.1069e-05 | 3.4028e+38 | 0 | 230.0007 | -0.12589 | 1380.0732 | 1380.047 | -0.0042879 | 0.016608 | 3.9434e-06 | 7.9684e-09 | 6.549e-08 | 1.5852e-08 | 74.701 | -62.7919 | 0.1221 | 3.4028e+38 | -18.0459 | 0.019712 | 0 | 0 | 0 | 0.00018993 | 0.017249 | 3.2328e-06 | -9.7759e-10 | 5.5695e-08 | 4.7665e-09 | 74.7586 | -103.1433 | -0.020329 | 3.4028e+38 | -16.0373 | 229.4752 | 4.3921e-10 | 0.00072765 | -7.0566e-10 | -9.8477e-05 | 3.4028e+38 | 3.4028e+38 | -0.066239 | 1380.0762 | 1380.0872 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -0.11176 | -0.10658 | 1380.0582 | -0.0013736 | 230.0021 | 6.0003 | -0.067954 | 1380.0785 | 1380.0826 | 49.9999 | -90.0029 | [17 empty] | 23.2421 | 66.1758 | 7048989.4734 | 89.0303 | 92.0143 |  | -0.0025608
104 | 230.0021 | 6.0003 | -0.068154 | 1380.0732 | 1380.075 | 49.9999 | -90.0028 | 0.016967 | 3.9818e-06 | 1.6905e-08 | 6.6341e-08 | 1.274e-08 | 74.7399 | -38.1044 | 0.018382 | 3.2777e-06 | 2.4857e-09 | 6.0169e-08 | 7.4757e-09 | 74.6847 | -73.2756 | 3.4028e+38 | 3.4028e+38 | -0.068154 | 1380.0732 | 1380.075 | 3.4028e+38 | 3.4028e+38 | 230.002 | 6.0003 | -0.074782 | 1380.08 | 1380.0809 | 49.9999 | -90.0029 | 0.016448 | 3.9551e-06 | 5.4086e-09 | 6.6432e-08 | 1.3863e-08 | 74.686 | -67.6178 | 0.018194 | 3.2e-06 | -1.8801e-09 | 5.959e-08 | 4.9567e-09 | 74.8037 | -117.135 | 3.4028e+38 | 3.4028e+38 | -0.074782 | 1380.08 | 1380.0809 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 6.0003 | -0.067136 | 1380.0811 | 1380.084 | 49.9999 | -90.0027 | 0.01541 | 3.994e-06 | 1.386e-08 | 6.1567e-08 | 8.4711e-09 | 74.7356 | -33.2456 | 0.016019 | 3.1518e-06 | 3.2448e-09 | 5.2544e-08 | 6.352e-09 | 74.7366 | -58.2712 | 3.4028e+38 | 3.4028e+38 | -0.067136 | 1380.0811 | 1380.084 | 3.4028e+38 | 3.4028e+38 | 230.0024 | 6.0003 | -0.064864 | 1380.0835 | 1380.0852 | 50 | -90.0029 | 0.01597 | 3.9739e-06 | 1.5466e-08 | 6.3488e-08 | 9.565e-09 | 74.7255 | -32.7323 | 0.017515 | 3.2668e-06 | 2.4297e-09 | 5.462e-08 | 6.3442e-09 | 74.7769 | -64.8003 | 3.4028e+38 | 3.4028e+38 | -0.064864 | 1380.0835 | 1380.0852 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 6.0003 | -0.068028 | 1380.0802 | 1380.0823 | 49.9999 | -90.0029 | 0.017408 | 3.97e-06 | 1.562e-08 | 6.905e-08 | 1.223e-08 | 74.7607 | -39.6471 | 0.019196 | 3.2133e-06 | 2.7558e-09 | 6.1019e-08 | 9.7393e-09 | 74.7367 | -73.0458 | 3.4028e+38 | 3.4028e+38 | -0.068028 | 1380.0802 | 1380.0823 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 6.0003 | -0.07026 | 1380.0782 | 1380.0773 | 50 | -90.0028 | -4.8232e-05 | 3.4028e+38 | 0 | 230.0069 | -0.12632 | 1380.1067 | 1380.0782 | -0.0038958 | 0.017307 | 3.999e-06 | 1.6398e-08 | 6.9159e-08 | 1.5732e-08 | 74.6625 | -40.6146 | 0.25845 | 3.4028e+38 | 14.8033 | 0.020907 | 0 | 0 | 0 | 0.00017549 | 0.018989 | 3.2813e-06 | 1.6632e-09 | 6.2377e-08 | 8.1857e-09 | 74.8013 | -89.3609 | 0.001499 | 3.4028e+38 | -52.4203 | 229.1409 | -8.3515e-10 | 0.00074002 | 1.5305e-09 | -0.00010583 | 3.4028e+38 | 3.4028e+38 | -0.067125 | 1380.0803 | 1380.0806 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -0.13118 | -0.13118 | 1380.0682 | -0.0048514 | 230.0022 | 6.0003 | -0.068593 | 1380.0796 | 1380.0815 | 49.9999 | -90.0029 | [17 empty] | 23.5962 | 73.43 | -139500741.0119 | 96.7175 | 98.0739 |  | -0.0028724
105 | 230.0021 | 6.0003 | 1195.1552 | 1380.0806 | 690.098 | 50 | -30.0031 | 0.017511 | 4.0042e-06 | 1.5884e-08 | 6.9999e-08 | -1.8102e-09 | 74.6726 | 6.0183 | 0.019451 | 3.2938e-06 | 1.1209e-08 | 6.2662e-08 | 1.0653e-08 | 74.7995 | -45.9124 | 3.4028e+38 | 3.4028e+38 | 1195.1552 | 1380.0806 | 690.098 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 6.0003 | 1195.1561 | 1380.0848 | 690.1098 | 49.9999 | -30.0031 | 0.015935 | 4.0668e-06 | 1.0957e-08 | 6.336e-08 | -3.4101e-09 | 74.6247 | 18.9188 | 0.017093 | 3.2818e-06 | 7.2737e-09 | 5.5968e-08 | 4.9597e-09 | 74.765 | -34.0194 | 3.4028e+38 | 3.4028e+38 | 1195.1561 | 1380.0848 | 690.1097 | 3.4028e+38 | 3.4028e+38 | 230.002 | 6.0003 | 1195.1562 | 1380.0854 | 690.1026 | 50 | -30.0029 | 0.016428 | 4.0591e-06 | 9.9177e-09 | 6.5819e-08 | -5.3433e-09 | 74.7254 | 30.071 | 0.018073 | 3.2669e-06 | 8.3015e-09 | 5.5622e-08 | 5.0857e-09 | 74.6171 | -37.1745 | 3.4028e+38 | 3.4028e+38 | 1195.1562 | 1380.0854 | 690.1025 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 6.0003 | 1195.1503 | 1380.0742 | 690.0978 | 50 | -30.0029 | 0.016687 | 3.9578e-06 | 1.3916e-08 | 6.5731e-08 | 6.6618e-09 | 74.7343 | -25.3516 | 0.01804 | 3.2204e-06 | 7.5453e-09 | 6.1865e-08 | 1.1216e-08 | 74.7995 | -57.6069 | 3.4028e+38 | 3.4028e+38 | 1195.1503 | 1380.0742 | 690.0978 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 6.0003 | 1195.1471 | 1380.074 | 690.0967 | 49.9999 | -30.0028 | 0.016236 | 3.9533e-06 | 1.5485e-08 | 6.4164e-08 | 7.841e-09 | 74.6266 | -28.3955 | 0.017595 | 3.3119e-06 | 5.3151e-09 | 5.9162e-08 | 1.0567e-08 | 74.7011 | -59.6026 | 3.4028e+38 | 3.4028e+38 | 1195.1471 | 1380.074 | 690.0967 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0023 | 6.0003 | 1195.1517 | 1380.0779 | 690.0992 | 49.9999 | -30.003 | 0.866 | 3.4028e+38 | 0 | 230.0011 | 1195.1163 | 1380.0726 | 690.1259 | -0.0038181 | 0.01666 | 3.9633e-06 | 1.3259e-08 | 6.6034e-08 | 6.0824e-09 | 74.7865 | -22.8543 | 0.21893 | 3.4028e+38 | -37.9002 | 0.019665 | 0 | 0 | 0 | -3.0463e-05 | 0.018422 | 3.3296e-06 | 5.7941e-09 | 5.965e-08 | 1.0007e-08 | 74.6894 | -59.9287 | 0.097135 | 3.4028e+38 | -82.9356 | 229.0547 | 7.1586e-10 | 0.00074124 | 1.6019e-09 | -9.3401e-05 | 3.4028e+38 | 3.4028e+38 | 1195.151 | 1380.0795 | 690.1002 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1195.1119 | 1195.1119 | 690.1289 | -0.0043437 | 230.002 | 6.0003 | 1195.153 | 1380.0798 | 690.101 | 50 | -30.0029 | [17 empty] | 22.3998 | 7.5481 | 4.9647 | 28.6619 | 109.9309 |  | -0.0027069
106 | 230.0019 | 6.0003 | 1195.1521 | 1380.08 | 690.1014 | 49.9999 | -30.0028 | 0.017654 | 3.9577e-06 | 1.5613e-08 | 6.9317e-08 | 3.9426e-09 | 74.6669 | -13.4102 | 0.019313 | 3.2668e-06 | 8.4834e-09 | 6.3156e-08 | 1.1468e-08 | 74.748 | -54.6102 | 3.4028e+38 | 3.4028e+38 | 1195.1521 | 1380.08 | 690.1014 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 6.0003 | 1195.1494 | 1380.0782 | 690.1014 | 50 | -30.0028 | 0.017553 | 4.0141e-06 | 1.5547e-08 | 7.1305e-08 | -9.543e-11 | 74.7866 | -9.8856 | 0.01937 | 3.3477e-06 | 1.132e-08 | 6.4527e-08 | 1.2318e-08 | 74.7769 | -47.7632 | 3.4028e+38 | 3.4028e+38 | 1195.1494 | 1380.0782 | 690.1014 | 3.4028e+38 | 3.4028e+38 | 230.0026 | 6.0003 | 1195.1516 | 1380.0819 | 690.1086 | 49.9999 | -30.0032 | 0.015308 | 3.9947e-06 | 8.8928e-09 | 5.979e-08 | 7.6785e-09 | 74.6892 | -37.181 | 0.016305 | 3.2888e-06 | 2.8471e-09 | 5.2721e-08 | 5.6988e-09 | 74.689 | -63.7694 | 3.4028e+38 | 3.4028e+38 | 1195.1516 | 1380.0819 | 690.1086 | 3.4028e+38 | 3.4028e+38 | 230.002 | 6.0003 | 1195.1516 | 1380.0804 | 690.1003 | 50 | -30.0028 | 0.017024 | 3.9755e-06 | 1.7438e-08 | 6.7529e-08 | 8.1447e-09 | 74.6681 | -21.9742 | 0.018831 | 3.3495e-06 | 5.7423e-09 | 6.3514e-08 | 1.0064e-08 | 74.6701 | -60.347 | 3.4028e+38 | 3.4028e+38 | 1195.1516 | 1380.0804 | 690.1003 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 6.0003 | 1195.1514 | 1380.0778 | 690.0996 | 49.9999 | -30.0027 | 0.017425 | 3.905e-06 | 1.9491e-08 | 6.8183e-08 | 7.6664e-09 | 74.7242 | -21.7269 | 0.019128 | 3.244e-06 | 7.4545e-09 | 6.208e-08 | 1.2182e-08 | 74.7403 | -56.7265 | 3.4028e+38 | 3.4028e+38 | 1195.1514 | 1380.0778 | 690.0996 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 6.0003 | 1195.1487 | 1380.0798 | 690.1082 | 50 | -30.0028 | 0.866 | 3.4028e+38 | 0 | 230.0014 | 1195.1189 | 1380.074 | 690.1271 | -0.0036382 | 0.017155 | 3.9736e-06 | 1.6661e-08 | 6.8093e-08 | 4.9357e-09 | 74.7522 | -16.4201 | 0.24695 | 3.4028e+38 | -48.9748 | 0.021184 | 0 | 0 | 0 | -3.4368e-05 | 0.019169 | 3.3436e-06 | 6.4977e-09 | 6.2697e-08 | 9.7603e-09 | 74.6726 | -54.038 | 0.094006 | 3.4028e+38 | -64.3847 | 229.1946 | -1.5312e-09 | 0.00074306 | 2.6264e-10 | -9.61e-05 | 3.4028e+38 | 3.4028e+38 | 1195.1473 | 1380.0745 | 690.0984 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1195.1288 | 1195.1288 | 690.1438 | -0.0029217 | 230.002 | 6.0003 | 1195.1512 | 1380.0797 | 690.1023 | 49.9999 | -30.0029 | [17 empty] | 21.766 | 5.3634 | 2.4212 | 26.754 | 107.8349 |  | -0.0025544
107 | 230.0023 | 6.0003 | 1195.1575 | 1380.0814 | 690.103 | 49.9999 | -30.0029 | 0.017624 | 3.9688e-06 | 1.8941e-08 | 7.0541e-08 | 3.6253e-09 | 74.579 | -9.937 | 0.019632 | 3.2614e-06 | 1.0072e-08 | 6.4275e-08 | 1.1258e-08 | 74.8016 | -48.4958 | 3.4028e+38 | 3.4028e+38 | 1195.1575 | 1380.0814 | 690.103 | 3.4028e+38 | 3.4028e+38 | 230.002 | 6.0003 | 1195.1552 | 1380.0808 | 690.098 | 49.9999 | -30.0027 | 0.016677 | 4.0062e-06 | 1.2241e-08 | 6.6929e-08 | -2.8324e-09 | 74.6987 | 19.6359 | 0.018233 | 3.2655e-06 | 8.6273e-09 | 5.8274e-08 | 7.1436e-09 | 74.8057 | -35.8964 | 3.4028e+38 | 3.4028e+38 | 1195.1552 | 1380.0808 | 690.0979 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 6.0003 | 1195.1447 | 1380.0743 | 690.1044 | 49.9999 | -30.0028 | 0.017862 | 3.9792e-06 | 1.906e-08 | 7.1144e-08 | 3.936e-09 | 74.73 | -11.9469 | 0.019293 | 3.3139e-06 | 9.5145e-09 | 6.4004e-08 | 1.1155e-08 | 74.7242 | -49.7578 | 3.4028e+38 | 3.4028e+38 | 1195.1447 | 1380.0743 | 690.1044 | 3.4028e+38 | 3.4028e+38 | 230.0024 | 6.0003 | 1195.1554 | 1380.0808 | 690.0998 | 49.9999 | -30.0028 | 0.016906 | 4.0518e-06 | 1.0502e-08 | 6.7469e-08 | -4.2809e-09 | 74.7245 | 24.1707 | 0.018744 | 3.179e-06 | 9.4539e-09 | 6.0724e-08 | 8.5173e-09 | 74.6568 | -36.8734 | 3.4028e+38 | 3.4028e+38 | 1195.1554 | 1380.0808 | 690.0997 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 6.0003 | 1195.1532 | 1380.083 | 690.1049 | 50 | -30.0029 | 0.014621 | 3.9708e-06 | 9.1614e-09 | 5.8086e-08 | 4.8047e-09 | 74.6943 | -27.9187 | 0.015685 | 3.1968e-06 | 3.1468e-09 | 5.0144e-08 | 4.7454e-09 | 74.6921 | -54.9184 | 3.4028e+38 | 3.4028e+38 | 1195.1532 | 1380.083 | 690.1049 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 6.0003 | 1195.1542 | 1380.0835 | 690.1036 | 49.9999 | -30.0029 | 0.866 | 3.4028e+38 | 0 | 230.0017 | 1195.1333 | 1380.0767 | 690.1154 | -0.0022572 | 0.014718 | 3.9586e-06 | 1.0383e-08 | 5.7707e-08 | 4.7073e-09 | 74.6614 | -27.6459 | 0.1435 | 3.4028e+38 | -44.5753 | 0.019746 | 0 | 0 | 0 | -1.2263e-05 | 0.015688 | 3.1866e-06 | 1.2946e-09 | 5.0044e-08 | 3.8755e-09 | 74.6194 | -59.774 | 0.04065 | 3.4028e+38 | 55.5757 | 229.1679 | -6.9787e-10 | 0.00076051 | -9.9777e-10 | -8.3265e-05 | 3.4028e+38 | 3.4028e+38 | 1195.1492 | 1380.0793 | 690.1045 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1195.1389 | 1195.1392 | 690.1227 | -0.0017444 | 230.0022 | 6.0003 | 1195.1532 | 1380.0801 | 690.102 | 49.9999 | -30.0028 | [17 empty] | 23.4754 | 8.6765 | 4.5533 | 30.4998 | 119.7316 |  | -0.0027885
108 | 230.0021 | 6.0003 | -828.1167 | 1380.0787 | 1104.0228 | 49.9999 | -126.8732 | 0.016223 | 3.9287e-06 | 9.4248e-09 | 6.4585e-08 | 1.4498e-08 | 74.6192 | -57.5332 | 0.017415 | 3.1428e-06 | -2.4251e-09 | 5.77e-08 | 4.2769e-09 | 74.7723 | -111.0483 | 3.4028e+38 | 3.4028e+38 | -828.1168 | 1380.0787 | 1104.0228 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 6.0003 | -828.1223 | 1380.078 | 1104.0193 | 50 | -126.8734 | 0.017719 | 3.956e-06 | 2.7739e-09 | 7.0243e-08 | 1.5383e-08 | 74.656 | -76.548 | 0.019488 | 3.1796e-06 | -5.241e-09 | 6.1964e-08 | 3.5774e-09 | 74.7558 | -144.3068 | 3.4028e+38 | 3.4028e+38 | -828.1223 | 1380.078 | 1104.0193 | 3.4028e+38 | 3.4028e+38 | 230.002 | 6.0003 | -828.1223 | 1380.0782 | 1104.0206 | 49.9999 | -126.8733 | 0.017525 | 3.978e-06 | 5.6838e-09 | 6.9498e-08 | 1.8701e-08 | 74.7403 | -74.8225 | 0.019356 | 3.2583e-06 | -5.0814e-09 | 6.1177e-08 | 5.1024e-09 | 74.7782 | -118.4583 | 3.4028e+38 | 3.4028e+38 | -828.1223 | 1380.0782 | 1104.0206 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 6.0003 | -828.1262 | 1380.0837 | 1104.0222 | 49.9999 | -126.8733 | 0.014722 | 4.056e-06 | -2.4852e-09 | 5.9746e-08 | 7.4001e-09 | 74.6842 | -118.2343 | 0.015891 | 3.1894e-06 | -3.5133e-09 | 5.0696e-08 | -7.2702e-10 | 74.7903 | 169.4692 | 3.4028e+38 | 3.4028e+38 | -828.1263 | 1380.0837 | 1104.0221 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 6.0003 | -828.1166 | 1380.0831 | 1104.0254 | 49.9999 | -126.8733 | 0.01435 | 3.9823e-06 | 2.2955e-09 | 5.8228e-08 | 7.7856e-09 | 74.6452 | -73.5724 | 0.015141 | 3.135e-06 | -2.3697e-09 | 4.8383e-08 | 1.0327e-09 | 74.7757 | -152.5296 | 3.4028e+38 | 3.4028e+38 | -828.1166 | 1380.0831 | 1104.0254 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0023 | 6.0003 | -828.1188 | 1380.0823 | 1104.0304 | 50 | -126.8734 | -0.60005 | 3.4028e+38 | 0 | 230.0036 | -828.1481 | 1380.0914 | 1104.0052 | 0.00098798 | 0.01442 | 4.0125e-06 | 1.2629e-09 | 5.7826e-08 | 6.2913e-09 | 74.7063 | -78.6845 | 0.022961 | 3.4028e+38 | -52.753 | 0.020069 | 0 | 0 | 0 | 0.00024696 | 0.015278 | 3.178e-06 | -2.3291e-09 | 4.8053e-08 | 6.7603e-10 | 74.6592 | -156.7715 | -0.02836 | 3.4028e+38 | 2.2509 | 229.8741 | 4.2847e-10 | 0.00072481 | 7.3501e-10 | -4.5776e-05 | 3.4028e+38 | 3.4028e+38 | -828.1202 | 1380.083 | 1104.0266 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -828.1412 | -828.1412 | 1104.0061 | 0.00056915 | 230.0021 | 6.0003 | -828.1208 | 1380.0804 | 1104.0221 | 49.9999 | -126.8733 | [17 empty] | 23.4287 | 55.8786 | 115.9422 | 79.5487 | 70.1252 |  | -0.0011834
109 | 230.0027 | 6.0003 | -828.126 | 1380.0731 | 1104.0142 | 50 | -126.8734 | 0.017267 | 4.0034e-06 | -1.6915e-10 | 6.8282e-08 | 1.4819e-08 | 74.6122 | -95.095 | 0.018827 | 3.1667e-06 | -5.9943e-09 | 6.172e-08 | 1.9512e-09 | 74.7288 | -152.7261 | 3.4028e+38 | 3.4028e+38 | -828.1261 | 1380.0731 | 1104.0142 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 6.0003 | -828.1162 | 1380.0789 | 1104.0247 | 49.9999 | -126.8732 | 0.015409 | 4.0239e-06 | 8.6488e-09 | 6.109e-08 | 1.1836e-08 | 74.7437 | -51.042 | 0.016534 | 3.1581e-06 | -1.4118e-09 | 5.2237e-08 | 2.9503e-09 | 74.7245 | -86.2062 | 3.4028e+38 | 3.4028e+38 | -828.1163 | 1380.0789 | 1104.0247 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 6.0003 | -828.1182 | 1380.0786 | 1104.0214 | 49.9999 | -126.8731 | 0.014545 | 4.0059e-06 | 4.9261e-09 | 5.8745e-08 | 8.3442e-09 | 74.7655 | -56.5799 | 0.015379 | 3.1876e-06 | -2.0139e-09 | 4.9389e-08 | 1.935e-09 | 74.7724 | -94.0477 | 3.4028e+38 | 3.4028e+38 | -828.1183 | 1380.0786 | 1104.0214 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 6.0003 | -828.1196 | 1380.0801 | 1104.0209 | 50 | -126.8734 | 0.017307 | 3.9748e-06 | 8.4182e-09 | 6.8768e-08 | 1.5554e-08 | 74.8052 | -66.2072 | 0.018911 | 3.2711e-06 | -3.6542e-09 | 6.0329e-08 | 4.7828e-09 | 74.7415 | -128.9176 | 3.4028e+38 | 3.4028e+38 | -828.1197 | 1380.0801 | 1104.0209 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 6.0003 | -828.1237 | 1380.0789 | 1104.0183 | 50 | -126.8736 | 0.015723 | 3.9508e-06 | -4.6949e-09 | 6.3643e-08 | 8.9934e-09 | 74.7042 | -115.5518 | 0.017091 | 3.2116e-06 | -5.194e-09 | 5.5508e-08 | -4.9732e-10 | 74.6944 | -178.5832 | 3.4028e+38 | 3.4028e+38 | -828.1237 | 1380.0789 | 1104.0182 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0022 | 6.0003 | -828.1256 | 1380.0841 | 1104.0253 | 49.9999 | -126.8733 | -0.60005 | 3.4028e+38 | 0 | 230.0003 | -828.1533 | 1380.0673 | 1103.9717 | -0.00075203 | 0.015703 | 4.0042e-06 | -4.7651e-09 | 6.2309e-08 | 8.6885e-09 | 74.6188 | -118.742 | -0.07501 | 3.4028e+38 | 20.6274 | 0.020532 | 0 | 0 | 0 | 0.00024036 | 0.017087 | 3.1463e-06 | -5.0191e-09 | 5.4222e-08 | -3.4005e-10 | 74.7837 | 170.6664 | -0.10173 | 3.4028e+38 | -6.6599 | 229.293 | -2.1993e-09 | 0.00073313 | -1.2417e-09 | -5.3385e-05 | 3.4028e+38 | 3.4028e+38 | -828.1221 | 1380.0867 | 1104.0265 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -828.1466 | -828.1466 | 1103.9836 | 8.2943e-06 | 230.0023 | 6.0003 | -828.1208 | 1380.0779 | 1104.0199 | 50 | -126.8734 | [17 empty] | 24.0597 | 67.576 | 115.7693 | 88.1847 | 84.4857 |  | -0.00080841
110 | 230.002 | 6.0003 | -828.1254 | 1380.08 | 1104.0188 | 50 | -126.8736 | 0.016612 | 3.983e-06 | -2.063e-09 | 6.6047e-08 | 1.3231e-08 | 74.651 | -99.8064 | 0.018273 | 3.1876e-06 | -5.8104e-09 | 5.6973e-08 | -3.9798e-10 | 74.7041 | 175.9224 | 3.4028e+38 | 3.4028e+38 | -828.1255 | 1380.08 | 1104.0187 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 6.0003 | -828.1209 | 1380.0793 | 1104.0211 | 50 | -126.8735 | 0.01756 | 3.9365e-06 | 3.3173e-10 | 6.921e-08 | 1.499e-08 | 74.6593 | -89.5392 | 0.019659 | 3.1895e-06 | -6.4865e-09 | 6.3109e-08 | 1.9333e-09 | 74.8002 | -156.6552 | 3.4028e+38 | 3.4028e+38 | -828.121 | 1380.0793 | 1104.0211 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 6.0003 | -828.1227 | 1380.0813 | 1104.0253 | 49.9999 | -126.8733 | 0.0144 | 4.0308e-06 | 2.5097e-09 | 5.8971e-08 | 7.9473e-09 | 74.6903 | -66.8255 | 0.01512 | 3.11e-06 | -1.7419e-09 | 4.8508e-08 | 7.8781e-10 | 74.7701 | -152.0338 | 3.4028e+38 | 3.4028e+38 | -828.1228 | 1380.0813 | 1104.0253 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 6.0003 | -828.1246 | 1380.0839 | 1104.0258 | 50 | -126.8735 | 0.014226 | 4.0927e-06 | 7.2884e-11 | 5.9092e-08 | 7.4277e-09 | 74.7013 | -87.5461 | 0.015379 | 3.1603e-06 | -3.0309e-09 | 4.8447e-08 | 2.0372e-11 | 74.7777 | 162.6713 | 3.4028e+38 | 3.4028e+38 | -828.1246 | 1380.0839 | 1104.0258 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 6.0003 | -828.121 | 1380.0841 | 1104.0254 | 50 | -126.8734 | 0.016825 | 3.9815e-06 | -2.731e-09 | 6.5429e-08 | 1.099e-08 | 74.6509 | -101.4461 | 0.018402 | 3.1804e-06 | -4.6141e-09 | 5.8592e-08 | 4.1987e-10 | 74.8076 | -175.3547 | 3.4028e+38 | 3.4028e+38 | -828.121 | 1380.0841 | 1104.0254 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0024 | 6.0003 | -828.1267 | 1380.083 | 1104.0225 | 50 | -126.8737 | -0.60006 | 3.4028e+38 | 0 | 230.0006 | -828.1588 | 1380.0706 | 1103.9689 | -0.0026734 | 0.017026 | 3.9469e-06 | -4.2311e-09 | 6.6281e-08 | 1.1286e-08 | 74.6763 | -116.6696 | -0.081036 | 3.4028e+38 | -3.4582 | 0.020841 | 0 | 0 | 0 | 0.00022926 | 0.018479 | 3.0904e-06 | -6.3971e-09 | 5.7067e-08 | -4.1962e-10 | 74.7763 | 175.9356 | -0.1126 | 3.4028e+38 | -84.7945 | 229.6088 | -1.7235e-09 | 0.00072711 | -3.1585e-10 | -5.9855e-05 | 3.4028e+38 | 3.4028e+38 | -828.1199 | 1380.0779 | 1104.021 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -828.1691 | -828.1691 | 1103.9666 | -0.0014406 | 230.002 | 6.0003 | -828.1229 | 1380.0817 | 1104.0233 | 50 | -126.8734 | [17 empty] | 22.0319 | 67.6682 | 114.9021 | 91.0744 | 89.7268 |  | -0.00073826
111 | 230.0033 | 50.0009 | 11500.3672 | 11500.373 | 0.3544 | 49.9999 | -0.0020149 | 0.017458 | 8.2398e-06 | 7.6247e-08 | 1.4445e-07 | 1.4424e-08 | 72.4119 | -11.6866 | 0.019618 | 8.0322e-06 | 5.0693e-08 | 1.5551e-07 | 8.042e-08 | 69.4184 | -57.6303 | 3.4028e+38 | 3.4028e+38 | 11500.3672 | 11500.373 | 0.3544 | 3.4028e+38 | 3.4028e+38 | 230.0031 | 50.001 | 11500.3438 | 11500.3496 | 0.43049 | 49.9999 | -0.002196 | 0.01445 | 8.483e-06 | 2.6308e-08 | 1.2157e-07 | -7.7594e-09 | 71.9417 | 9.1021 | 0.015541 | 7.8792e-06 | 3.2392e-08 | 1.2047e-07 | 2.5376e-08 | 72.0927 | -38.2854 | 3.4028e+38 | 3.4028e+38 | 11500.3438 | 11500.3496 | 0.43049 | 3.4028e+38 | 3.4028e+38 | 230.0031 | 50.0008 | 11500.3271 | 11500.335 | 0.41726 | 50 | -0.0020788 | 0.01605 | 8.1808e-06 | 6.0232e-08 | 1.3173e-07 | 2.9853e-08 | 73.539 | -27.2195 | 0.017656 | 8.195e-06 | 2.4781e-08 | 1.4317e-07 | 6.7109e-08 | 71.336 | -68.4186 | 3.4028e+38 | 3.4028e+38 | 11500.3271 | 11500.335 | 0.41726 | 3.4028e+38 | 3.4028e+38 | 230.0025 | 50.001 | 11500.3604 | 11500.3672 | 0.40977 | 49.9999 | -0.0020293 | 0.016576 | 8.5692e-06 | 5.3835e-08 | 1.4342e-07 | -1.4866e-08 | 71.4454 | 19.2146 | 0.018551 | 7.8315e-06 | 7.0106e-08 | 1.4264e-07 | 5.6673e-08 | 72.5757 | -41.2866 | 3.4028e+38 | 3.4028e+38 | 11500.3604 | 11500.3672 | 0.40977 | 3.4028e+38 | 3.4028e+38 | 230.0037 | 50.001 | 11500.334 | 11500.3418 | 0.43691 | 49.9999 | -0.0021497 | 0.014214 | 8.4215e-06 | 2.6315e-08 | 1.212e-07 | 3.2694e-09 | 72.3645 | -6.9221 | 0.015117 | 7.9903e-06 | 2.288e-08 | 1.2231e-07 | 2.4699e-08 | 68.0652 | -56.4369 | 3.4028e+38 | 3.4028e+38 | 11500.334 | 11500.3418 | 0.43691 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0036 | 50.0009 | 11500.375 | 11500.3828 | 0.45395 | 49.9999 | -0.0022616 | 1 | 3.4028e+38 | 0 | 230.0013 | 11500.2227 | 11500.2275 | 0.70134 | -0.012247 | 0.015776 | 8.1505e-06 | 5.4269e-08 | 1.2968e-07 | 2.8313e-08 | 72.385 | -27.5518 | 0.45589 | 3.4028e+38 | 5.4924 | 0.019683 | 0 | 0 | 0 | -0.00017038 | 0.018574 | 8.2584e-06 | 2.9902e-08 | 1.532e-07 | 8.0218e-08 | 69.9963 | -63.7148 | 0.30244 | 3.4028e+38 | -52.7358 | 229.729 | 4.391e-09 | 0.0018377 | 7.442e-09 | -0.00078517 | 3.4028e+38 | 3.4028e+38 | 11500.3047 | 11500.3096 | 0.41268 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 11500.4326 | 11500.4326 | 0.91261 | 0.015137 | 230.0031 | 50.0009 | 11500.3465 | 11500.3533 | 0.40977 | 49.9999 | -0.0020937 | [17 empty] | 29.7886 | -4.6659 | 22.8541 | 23.4547 | 3792122.1268 |  | -0.0016677
112 | 230.0048 | 50.0011 | 5749.981 | 11500.4834 | 9959.8916 | 49.9999 | -60.0018 | 0.015264 | 8.5942e-06 | 4.7802e-08 | 1.3146e-07 | 2.7135e-08 | 72.3756 | -29.5812 | 0.016507 | 7.7366e-06 | 1.1813e-08 | 1.2767e-07 | 4.8084e-08 | 72.1771 | -74.7591 | 3.4028e+38 | 3.4028e+38 | 5749.981 | 11500.4834 | 9959.8916 | 3.4028e+38 | 3.4028e+38 | 230.0049 | 50.0009 | 5749.9688 | 11500.4443 | 9959.8359 | 49.9999 | -60.0017 | 0.015678 | 8.3452e-06 | 2.3809e-08 | 1.3081e-07 | 4.6369e-08 | 72.0962 | -66.4772 | 0.017114 | 7.9494e-06 | -3.3085e-08 | 1.324e-07 | 1.9227e-08 | 73.4451 | -155.7281 | 3.4028e+38 | 3.4028e+38 | 5749.9683 | 11500.4443 | 9959.8359 | 3.4028e+38 | 3.4028e+38 | 230.0045 | 50.001 | 5749.9316 | 11500.4385 | 9959.8516 | 50 | -60.0017 | 0.016419 | 8.1752e-06 | 3.0429e-08 | 1.3591e-07 | 5.7858e-08 | 72.1045 | -65.5453 | 0.017758 | 7.991e-06 | -3.9171e-08 | 1.4229e-07 | 3.0652e-08 | 69.9499 | -142.2598 | 3.4028e+38 | 3.4028e+38 | 5749.9312 | 11500.4385 | 9959.8516 | 3.4028e+38 | 3.4028e+38 | 230.005 | 50.0009 | 5749.9897 | 11500.459 | 9959.8545 | 49.9999 | -60.0015 | 0.016208 | 8.3983e-06 | 6.0531e-08 | 1.3855e-07 | 4.0166e-08 | 71.9779 | -32.8045 | 0.017789 | 7.8217e-06 | 1.0206e-08 | 1.3922e-07 | 6.6364e-08 | 72.7794 | -78.3426 | 3.4028e+38 | 3.4028e+38 | 5749.9897 | 11500.459 | 9959.8545 | 3.4028e+38 | 3.4028e+38 | 230.0047 | 50.001 | 5749.9609 | 11500.4531 | 9959.8721 | 49.9999 | -60.0016 | 0.017682 | 8.1032e-06 | 5.7513e-08 | 1.4361e-07 | 6.8231e-08 | 73.0203 | -52.5408 | 0.019545 | 8.1501e-06 | -2.6819e-08 | 1.5997e-07 | 6.0629e-08 | 71.4465 | -112.5372 | 3.4028e+38 | 3.4028e+38 | 5749.9604 | 11500.4531 | 9959.8721 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0047 | 50.0009 | 5749.9492 | 11500.4463 | 9959.8711 | 50 | -60.0015 | 0.49998 | 3.4028e+38 | 0 | 230.0031 | 5749.52 | 11500.3506 | 9959.8066 | -0.023112 | 0.015782 | 8.343e-06 | 2.5982e-08 | 1.316e-07 | 4.9753e-08 | 72.0041 | -67.6829 | 0.15419 | 3.4028e+38 | -153.5744 | 0.019882 | 0 | 0 | 0 | 0.00075823 | 0.015566 | 7.839e-06 | -1.2326e-08 | 1.1836e-07 | 1.8119e-08 | 71.9351 | -100.3125 | -0.036731 | 3.4028e+38 | 9.1433 | 229.0753 | -6.065e-10 | 0.0017686 | -2.1904e-09 | -0.00080538 | 3.4028e+38 | 3.4028e+38 | 5749.9409 | 11500.46 | 9959.8838 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 5749.8242 | 5749.8242 | 9960.1172 | -0.005908 | 230.0048 | 50.001 | 5749.9664 | 11500.4557 | 9959.8611 | 49.9999 | -60.0017 | [17 empty] | 36.046 | 40.5679 | -37.609 | 76.1876 | 116.4704 |  | -0.0039219
113 | 230.001 | 50.0008 | 9196.7939 | 11500.2178 | -6904.7705 | 49.9999 | 36.8984 | 0.015043 | 8.0759e-06 | 3.0296e-08 | 1.2133e-07 | -6.8163e-09 | 72.6791 | 15.0917 | 0.016027 | 8.2442e-06 | 3.6771e-08 | 1.3221e-07 | 2.8416e-08 | 72.1628 | -40.8724 | 3.4028e+38 | 3.4028e+38 | 9196.7939 | 11500.2178 | -6904.771 | 3.4028e+38 | 3.4028e+38 | 230.0013 | 50.0007 | 9196.8174 | 11500.2285 | -6904.7725 | 49.9999 | 36.8986 | 0.017158 | 8.3123e-06 | 4.5268e-08 | 1.3955e-07 | -2.2355e-08 | 71.6925 | 22.4561 | 0.019105 | 8.0831e-06 | 7.5624e-08 | 1.5426e-07 | 6.2915e-08 | 71.7939 | -40.8537 | 3.4028e+38 | 3.4028e+38 | 9196.8174 | 11500.2285 | -6904.7729 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 50.0008 | 9196.8408 | 11500.2646 | -6904.7803 | 50 | 36.8983 | 0.014104 | 8.1722e-06 | 1.4128e-08 | 1.1544e-07 | -1.2884e-08 | 72.3427 | 35.3114 | 0.015439 | 8.1838e-06 | 2.9784e-08 | 1.264e-07 | 1.7156e-08 | 72.4831 | -29.0416 | 3.4028e+38 | 3.4028e+38 | 9196.8408 | 11500.2646 | -6904.7808 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 50.0008 | 9196.8262 | 11500.2422 | -6904.7974 | 49.9999 | 36.8985 | 0.015445 | 8.0639e-06 | 3.8532e-08 | 1.2505e-07 | -4.8225e-09 | 73.3858 | 6.81 | 0.016688 | 8.2667e-06 | 4.51e-08 | 1.3877e-07 | 4.9231e-08 | 72.3292 | -47.5417 | 3.4028e+38 | 3.4028e+38 | 9196.8262 | 11500.2422 | -6904.7979 | 3.4028e+38 | 3.4028e+38 | 230.001 | 50.0008 | 9196.8164 | 11500.2471 | -6904.7939 | 50 | 36.8984 | 0.014312 | 8.2135e-06 | 1.9786e-08 | 1.1745e-07 | -1.1421e-08 | 73.4763 | 36.0956 | 0.01557 | 8.2569e-06 | 3.1901e-08 | 1.2714e-07 | 2.2036e-08 | 72.3143 | -31.5445 | 3.4028e+38 | 3.4028e+38 | 9196.8164 | 11500.2471 | -6904.7944 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0016 | 50.0007 | 9196.8418 | 11500.2393 | -6904.7905 | 49.9999 | 36.8984 | 0.79971 | 3.4028e+38 | 0 | 230.0033 | 9197.0283 | 11500.3232 | -6904.5879 | 0.014398 | 0.014365 | 8.382e-06 | 7.7955e-09 | 1.1954e-07 | -2.9952e-08 | 72.6156 | 78.4241 | 0.045939 | 3.4028e+38 | -4.1403 | 0.019761 | 0 | 0 | 0 | -0.00067925 | 0.015983 | 8.0072e-06 | 4.0086e-08 | 1.2805e-07 | 4.8486e-09 | 72.0127 | -6.995 | 0.30058 | 3.4028e+38 | 18.803 | 229.3062 | 5.5894e-09 | 0.001835 | -4.0332e-10 | -0.00026258 | 3.4028e+38 | 3.4028e+38 | 9196.833 | 11500.2383 | -6904.7803 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 9197.1719 | 9197.1719 | -6904.5464 | 0.025567 | 230.0013 | 50.0008 | 9196.8189 | 11500.24 | -6904.7829 | 49.9999 | 36.8984 | [17 empty] | 19.2999 | 34.6761 | 56.2955 | 53.8109 | 70.2041 |  | 0.00035869
114 | 230.003 | 50.0009 | 11500.3545 | 11500.335 | 0.38636 | 49.9999 | -0.0019249 | 0.01767 | 8.3794e-06 | 7.2082e-08 | 1.4725e-07 | 2.8045e-09 | 72.5397 | -1.1491 | 0.019491 | 8.0979e-06 | 6.2628e-08 | 1.569e-07 | 7.8106e-08 | 71.7613 | -51.0705 | 3.4028e+38 | 3.4028e+38 | 11500.3545 | 11500.335 | 0.38636 | 3.4028e+38 | 3.4028e+38 | 230.0033 | 50.0008 | 11500.3223 | 11500.3301 | 0.39602 | 50 | -0.0018518 | 0.016989 | 8.0026e-06 | 6.6563e-08 | 1.3559e-07 | 2.4761e-08 | 72.0706 | -19.7451 | 0.018651 | 8.2441e-06 | 3.813e-08 | 1.5498e-07 | 8.1446e-08 | 72.1551 | -64.4247 | 3.4028e+38 | 3.4028e+38 | 11500.3223 | 11500.3301 | 0.39602 | 3.4028e+38 | 3.4028e+38 | 230.0024 | 50.0009 | 11500.3193 | 11500.3555 | 0.36471 | 50 | -0.001817 | 0.017876 | 8.3045e-06 | 6.4485e-08 | 1.4551e-07 | -4.4796e-09 | 72.1037 | -2.4385 | 0.019676 | 7.9297e-06 | 6.5245e-08 | 1.5737e-07 | 8.5551e-08 | 71.2521 | -50.7553 | 3.4028e+38 | 3.4028e+38 | 11500.3193 | 11500.3555 | 0.36471 | 3.4028e+38 | 3.4028e+38 | 230.0029 | 50.0008 | 11500.3125 | 11500.3184 | 0.40165 | 49.9999 | -0.002001 | 0.015574 | 8.1879e-06 | 5.4762e-08 | 1.2953e-07 | 2.8572e-08 | 71.7419 | -26.8307 | 0.01668 | 8.2645e-06 | 1.9238e-08 | 1.4173e-07 | 6.1793e-08 | 72.7242 | -68.4488 | 3.4028e+38 | 3.4028e+38 | 11500.3125 | 11500.3184 | 0.40165 | 3.4028e+38 | 3.4028e+38 | 230.0031 | 50.0008 | 11500.3369 | 11500.3438 | 0.44895 | 49.9999 | -0.0022367 | 0.015961 | 8.0968e-06 | 5.6521e-08 | 1.2941e-07 | 2.9924e-08 | 72.7656 | -26.9712 | 0.017309 | 8.1937e-06 | 2.0705e-08 | 1.4218e-07 | 6.2343e-08 | 71.5373 | -68.043 | 3.4028e+38 | 3.4028e+38 | 11500.3369 | 11500.3438 | 0.44895 | 3.4028e+38 | 3.4028e+38 | 123 | 230.003 | 50.0007 | 11500.3076 | 11500.3135 | 0.4153 | 50 | -0.001875 | 1 | 3.4028e+38 | 0 | 230.0026 | 11500.2764 | 11500.2832 | 0.84807 | -0.0052663 | 0.017524 | 8.1552e-06 | 7.0005e-08 | 1.4335e-07 | 1.1323e-08 | 72.2569 | -7.7479 | 0.49517 | 3.4028e+38 | -64.9547 | 0.019812 | 0 | 0 | 0 | -0.00020838 | 0.019483 | 8.0969e-06 | 6.4898e-08 | 1.5472e-07 | 7.4563e-08 | 72.5519 | -47.5768 | 0.4148 | 3.4028e+38 | -4.4419 | 229.4862 | 1.1199e-08 | 0.0018005 | 7.7173e-09 | -0.00075344 | 3.4028e+38 | 3.4028e+38 | 11500.3477 | 11500.3545 | 0.37601 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 11500.6504 | 11500.6504 | 0.85969 | 0.02581 | 230.003 | 50.0009 | 11500.3291 | 11500.3365 | 0.39954 | 49.9999 | -0.0019663 | [17 empty] | 26.7669 | 2.8437 | 28.1053 | 28.7565 | 2018519.39 |  | -0.0013068
115 | 230.0033 | 50.0009 | 11500.3457 | 11500.3525 | 0.43493 | 50 | -0.0021209 | 0.015556 | 8.6191e-06 | 3.5054e-08 | 1.3371e-07 | -2.0774e-08 | 71.795 | 30.0727 | 0.016689 | 7.9061e-06 | 4.9344e-08 | 1.3221e-07 | 3.8498e-08 | 72.9711 | -34.9127 | 3.4028e+38 | 3.4028e+38 | 11500.3457 | 11500.3525 | 0.43493 | 3.4028e+38 | 3.4028e+38 | 230.0029 | 50.0008 | 11500.3135 | 11500.3223 | 0.43514 | 49.9999 | -0.0020903 | 0.014345 | 8.3437e-06 | 3.0389e-08 | 1.2005e-07 | 5.6563e-09 | 71.9648 | -13.8165 | 0.015123 | 8.0115e-06 | 2.2462e-08 | 1.2727e-07 | 2.9445e-08 | 70.6402 | -51.7241 | 3.4028e+38 | 3.4028e+38 | 11500.3135 | 11500.3223 | 0.43514 | 3.4028e+38 | 3.4028e+38 | 230.0028 | 50.001 | 11500.3477 | 11500.3564 | 0.41212 | 49.9999 | -0.0020532 | 0.015839 | 8.6013e-06 | 3.9925e-08 | 1.3608e-07 | -2.0223e-08 | 72.0068 | 30.7135 | 0.017381 | 7.8718e-06 | 5.5668e-08 | 1.3376e-07 | 3.6865e-08 | 70.9368 | -35.2026 | 3.4028e+38 | 3.4028e+38 | 11500.3477 | 11500.3564 | 0.41212 | 3.4028e+38 | 3.4028e+38 | 230.0029 | 50.0008 | 11500.3232 | 11500.3311 | 0.37708 | 49.9999 | -0.0018923 | 0.017789 | 8.4076e-06 | 6.2175e-08 | 1.4709e-07 | -1.0123e-08 | 71.9214 | 13.2193 | 0.019676 | 7.9978e-06 | 7.0095e-08 | 1.5607e-07 | 7.2989e-08 | 73.2818 | -42.8691 | 3.4028e+38 | 3.4028e+38 | 11500.3232 | 11500.3311 | 0.37708 | 3.4028e+38 | 3.4028e+38 | 230.0032 | 50.0008 | 11500.3184 | 11500.3359 | 0.4013 | 50 | -0.0020363 | 0.015851 | 8.0863e-06 | 5.4652e-08 | 1.2847e-07 | 2.8968e-08 | 72.6058 | -26.6923 | 0.017224 | 7.9024e-06 | 2.3904e-08 | 1.3583e-07 | 6.4699e-08 | 70.356 | -70.0306 | 3.4028e+38 | 3.4028e+38 | 11500.3184 | 11500.3359 | 0.4013 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0031 | 50.0007 | 11500.3105 | 11500.3105 | 0.35643 | 50 | -0.0017758 | 1 | 3.4028e+38 | 0 | 230.0013 | 11500.2051 | 11500.2119 | 0.73337 | -0.01536 | 0.015877 | 8.0671e-06 | 5.5005e-08 | 1.2822e-07 | 2.7561e-08 | 71.6067 | -27.4029 | 0.41539 | 3.4028e+38 | -84.341 | 0.019987 | 0 | 0 | 0 | -0.00014476 | 0.017029 | 8.2327e-06 | 1.928e-08 | 1.3834e-07 | 5.4637e-08 | 72.6511 | -73.558 | 0.11177 | 3.4028e+38 | -1.0941 | 229.8598 | -4.4659e-10 | 0.0018875 | 3.204e-09 | -0.00078849 | 3.4028e+38 | 3.4028e+38 | 11500.3164 | 11500.3232 | 0.41114 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 11500.2852 | 11500.2852 | 0.62156 | -0.012861 | 230.003 | 50.0009 | 11500.3297 | 11500.3396 | 0.41211 | 49.9999 | -0.0020386 | [17 empty] | 27.3773 | 1.615 | 27.4041 | 28.2758 | 1720264.6328 |  | -0.0012838
116 | 230.0025 | 50.001 | 11500.3301 | 11500.3369 | 0.3776 | 50 | -0.0021016 | 0.017227 | 8.397e-06 | 6.2816e-08 | 1.4719e-07 | -1.5419e-08 | 71.5919 | 15.5428 | 0.019221 | 7.9823e-06 | 7.4379e-08 | 1.5054e-07 | 6.9209e-08 | 72.8023 | -41.8739 | 3.4028e+38 | 3.4028e+38 | 11500.3301 | 11500.3369 | 0.3776 | 3.4028e+38 | 3.4028e+38 | 230.0029 | 50.0009 | 11500.3408 | 11500.3486 | 0.40432 | 50 | -0.0020262 | 0.017752 | 8.2192e-06 | 7.2527e-08 | 1.4592e-07 | 8.7918e-09 | 71.3154 | -6.8748 | 0.019325 | 8.1085e-06 | 5.8797e-08 | 1.5742e-07 | 8.1065e-08 | 71.9952 | -54.4186 | 3.4028e+38 | 3.4028e+38 | 11500.3408 | 11500.3486 | 0.40432 | 3.4028e+38 | 3.4028e+38 | 230.003 | 50.0009 | 11500.3311 | 11500.3369 | 0.43445 | 49.9999 | -0.0021645 | 0.014949 | 8.2033e-06 | 3.9257e-08 | 1.226e-07 | 2.3814e-08 | 71.992 | -29.0506 | 0.016274 | 8.1806e-06 | 1.5551e-08 | 1.3309e-07 | 4.6558e-08 | 71.5753 | -70.9098 | 3.4028e+38 | 3.4028e+38 | 11500.3311 | 11500.3369 | 0.43445 | 3.4028e+38 | 3.4028e+38 | 230.003 | 50.0008 | 11500.3252 | 11500.3516 | 0.44025 | 49.9999 | -0.0021934 | 0.014341 | 8.3222e-06 | 3.0963e-08 | 1.1993e-07 | 1.1528e-08 | 72.2342 | -23.8722 | 0.015448 | 8.2181e-06 | 1.9076e-08 | 1.2985e-07 | 3.1628e-08 | 71.9713 | -64.268 | 3.4028e+38 | 3.4028e+38 | 11500.3252 | 11500.3516 | 0.44025 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 50.0008 | 11500.2725 | 11500.2812 | 0.40022 | 50 | -0.001994 | 0.017241 | 8.1508e-06 | 6.906e-08 | 1.4202e-07 | 2.1736e-08 | 73.2591 | -14.2984 | 0.019091 | 8.2806e-06 | 4.3878e-08 | 1.5823e-07 | 8.3772e-08 | 69.4778 | -59.7549 | 3.4028e+38 | 3.4028e+38 | 11500.2725 | 11500.2812 | 0.40022 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0026 | 50.0008 | 11500.3105 | 11500.3193 | 0.39891 | 50 | -0.00219 | 1 | 3.4028e+38 | 0 | 230.0053 | 11500.4346 | 11500.4424 | 0.94696 | 0.016146 | 0.017705 | 8.4399e-06 | 6.9273e-08 | 1.4821e-07 | -6.3202e-09 | 71.2475 | 5.6009 | 0.43484 | 3.4028e+38 | 25.6886 | 0.020219 | 0 | 0 | 0 | -0.00020988 | 0.018971 | 7.9413e-06 | 6.6191e-08 | 1.484e-07 | 5.9515e-08 | 72.388 | -41.96 | 0.46711 | 3.4028e+38 | -70.2154 | 228.9936 | 4.4995e-09 | 0.0017918 | 4.5006e-09 | -0.00075419 | 3.4028e+38 | 3.4028e+38 | 11500.3232 | 11500.3311 | 0.39747 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 11500.5889 | 11500.5889 | 0.87347 | 0.019362 | 230.0027 | 50.0009 | 11500.3199 | 11500.3311 | 0.41137 | 50 | -0.0020959 | [17 empty] | 26.1763 | -0.3787 | 24.5175 | 25.4916 | 1531105.5137 |  | -0.0012862
117 | 230.0042 | 50.0009 | 5749.9053 | 11500.4355 | 9959.8506 | 49.9999 | -60.0017 | 0.015506 | 8.455e-06 | 5.3324e-08 | 1.324e-07 | 3.1115e-08 | 72.391 | -31.4477 | 0.01682 | 7.7671e-06 | 1.3916e-08 | 1.3066e-07 | 5.2812e-08 | 71.7963 | -73.8342 | 3.4028e+38 | 3.4028e+38 | 5749.9053 | 11500.4355 | 9959.8506 | 3.4028e+38 | 3.4028e+38 | 230.0046 | 50.0011 | 5749.9644 | 11500.4629 | 9959.9014 | 49.9999 | -60.0019 | 0.014918 | 8.5266e-06 | 1.8524e-08 | 1.2726e-07 | 4.1583e-08 | 71.7399 | -64.0255 | 0.015742 | 7.8015e-06 | -2.5962e-08 | 1.2788e-07 | 1.7102e-08 | 72.2959 | -149.3708 | 3.4028e+38 | 3.4028e+38 | 5749.9639 | 11500.4629 | 9959.9014 | 3.4028e+38 | 3.4028e+38 | 230.0045 | 50.0008 | 5749.9771 | 11500.4355 | 9959.8164 | 49.9999 | -60.0016 | 0.018047 | 8.0591e-06 | 6.1868e-08 | 1.4433e-07 | 6.7589e-08 | 72.877 | -47.6014 | 0.019671 | 8.0286e-06 | -1.8462e-08 | 1.5766e-07 | 7.0207e-08 | 72.1764 | -103.7079 | 3.4028e+38 | 3.4028e+38 | 5749.9766 | 11500.4355 | 9959.8164 | 3.4028e+38 | 3.4028e+38 | 230.0048 | 50.0008 | 5749.9644 | 11500.4219 | 9959.833 | 50 | -60.0017 | 0.017661 | 8.161e-06 | 6.844e-08 | 1.4222e-07 | 5.5363e-08 | 72.5884 | -38.0937 | 0.019189 | 7.9714e-06 | -2.2101e-09 | 1.5316e-07 | 6.8251e-08 | 72.227 | -92.5751 | 3.4028e+38 | 3.4028e+38 | 5749.9639 | 11500.4219 | 9959.833 | 3.4028e+38 | 3.4028e+38 | 230.0046 | 50.001 | 5749.9165 | 11500.4102 | 9959.8477 | 49.9999 | -60.0016 | 0.016098 | 8.3654e-06 | 2.5829e-08 | 1.342e-07 | 5.8467e-08 | 72.6793 | -66.0889 | 0.017209 | 8.0787e-06 | -3.8644e-08 | 1.414e-07 | 2.6138e-08 | 71.9627 | -141.161 | 3.4028e+38 | 3.4028e+38 | 5749.916 | 11500.4102 | 9959.8477 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0049 | 50.0008 | 5749.9551 | 11500.3926 | 9959.833 | 49.9999 | -60.0017 | 0.49998 | 3.4028e+38 | 0 | 230.003 | 5749.5088 | 11500.3701 | 9959.7988 | -0.035777 | 0.017111 | 8.198e-06 | 3.4277e-08 | 1.3867e-07 | 6.497e-08 | 72.9377 | -62.1843 | 0.27068 | 3.4028e+38 | -41.0504 | 0.020202 | 0 | 0 | 0 | 0.00071433 | 0.018713 | 8.1326e-06 | -3.9288e-08 | 1.5565e-07 | 4.6614e-08 | 72.1088 | -117.342 | -0.19131 | 3.4028e+38 | -2.5491 | 229.1822 | -2.9918e-09 | 0.0018677 | 5.3621e-09 | -0.00086218 | 3.4028e+38 | 3.4028e+38 | 5749.9395 | 11500.3994 | 9959.7861 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 5749.4961 | 5749.4961 | 9959.9258 | -0.037819 | 230.0046 | 50.0009 | 5749.9455 | 11500.4332 | 9959.8498 | 49.9999 | -60.0017 | [17 empty] | 33.5925 | 32.4442 | -38.5431 | 65.415 | 102.6722 |  | -0.0035607
118 | 230.0046 | 50.0008 | 5749.9146 | 11500.4189 | 9959.8145 | 50 | -60.0017 | 0.015688 | 8.4309e-06 | 2.1991e-08 | 1.3187e-07 | 4.9838e-08 | 72.4709 | -65.6711 | 0.017391 | 7.9965e-06 | -3.8161e-08 | 1.3498e-07 | 1.7625e-08 | 71.7993 | -148.9009 | 3.4028e+38 | 3.4028e+38 | 5749.9141 | 11500.4189 | 9959.8145 | 3.4028e+38 | 3.4028e+38 | 230.0045 | 50.001 | 5749.959 | 11500.4424 | 9959.8711 | 49.9999 | -60.0018 | 0.014431 | 8.6643e-06 | 3.317e-08 | 1.2441e-07 | 2.6288e-08 | 71.7981 | -36.4781 | 0.015215 | 7.6753e-06 | 2.8672e-09 | 1.1605e-07 | 2.7982e-08 | 72.1488 | -83.91 | 3.4028e+38 | 3.4028e+38 | 5749.959 | 11500.4424 | 9959.8711 | 3.4028e+38 | 3.4028e+38 | 230.0047 | 50.0008 | 5749.9429 | 11500.3955 | 9959.7998 | 49.9999 | -60.0016 | 0.017635 | 8.033e-06 | 4.7969e-08 | 1.4213e-07 | 6.6992e-08 | 72.2673 | -60.3023 | 0.019434 | 8.1785e-06 | -3.7439e-08 | 1.52e-07 | 4.7237e-08 | 70.9014 | -137.6516 | 3.4028e+38 | 3.4028e+38 | 5749.9424 | 11500.3955 | 9959.7998 | 3.4028e+38 | 3.4028e+38 | 230.0046 | 50.001 | 5749.9565 | 11500.4688 | 9959.8535 | 50 | -60.0016 | 0.015086 | 8.4799e-06 | 1.9486e-08 | 1.2591e-07 | 3.5219e-08 | 71.823 | -63.9856 | 0.016061 | 7.9339e-06 | -2.6952e-08 | 1.2375e-07 | 1.7873e-08 | 71.9413 | -139.0619 | 3.4028e+38 | 3.4028e+38 | 5749.9561 | 11500.4688 | 9959.8535 | 3.4028e+38 | 3.4028e+38 | 230.0042 | 50.0008 | 5749.9106 | 11500.3896 | 9959.8154 | 49.9999 | -60.0015 | 0.01734 | 8.0621e-06 | 3.5515e-08 | 1.3883e-07 | 6.683e-08 | 72.2069 | -64.1199 | 0.018946 | 8.2118e-06 | -3.6887e-08 | 1.5503e-07 | 3.518e-08 | 72.6346 | -138.7353 | 3.4028e+38 | 3.4028e+38 | 5749.9102 | 11500.3896 | 9959.8154 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0044 | 50.0009 | 5749.9497 | 11500.4053 | 9959.8086 | 49.9999 | -60.0016 | 0.49998 | 3.4028e+38 | 0 | 230.0037 | 5749.6621 | 11500.3916 | 9959.8467 | -0.012039 | 0.014586 | 8.5135e-06 | 2.0719e-08 | 1.2418e-07 | 3.2008e-08 | 72.393 | -53.9412 | 0.18568 | 3.4028e+38 | -63.5807 | 0.020382 | 0 | 0 | 0 | 0.00075376 | 0.015375 | 7.7356e-06 | 2.2524e-09 | 1.1861e-07 | 2.7957e-08 | 72.1824 | -77.7366 | 0.059516 | 3.4028e+38 | -79.4129 | 228.6723 | -3.269e-10 | 0.0017999 | 1.3188e-09 | -0.00081744 | 3.4028e+38 | 3.4028e+38 | 5749.9604 | 11500.4512 | 9959.8486 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 5749.8066 | 5749.7832 | 9960.1406 | -0.012439 | 230.0045 | 50.0009 | 5749.9367 | 11500.423 | 9959.8309 | 50 | -60.0016 | [17 empty] | 33.6857 | 33.167 | -37.6458 | 66.6535 | 102.79 |  | -0.0035003
119 | 230.0045 | 50.0009 | 5749.9341 | 11500.4287 | 9959.8135 | 50 | -60.0016 | 0.017537 | 8.1928e-06 | 6.7561e-08 | 1.4272e-07 | 5.5633e-08 | 72.3895 | -38.0152 | 0.018891 | 7.9112e-06 | 1.3259e-09 | 1.5207e-07 | 6.8963e-08 | 71.7168 | -90.1465 | 3.4028e+38 | 3.4028e+38 | 5749.9341 | 11500.4287 | 9959.8135 | 3.4028e+38 | 3.4028e+38 | 230.0046 | 50.0009 | 5749.9268 | 11500.4307 | 9959.8662 | 50 | -60.0016 | 0.01603 | 8.3206e-06 | 2.5941e-08 | 1.3594e-07 | 5.9862e-08 | 72.8879 | -64.2624 | 0.01769 | 7.9077e-06 | -3.9553e-08 | 1.4009e-07 | 3.2143e-08 | 72.5581 | -140.3088 | 3.4028e+38 | 3.4028e+38 | 5749.9263 | 11500.4307 | 9959.8662 | 3.4028e+38 | 3.4028e+38 | 230.0048 | 50.001 | 5749.9141 | 11500.416 | 9959.8555 | 49.9999 | -60.0016 | 0.01486 | 8.5171e-06 | 2.0737e-08 | 1.2473e-07 | 4.27e-08 | 72.2773 | -66.7263 | 0.016067 | 7.8685e-06 | -2.5869e-08 | 1.2651e-07 | 1.7095e-08 | 73.3095 | -155.1883 | 3.4028e+38 | 3.4028e+38 | 5749.9136 | 11500.416 | 9959.8555 | 3.4028e+38 | 3.4028e+38 | 230.0041 | 50.0008 | 5749.9355 | 11500.3789 | 9959.8252 | 49.9999 | -60.0017 | 0.017617 | 8.0315e-06 | 5.3245e-08 | 1.4223e-07 | 6.6112e-08 | 72.1432 | -51.8296 | 0.019356 | 8.051e-06 | -2.8669e-08 | 1.5715e-07 | 5.9199e-08 | 72.6698 | -111.2492 | 3.4028e+38 | 3.4028e+38 | 5749.9351 | 11500.3789 | 9959.8252 | 3.4028e+38 | 3.4028e+38 | 230.0049 | 50.0011 | 5749.9277 | 11500.4424 | 9959.8799 | 49.9999 | -60.0015 | 0.014623 | 8.6938e-06 | 3.1941e-08 | 1.2713e-07 | 2.7442e-08 | 73.3075 | -44.6182 | 0.015735 | 7.7429e-06 | -4.5566e-10 | 1.203e-07 | 2.4111e-08 | 72.0921 | -97.6499 | 3.4028e+38 | 3.4028e+38 | 5749.9272 | 11500.4424 | 9959.8799 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0048 | 50.001 | 5749.9727 | 11500.4658 | 9959.8809 | 49.9999 | -60.0017 | 0.49998 | 3.4028e+38 | 0 | 230.0055 | 5749.7832 | 11500.4873 | 9959.9404 | 0.00012085 | 0.014359 | 8.562e-06 | 2.6461e-08 | 1.23e-07 | 2.9156e-08 | 72.924 | -49.6162 | 0.20024 | 3.4028e+38 | -84.7816 | 0.020034 | 0 | 0 | 0 | 0.00075723 | 0.015532 | 7.8081e-06 | -1.2273e-08 | 1.1995e-07 | 1.7994e-08 | 71.3395 | -123.3724 | -0.10323 | 3.4028e+38 | -73.4198 | 229.2535 | -2.8788e-09 | 0.0018033 | -3.1691e-09 | -0.00082855 | 3.4028e+38 | 3.4028e+38 | 5749.9199 | 11500.4736 | 9959.8662 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 5749.7217 | 5749.7217 | 9959.8174 | -0.0046764 | 230.0046 | 50.0009 | 5749.9276 | 11500.4193 | 9959.848 | 49.9999 | -60.0016 | [17 empty] | 34.3213 | 28.1621 | -43.9437 | 60.405 | 98.1871 |  | -0.0034367
120 | 230.0013 | 50.0008 | 9196.8203 | 11500.209 | -6904.7358 | 49.9999 | 36.8983 | 0.015153 | 8.0214e-06 | 3.5877e-08 | 1.2333e-07 | -5.3123e-09 | 72.3019 | 8.4699 | 0.016363 | 8.2894e-06 | 3.9762e-08 | 1.3632e-07 | 4.3549e-08 | 71.6853 | -45.8037 | 3.4028e+38 | 3.4028e+38 | 9196.8203 | 11500.209 | -6904.7363 | 3.4028e+38 | 3.4028e+38 | 230.001 | 50.001 | 9196.8486 | 11500.2627 | -6904.7588 | 49.9999 | 36.8983 | 0.017267 | 8.6275e-06 | 2.2947e-08 | 1.4717e-07 | -5.2426e-08 | 71.9326 | 68.323 | 0.01911 | 7.8896e-06 | 8.3023e-08 | 1.5146e-07 | 2.4777e-08 | 71.4914 | -17.2631 | 3.4028e+38 | 3.4028e+38 | 9196.8486 | 11500.2627 | -6904.7593 | 3.4028e+38 | 3.4028e+38 | 230.0011 | 50.0008 | 9196.8242 | 11500.2402 | -6904.7861 | 50 | 36.8986 | 0.017436 | 8.3398e-06 | 4.363e-08 | 1.446e-07 | -2.7108e-08 | 72.9304 | 29.2835 | 0.019451 | 8.0831e-06 | 8.5116e-08 | 1.5854e-07 | 6.3335e-08 | 72.0566 | -36.9095 | 3.4028e+38 | 3.4028e+38 | 9196.8242 | 11500.2402 | -6904.7866 | 3.4028e+38 | 3.4028e+38 | 230.0007 | 50.0009 | 9196.8555 | 11500.2461 | -6904.7598 | 50 | 36.8982 | 0.014901 | 8.5135e-06 | 6.4302e-09 | 1.2931e-07 | -3.7522e-08 | 71.8641 | 84.0353 | 0.016545 | 8.0751e-06 | 4.6475e-08 | 1.3381e-07 | 5.0833e-09 | 71.2931 | -3.6481 | 3.4028e+38 | 3.4028e+38 | 9196.8555 | 11500.2461 | -6904.7603 | 3.4028e+38 | 3.4028e+38 | 230.0013 | 50.0009 | 9196.8525 | 11500.2549 | -6904.7769 | 49.9999 | 36.8983 | 0.015716 | 8.5373e-06 | 4.4474e-09 | 1.3714e-07 | -5.1922e-08 | 72.7094 | 81.4129 | 0.017522 | 7.9352e-06 | 6.2687e-08 | 1.4059e-07 | 9.884e-09 | 71.6691 | -8.9031 | 3.4028e+38 | 3.4028e+38 | 9196.8525 | 11500.2549 | -6904.7773 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0012 | 50.0009 | 9196.8418 | 11500.2568 | -6904.7822 | 50 | 36.8984 | 0.79971 | 3.4028e+38 | 0 | 230.0059 | 9197.2686 | 11500.4844 | -6904.4395 | 0.040957 | 0.016873 | 8.6382e-06 | 1.6437e-08 | 1.4485e-07 | -5.1952e-08 | 72.1789 | 72.4432 | 0.084198 | 3.4028e+38 | -69.9508 | 0.020849 | 0 | 0 | 0 | -0.00072607 | 0.018379 | 7.9213e-06 | 7.3009e-08 | 1.4866e-07 | 2.0173e-08 | 73.0417 | -13.7963 | 0.49594 | 3.4028e+38 | -80.6531 | 229.1061 | 1.5849e-08 | 0.0018186 | 7.3231e-09 | -0.00029138 | 3.4028e+38 | 3.4028e+38 | 9196.8818 | 11500.2861 | -6904.7896 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 9197.2568 | 9197.2568 | -6904.478 | 0.029745 | 230.0011 | 50.0009 | 9196.8402 | 11500.2426 | -6904.7635 | 49.9999 | 36.8983 | [17 empty] | 19.3692 | 34.2264 | 60.48 | 52.6329 | 60.1913 |  | 1.008e-05
121 | 230.0009 | 50.0009 | 9196.8311 | 11500.2393 | -6904.7886 | 50 | 36.8984 | 0.01744 | 8.5055e-06 | 3.3798e-08 | 1.4657e-07 | -4.2965e-08 | 71.156 | 47.033 | 0.019555 | 7.9869e-06 | 8.908e-08 | 1.5649e-07 | 5.323e-08 | 72.0261 | -30.2249 | 3.4028e+38 | 3.4028e+38 | 9196.8311 | 11500.2393 | -6904.7891 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 50.0009 | 9196.8477 | 11500.2412 | -6904.7412 | 49.9999 | 36.8983 | 0.016061 | 8.5619e-06 | 5.166e-09 | 1.3862e-07 | -5.1899e-08 | 72.2205 | 81.5605 | 0.017572 | 7.9742e-06 | 5.9674e-08 | 1.3892e-07 | 6.9026e-09 | 71.9478 | -6.5771 | 3.4028e+38 | 3.4028e+38 | 9196.8477 | 11500.2412 | -6904.7417 | 3.4028e+38 | 3.4028e+38 | 230.0011 | 50.0009 | 9196.8809 | 11500.291 | -6904.8003 | 50 | 36.8983 | 0.017555 | 8.5757e-06 | 3.0611e-08 | 1.513e-07 | -4.6164e-08 | 71.9129 | 53.0812 | 0.019664 | 7.8783e-06 | 8.3046e-08 | 1.5562e-07 | 4.0688e-08 | 72.6666 | -27.71 | 3.4028e+38 | 3.4028e+38 | 9196.8809 | 11500.291 | -6904.8008 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 50.0008 | 9196.834 | 11500.2314 | -6904.7515 | 49.9999 | 36.8983 | 0.014209 | 8.1604e-06 | 1.5201e-08 | 1.1793e-07 | -1.2998e-08 | 72.686 | 36.292 | 0.015367 | 8.2457e-06 | 3.1294e-08 | 1.276e-07 | 1.8393e-08 | 71.0466 | -30.7447 | 3.4028e+38 | 3.4028e+38 | 9196.834 | 11500.2314 | -6904.752 | 3.4028e+38 | 3.4028e+38 | 230.0013 | 50.0009 | 9196.876 | 11500.2842 | -6904.8042 | 49.9999 | 36.8983 | 0.016229 | 8.6664e-06 | 9.3156e-09 | 1.3683e-07 | -5.0146e-08 | 71.634 | 85.3739 | 0.017919 | 8.0117e-06 | 6.5807e-08 | 1.4085e-07 | 8.0255e-09 | 72.2624 | -7.9404 | 3.4028e+38 | 3.4028e+38 | 9196.876 | 11500.2842 | -6904.8047 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0012 | 50.0009 | 9196.8613 | 11500.2607 | -6904.7612 | 49.9999 | 36.8983 | 0.79971 | 3.4028e+38 | 0 | 230.0037 | 9197.0801 | 11500.3662 | -6904.6025 | 0.0091074 | 0.014487 | 8.4283e-06 | 1.1948e-08 | 1.1957e-07 | -1.7276e-08 | 73.6361 | 55.333 | 0.12545 | 3.4028e+38 | 9.7045 | 0.020048 | 0 | 0 | 0 | -0.00065666 | 0.015801 | 8.2852e-06 | 3.2584e-08 | 1.3314e-07 | 3.4819e-08 | 71.8835 | -44.9341 | 0.26212 | 3.4028e+38 | -56.2523 | 229.9657 | 8.1247e-10 | 0.0018739 | 5.6042e-09 | -0.00031215 | 3.4028e+38 | 3.4028e+38 | 9196.79 | 11500.207 | -6904.7817 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 9197.0156 | 9197.0156 | -6904.376 | 0.01163 | 230.0012 | 50.0009 | 9196.8539 | 11500.2574 | -6904.7771 | 50 | 36.8984 | [17 empty] | 18.8392 | 31.5395 | 59.7424 | 50.0786 | 55.4513 |  | -0.00010902
122 | 230.0016 | 50.0009 | 9196.8672 | 11500.2715 | -6904.7725 | 49.9999 | 36.8985 | 0.016533 | 8.6581e-06 | 9.6475e-09 | 1.4208e-07 | -5.2785e-08 | 71.9624 | 79.5959 | 0.017997 | 7.9901e-06 | 6.4631e-08 | 1.4371e-07 | 1.2786e-08 | 72.1985 | -9.2954 | 3.4028e+38 | 3.4028e+38 | 9196.8672 | 11500.2715 | -6904.7729 | 3.4028e+38 | 3.4028e+38 | 230.0009 | 50.0009 | 9196.834 | 11500.2559 | -6904.811 | 49.9999 | 36.8986 | 0.017653 | 8.3656e-06 | 3.79e-08 | 1.4618e-07 | -4.4178e-08 | 72.0915 | 49.7638 | 0.019765 | 8.0294e-06 | 8.1968e-08 | 1.5729e-07 | 4.8593e-08 | 72.2691 | -28.2293 | 3.4028e+38 | 3.4028e+38 | 9196.834 | 11500.2559 | -6904.8115 | 3.4028e+38 | 3.4028e+38 | 230.0009 | 50.001 | 9196.8477 | 11500.2363 | -6904.7461 | 49.9999 | 36.8983 | 0.016319 | 8.594e-06 | 1.0844e-08 | 1.4017e-07 | -4.9741e-08 | 72.1541 | 77.4495 | 0.01804 | 7.9372e-06 | 6.1196e-08 | 1.4052e-07 | 6.2985e-09 | 72.052 | -9.5406 | 3.4028e+38 | 3.4028e+38 | 9196.8477 | 11500.2363 | -6904.7466 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 50.0007 | 9196.8271 | 11500.2061 | -6904.7583 | 49.9999 | 36.8984 | 0.015805 | 8.0279e-06 | 4.2055e-08 | 1.2559e-07 | -4.0785e-09 | 72.4762 | 5.7282 | 0.017035 | 8.3651e-06 | 4.6808e-08 | 1.3952e-07 | 5.0299e-08 | 72.8986 | -47.572 | 3.4028e+38 | 3.4028e+38 | 9196.8271 | 11500.2061 | -6904.7588 | 3.4028e+38 | 3.4028e+38 | 230.0008 | 50.0008 | 9196.8047 | 11500.2207 | -6904.8086 | 50 | 36.8986 | 0.017637 | 8.4385e-06 | 3.8825e-08 | 1.4705e-07 | -4.5126e-08 | 72.2028 | 52.2961 | 0.019739 | 7.8556e-06 | 8.5001e-08 | 1.5962e-07 | 4.603e-08 | 72.8538 | -28.4874 | 3.4028e+38 | 3.4028e+38 | 9196.8047 | 11500.2207 | -6904.8091 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0013 | 50.001 | 9196.8564 | 11500.2861 | -6904.8242 | 49.9999 | 36.8987 | 0.79971 | 3.4028e+38 | 0 | 230.0031 | 9197.2207 | 11500.334 | -6904.2856 | 0.033798 | 0.017667 | 8.3619e-06 | 3.981e-08 | 1.4913e-07 | -4.12e-08 | 72.8373 | 38.8609 | 0.29655 | 3.4028e+38 | -38.7465 | 0.021111 | 0 | 0 | 0 | -0.00072654 | 0.019324 | 8.1645e-06 | 7.777e-08 | 1.5762e-07 | 5.4048e-08 | 72.1891 | -36.3491 | 0.54184 | 3.4028e+38 | -36.9874 | 229.4203 | 9.5013e-09 | 0.0018503 | 5.9569e-09 | -0.00033138 | 3.4028e+38 | 3.4028e+38 | 9196.8164 | 11500.25 | -6904.8037 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 9197.1084 | 9197.1084 | -6904.2729 | 0.020612 | 230.0012 | 50.0009 | 9196.8361 | 11500.2381 | -6904.7793 | 49.9999 | 36.8985 | [17 empty] | 18.8818 | 32.934 | 58.0865 | 50.0816 | 59.9451 |  | 0.00012202
123 | 230.0043 | 100.0011 | 23000.6816 | 23000.6973 | 1.0992 | 49.9999 | -0.0029707 | 0.01795 | 1.4787e-05 | 1.2923e-07 | 2.6543e-07 | -9.3732e-08 | 49.9532 | 36.5626 | 0.019792 | 1.3331e-05 | 1.3177e-07 | 2.6383e-07 | 9.1273e-08 | 50.0526 | -35.1634 | 3.4028e+38 | 3.4028e+38 | 23000.6816 | 23000.6973 | 1.0992 | 3.4028e+38 | 3.4028e+38 | 230.0041 | 100.0011 | 23000.6348 | 23000.6504 | 1.1222 | 49.9999 | -0.0027955 | 0.01772 | 1.4644e-05 | 1.3651e-07 | 2.5807e-07 | -6.6103e-08 | 49.9403 | 25.8371 | 0.019202 | 1.3408e-05 | 1.1222e-07 | 2.6426e-07 | 1.1018e-07 | 50.0598 | -44.5675 | 3.4028e+38 | 3.4028e+38 | 23000.6348 | 23000.6504 | 1.1222 | 3.4028e+38 | 3.4028e+38 | 230.0041 | 100.0013 | 23000.6699 | 23000.6855 | 1.2114 | 49.9999 | -0.0027609 | 0.017611 | 1.501e-05 | 1.1229e-07 | 2.6435e-07 | -1.0706e-07 | 50.1092 | 42.118 | 0.019639 | 1.3168e-05 | 1.4117e-07 | 2.5797e-07 | 7.1858e-08 | 49.996 | -27.5811 | 3.4028e+38 | 3.4028e+38 | 23000.6699 | 23000.6855 | 1.2114 | 3.4028e+38 | 3.4028e+38 | 230.0045 | 100.0012 | 23000.7031 | 23000.6309 | 1.1678 | 49.9999 | -0.002909 | 0.017427 | 1.455e-05 | 1.3038e-07 | 2.5217e-07 | -5.1455e-08 | 49.913 | 17.3105 | 0.018915 | 1.3593e-05 | 9.4687e-08 | 2.5712e-07 | 1.1478e-07 | 49.9928 | -52.9968 | 3.4028e+38 | 3.4028e+38 | 23000.7031 | 23000.6309 | 1.1678 | 3.4028e+38 | 3.4028e+38 | 230.0041 | 100.0012 | 23000.6484 | 23000.6602 | 1.111 | 50 | -0.0029072 | 0.015893 | 1.4524e-05 | 1.0884e-07 | 2.3082e-07 | -1.5712e-08 | 49.9904 | 8.302 | 0.016923 | 1.3745e-05 | 5.1806e-08 | 2.2906e-07 | 9.2194e-08 | 50.0287 | -61.4505 | 3.4028e+38 | 3.4028e+38 | 23000.6484 | 23000.6602 | 1.111 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0042 | 100.0011 | 23000.6562 | 23000.6719 | 1.1723 | 49.9999 | -0.0028258 | 1 | 3.4028e+38 | 0 | 230.0034 | 23000.5527 | 23000.5645 | 1.7012 | 0.0067176 | 0.014924 | 1.4742e-05 | 6.9824e-08 | 2.1551e-07 | -2.0203e-08 | 50.0413 | 17.2881 | 0.28555 | 3.4028e+38 | 157.8422 | 0.02027 | 0 | 0 | 0 | -0.00072639 | 0.015374 | 1.3476e-05 | 4.5634e-08 | 2.072e-07 | 2.237e-08 | 49.9744 | -22.96 | 0.25437 | 3.4028e+38 | 144.3038 | 229.9647 | 2.1685e-09 | 0.0030767 | -2.053e-09 | -0.0012016 | 3.4028e+38 | 3.4028e+38 | 23000.7168 | 23000.7285 | 1.1277 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 23001.1113 | 23001.1113 | 1.6603 | 0.028143 | 230.0042 | 100.0012 | 23000.6676 | 23000.6648 | 1.1423 | 49.9999 | -0.0028686 | [17 empty] | 33.773 | -4.9689 | 27.7999 | 27.6866 | 5517812.2544 |  | -0.002432
124 | 230.0074 | 100.0013 | 11499.7783 | 23001.0703 | 19919.9453 | 49.9999 | -60.002 | 0.014551 | 1.5222e-05 | 5.9193e-08 | 2.2112e-07 | 4.9589e-08 | 49.9955 | -38.3849 | 0.015613 | 1.3109e-05 | -2.4794e-08 | 2.0689e-07 | 5.7207e-08 | 50.0377 | -111.5572 | 3.4028e+38 | 3.4028e+38 | 11499.7773 | 23001.0703 | 19919.9453 | 3.4028e+38 | 3.4028e+38 | 230.0074 | 100.0013 | 11499.8301 | 23001.0684 | 19919.9707 | 50 | -60.0021 | 0.017971 | 1.4486e-05 | 1.4131e-07 | 2.6039e-07 | 6.3043e-08 | 50 | -23.8678 | 0.019309 | 1.359e-05 | -1.5675e-08 | 2.6451e-07 | 1.6092e-07 | 49.9841 | -95.6725 | 3.4028e+38 | 3.4028e+38 | 11499.8291 | 23001.0684 | 19919.9707 | 3.4028e+38 | 3.4028e+38 | 230.0079 | 100.0014 | 11499.8799 | 23000.9902 | 19919.9043 | 49.9999 | -60.0021 | 0.01439 | 1.5341e-05 | 6.3598e-08 | 2.2183e-07 | 2.3097e-08 | 49.8976 | -17.5639 | 0.015415 | 1.2974e-05 | -1.0729e-09 | 1.9829e-07 | 5.1836e-08 | 49.9916 | -86.1616 | 3.4028e+38 | 3.4028e+38 | 11499.8789 | 23000.9902 | 19919.9043 | 3.4028e+38 | 3.4028e+38 | 230.0074 | 100.0012 | 11499.8506 | 23000.9844 | 19919.8926 | 49.9999 | -60.002 | 0.01789 | 1.4449e-05 | 1.3375e-07 | 2.5696e-07 | 7.7198e-08 | 49.9811 | -30.9289 | 0.019377 | 1.3616e-05 | -2.9929e-08 | 2.6744e-07 | 1.5674e-07 | 50.028 | -101.4874 | 3.4028e+38 | 3.4028e+38 | 11499.8496 | 23000.9844 | 19919.8926 | 3.4028e+38 | 3.4028e+38 | 230.0077 | 100.0013 | 11499.8369 | 23001.0781 | 19919.9473 | 50 | -60.0019 | 0.014707 | 1.5214e-05 | 5.7089e-08 | 2.2624e-07 | 5.0494e-08 | 50.0128 | -41.4736 | 0.015558 | 1.3162e-05 | -2.7074e-08 | 2.0693e-07 | 5.7542e-08 | 49.9592 | -115.6252 | 3.4028e+38 | 3.4028e+38 | 11499.8359 | 23001.0781 | 19919.9473 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0075 | 100.0014 | 11499.8662 | 23001.0469 | 19919.8848 | 50 | -60.0019 | 0.49997 | 3.4028e+38 | 0 | 230.0072 | 11499.3994 | 23000.9941 | 19920.1191 | -0.0021667 | 0.014718 | 1.5176e-05 | 5.6513e-08 | 2.2345e-07 | 5.0097e-08 | 49.9685 | -42.5366 | 0.25263 | 3.4028e+38 | -154.1954 | 0.020208 | 0 | 0 | 0 | 0.0011977 | 0.015623 | 1.318e-05 | -2.6814e-08 | 2.0737e-07 | 5.8491e-08 | 49.9836 | -113.8137 | -0.12207 | 3.4028e+38 | 121.6155 | 230.0033 | -5.655e-09 | 0.0030205 | -5.3193e-09 | -0.0016466 | 3.4028e+38 | 3.4028e+38 | 11499.8057 | 23001.0332 | 19919.9297 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 11499.4189 | 11499.4189 | 19920.1133 | 0.0020566 | 230.0075 | 100.0013 | 11499.8352 | 23001.0383 | 19919.932 | 49.9999 | -60.002 | [17 empty] | 47.5696 | 34.9968 | -44.4512 | 81.9976 | 126.8389 |  | -0.0042355
125 | 230.0004 | 100.0013 | 18393.7344 | 23000.2969 | -13809.1602 | 50 | 36.8975 | 0.017062 | 1.5338e-05 | -6.5106e-09 | 2.5657e-07 | -1.325e-07 | 49.9685 | 93.6823 | 0.01851 | 1.3116e-05 | 1.2431e-07 | 2.4036e-07 | -5.9907e-08 | 50.0088 | 27.0373 | 3.4028e+38 | 3.4028e+38 | 18393.7324 | 23000.2969 | -13809.1611 | 3.4028e+38 | 3.4028e+38 | 230.0003 | 100.0012 | 18393.6699 | 23000.2832 | -13809.1846 | 50 | 36.8976 | 0.017635 | 1.4896e-05 | 6.7263e-08 | 2.627e-07 | -1.4166e-07 | 50.024 | 67.2048 | 0.019306 | 1.3383e-05 | 1.4948e-07 | 2.5921e-07 | 1.4907e-08 | 50.0434 | -1.5193 | 3.4028e+38 | 3.4028e+38 | 18393.6699 | 23000.2832 | -13809.1855 | 3.4028e+38 | 3.4028e+38 | 230.0005 | 100.0013 | 18393.7285 | 23000.2656 | -13809.1211 | 49.9999 | 36.8975 | 0.014831 | 1.5071e-05 | -6.4664e-09 | 2.2645e-07 | -7.8378e-08 | 50.0232 | 94.6517 | 0.016263 | 1.3512e-05 | 6.8997e-08 | 2.1973e-07 | -4.4646e-08 | 49.9974 | 28.2946 | 3.4028e+38 | 3.4028e+38 | 18393.7266 | 23000.2656 | -13809.1221 | 3.4028e+38 | 3.4028e+38 | 230.0003 | 100.0013 | 18393.7695 | 23000.3164 | -13809.127 | 49.9999 | 36.8975 | 0.017624 | 1.5315e-05 | 1.7667e-08 | 2.6925e-07 | -1.5263e-07 | 50.0048 | 85.5103 | 0.019579 | 1.3069e-05 | 1.5006e-07 | 2.5588e-07 | -4.2129e-08 | 50.0293 | 16.1095 | 3.4028e+38 | 3.4028e+38 | 18393.7695 | 23000.3164 | -13809.1279 | 3.4028e+38 | 3.4028e+38 | 230.0002 | 100.0011 | 18393.7109 | 23000.2832 | -13809.1572 | 50 | 36.8977 | 0.017753 | 1.5318e-05 | 2.9704e-08 | 2.7047e-07 | -1.5546e-07 | 49.9839 | 80.1385 | 0.019546 | 1.31e-05 | 1.5366e-07 | 2.5827e-07 | -2.6787e-08 | 49.9436 | 12.5236 | 3.4028e+38 | 3.4028e+38 | 18393.7109 | 23000.2832 | -13809.1582 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0003 | 100.0011 | 18393.6895 | 23000.3281 | -13809.1641 | 49.9999 | 36.8975 | 0.79972 | 3.4028e+38 | 0 | 230.0043 | 18394.5586 | 23000.6738 | -13808.6895 | 0.10626 | 0.017261 | 1.5412e-05 | 8.9557e-09 | 2.6234e-07 | -1.4673e-07 | 49.9484 | 89.2668 | 0.0097497 | 3.4028e+38 | -156.6121 | 0.020123 | 0 | 0 | 0 | -0.0017344 | 0.018823 | 1.3119e-05 | 1.2963e-07 | 2.4578e-07 | -5.5897e-08 | 50.0328 | 23.5754 | 0.50371 | 3.4028e+38 | 92.2277 | 230.1406 | 1.4093e-08 | 0.0030224 | -1.2305e-08 | -0.00024441 | 3.4028e+38 | 3.4028e+38 | 18393.7617 | 23000.2617 | -13809.083 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 18394.5234 | 18394.5469 | -13808.8896 | 0.08179 | 230.0003 | 100.0012 | 18393.7227 | 23000.2891 | -13809.15 | 49.9999 | 36.8975 | [17 empty] | 16.1739 | 27.5395 | 55.3032 | 42.6132 | 42.0029 |  | -0.00033338
126 | 230.004 | 100.0012 | 23000.6328 | 23000.6816 | 1.1977 | 49.9999 | -0.0029656 | 0.014384 | 1.5002e-05 | 5.3962e-08 | 2.1579e-07 | -3.0361e-08 | 50.0083 | 29.4877 | 0.015742 | 1.3616e-05 | 4.7302e-08 | 2.1435e-07 | 3.0509e-08 | 50.0207 | -34.3236 | 3.4028e+38 | 3.4028e+38 | 23000.6328 | 23000.6816 | 1.1977 | 3.4028e+38 | 3.4028e+38 | 230.0037 | 100.0011 | 23000.6094 | 23000.5957 | 1.1889 | 49.9999 | -0.0029617 | 0.014881 | 1.4697e-05 | 7.9734e-08 | 2.1856e-07 | -1.4086e-08 | 49.9762 | 7.9413 | 0.016064 | 1.3726e-05 | 3.9621e-08 | 2.2053e-07 | 6.7824e-08 | 50.0267 | -59.9818 | 3.4028e+38 | 3.4028e+38 | 23000.6094 | 23000.5957 | 1.1889 | 3.4028e+38 | 3.4028e+38 | 230.0036 | 100.0012 | 23000.6016 | 23000.6133 | 1.1851 | 49.9999 | -0.0029521 | 0.017925 | 1.4743e-05 | 1.3473e-07 | 2.613e-07 | -7.4875e-08 | 51.0011 | 30.8597 | 0.019431 | 1.3439e-05 | 1.1972e-07 | 2.6543e-07 | 1.0475e-07 | 49.9949 | -40.1778 | 3.4028e+38 | 3.4028e+38 | 23000.6016 | 23000.6133 | 1.1851 | 3.4028e+38 | 3.4028e+38 | 230.0039 | 100.0009 | 23000.5293 | 23000.5469 | 1.1675 | 49.9999 | -0.0026521 | 0.014473 | 1.4924e-05 | 6.1106e-08 | 2.1632e-07 | -2.4804e-08 | 49.9619 | 21.5965 | 0.015566 | 1.3629e-05 | 4.2512e-08 | 2.1213e-07 | 3.9768e-08 | 50.0289 | -44.1912 | 3.4028e+38 | 3.4028e+38 | 23000.5293 | 23000.5469 | 1.1675 | 3.4028e+38 | 3.4028e+38 | 230.0042 | 100.0009 | 23000.6074 | 23000.5547 | 1.1567 | 49.9999 | -0.0028814 | 0.015661 | 1.4588e-05 | 9.8699e-08 | 2.2843e-07 | -1.5173e-08 | 50.0141 | 8.4284 | 0.016529 | 1.3768e-05 | 4.8674e-08 | 2.2882e-07 | 8.9742e-08 | 50.0238 | -61.4411 | 3.4028e+38 | 3.4028e+38 | 23000.6074 | 23000.5547 | 1.1567 | 3.4028e+38 | 3.4028e+38 | 123 | 230.004 | 100.0008 | 23000.5508 | 23000.5645 | 1.212 | 50 | -0.0026738 | 1 | 3.4028e+38 | 0 | 230.0019 | 23000.4121 | 23000.4258 | 2.0756 | -0.011873 | 0.016715 | 1.4625e-05 | 1.2788e-07 | 2.4699e-07 | -3.1563e-08 | 49.9359 | 13.8444 | 0.52778 | 3.4028e+38 | 129.6556 | 0.020016 | 0 | 0 | 0 | -0.00078372 | 0.018969 | 1.3691e-05 | 9.7871e-08 | 2.5969e-07 | 1.1521e-07 | 49.9421 | -49.0804 | 0.38517 | 3.4028e+38 | 82.7402 | 229.9863 | 1.2151e-08 | 0.0031189 | 1.0294e-08 | -0.0012887 | 3.4028e+38 | 3.4028e+38 | 23000.5742 | 23000.5879 | 1.1015 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 23000.7012 | 23000.7012 | 2.184 | 0.013379 | 230.0039 | 100.0011 | 23000.5961 | 23000.5984 | 1.1792 | 49.9999 | -0.0028826 | [17 empty] | 31.6836 | -5.6443 | 24.5632 | 24.6704 | 6153862.053 |  | -0.002472
127 | 230.004 | 100.0008 | 23000.5586 | 23000.6543 | 1.1149 | 50 | -0.0027773 | 0.017339 | 1.4551e-05 | 1.3536e-07 | 2.5629e-07 | -5.1891e-08 | 50.0439 | 22.9301 | 0.019168 | 1.3651e-05 | 9.8725e-08 | 2.6171e-07 | 1.1681e-07 | 49.9734 | -49.1533 | 3.4028e+38 | 3.4028e+38 | 23000.5586 | 23000.6543 | 1.1149 | 3.4028e+38 | 3.4028e+38 | 230.0037 | 100.0009 | 23000.623 | 23000.6367 | 1.1596 | 50 | -0.0028683 | 0.01799 | 1.4859e-05 | 1.3279e-07 | 2.6724e-07 | -9.1686e-08 | 49.9833 | 36.2061 | 0.01977 | 1.3323e-05 | 1.3136e-07 | 2.6329e-07 | 9.0332e-08 | 50.0037 | -34.6518 | 3.4028e+38 | 3.4028e+38 | 23000.623 | 23000.6367 | 1.1596 | 3.4028e+38 | 3.4028e+38 | 230.0037 | 100.001 | 23000.5684 | 23000.582 | 1.143 | 50.0001 | -0.0028474 | 0.017855 | 1.4728e-05 | 1.3436e-07 | 2.6248e-07 | -8.0536e-08 | 49.9669 | 31.1248 | 0.019621 | 1.344e-05 | 1.2053e-07 | 2.5877e-07 | 1.0274e-07 | 49.9883 | -40.3061 | 3.4028e+38 | 3.4028e+38 | 23000.5684 | 23000.582 | 1.143 | 3.4028e+38 | 3.4028e+38 | 230.0035 | 100.001 | 23000.5918 | 23000.6055 | 1.1173 | 49.9999 | -0.0027917 | 0.017368 | 1.4592e-05 | 1.3556e-07 | 2.5335e-07 | -5.04e-08 | 49.9727 | 20.7702 | 0.01886 | 1.3688e-05 | 9.563e-08 | 2.6066e-07 | 1.1897e-07 | 49.9929 | -50.5006 | 3.4028e+38 | 3.4028e+38 | 23000.5918 | 23000.6055 | 1.1173 | 3.4028e+38 | 3.4028e+38 | 230.004 | 100.001 | 23000.5664 | 23000.584 | 1.1872 | 49.9999 | -0.0029573 | 0.015019 | 1.4714e-05 | 8.0426e-08 | 2.2228e-07 | -1.2947e-08 | 50.0085 | 10.0908 | 0.015872 | 1.3762e-05 | 3.9145e-08 | 2.2062e-07 | 6.6904e-08 | 49.9878 | -59.5853 | 3.4028e+38 | 3.4028e+38 | 23000.5664 | 23000.584 | 1.1872 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0044 | 100.0009 | 23000.6465 | 23000.6582 | 1.1648 | 50 | -0.0029382 | 1 | 3.4028e+38 | 0 | 230.0032 | 23000.5273 | 23000.5391 | 1.7196 | -0.0026778 | 0.014687 | 1.4722e-05 | 7.2761e-08 | 2.1469e-07 | -1.4638e-08 | 50.0063 | 13.2331 | 0.33112 | 3.4028e+38 | 121.183 | 0.020227 | 0 | 0 | 0 | -0.00072461 | 0.015607 | 1.3645e-05 | 4.1339e-08 | 2.1196e-07 | 4.3012e-08 | 50.1009 | -46.1682 | 0.19482 | 3.4028e+38 | -103.9812 | 230.1202 | -3.2446e-09 | 0.003135 | -1.4064e-09 | -0.0012313 | 3.4028e+38 | 3.4028e+38 | 23000.6875 | 23000.7031 | 1.1628 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 23000.6719 | 23000.6719 | 1.5477 | 0.021525 | 230.0038 | 100.0009 | 23000.5816 | 23000.6125 | 1.1444 | 50 | -0.0028484 | [17 empty] | 30.7551 | -5.4532 | 24.717 | 26.0644 | 7238049.4639 |  | -0.0025023
128 | 230.0043 | 100.0011 | 23000.6523 | 23000.668 | 1.1595 | 49.9999 | -0.0028883 | 0.017771 | 1.5152e-05 | 1.0652e-07 | 2.6767e-07 | -1.1402e-07 | 50.0516 | 46.3983 | 0.01931 | 1.3202e-05 | 1.4177e-07 | 2.5395e-07 | 6.2276e-08 | 50.0487 | -23.6745 | 3.4028e+38 | 3.4028e+38 | 23000.6523 | 23000.668 | 1.1595 | 3.4028e+38 | 3.4028e+38 | 230.0043 | 100.0009 | 23000.5996 | 23000.6152 | 1.1214 | 50 | -0.002769 | 0.015838 | 1.4624e-05 | 9.7075e-08 | 2.2689e-07 | -1.3079e-08 | 49.9849 | 7.6711 | 0.016602 | 1.381e-05 | 4.5948e-08 | 2.2929e-07 | 8.2908e-08 | 50.0188 | -61.9513 | 3.4028e+38 | 3.4028e+38 | 23000.5996 | 23000.6152 | 1.1214 | 3.4028e+38 | 3.4028e+38 | 230.004 | 100.0009 | 23000.582 | 23000.5938 | 1.1736 | 50 | -0.0029235 | 0.015471 | 1.4649e-05 | 9.8832e-08 | 2.2286e-07 | -1.2998e-08 | 50.022 | 8.0004 | 0.01649 | 1.3745e-05 | 4.3714e-08 | 2.277e-07 | 8.3053e-08 | 49.9667 | -62.3896 | 3.4028e+38 | 3.4028e+38 | 23000.582 | 23000.5938 | 1.1736 | 3.4028e+38 | 3.4028e+38 | 230.0038 | 100.0009 | 23000.5664 | 23000.5508 | 1.1143 | 49.9999 | -0.0027758 | 0.017963 | 1.4848e-05 | 1.3303e-07 | 2.666e-07 | -8.762e-08 | 50.0497 | 33.1843 | 0.019553 | 1.3355e-05 | 1.2686e-07 | 2.6117e-07 | 9.7103e-08 | 50.0383 | -38.3298 | 3.4028e+38 | 3.4028e+38 | 23000.5664 | 23000.5508 | 1.1143 | 3.4028e+38 | 3.4028e+38 | 230.0037 | 100.0009 | 23000.5781 | 23000.5957 | 1.0872 | 50 | -0.0029847 | 0.017402 | 1.4663e-05 | 1.3514e-07 | 2.6312e-07 | -8.2238e-08 | 50.0051 | 38.0254 | 0.019531 | 1.3519e-05 | 1.1651e-07 | 2.6409e-07 | 9.8737e-08 | 49.968 | -38.0789 | 3.4028e+38 | 3.4028e+38 | 23000.5781 | 23000.5957 | 1.0872 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0038 | 100.001 | 23000.6094 | 23000.6211 | 1.1974 | 49.9999 | -0.0029579 | 1 | 3.4028e+38 | 0 | 230.0087 | 23001.0996 | 23001.1133 | 1.8066 | 0.044319 | 0.015407 | 1.5415e-05 | 4.7107e-08 | 2.3748e-07 | -8.6314e-08 | 50.0137 | 58.0791 | 0.18357 | 3.4028e+38 | 112.2246 | 0.019674 | 0 | 0 | 0 | -0.00073391 | 0.015666 | 1.3627e-05 | 4.2133e-08 | 2.1345e-07 | 4.2571e-08 | 49.9988 | -57.0555 | 0.18267 | 3.4028e+38 | 132.1543 | 230.0359 | -1.3499e-09 | 0.00317 | 7.674e-09 | -0.0012878 | 3.4028e+38 | 3.4028e+38 | 23000.5273 | 23000.543 | 1.1592 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 23000.6543 | 23000.6543 | 2.1988 | 0.042824 | 230.004 | 100.0009 | 23000.5957 | 23000.6047 | 1.1312 | 49.9999 | -0.0028682 | [17 empty] | 31.8392 | -5.305 | 25.5893 | 25.9864 | 4712521.5149 |  | -0.002375
129 | 230.0068 | 100.001 | 11499.8047 | 23000.9355 | 19919.7988 | 50 | -60.0019 | 0.016818 | 1.5035e-05 | 1.3024e-07 | 2.481e-07 | -2.8597e-09 | 49.9745 | 3.337 | 0.018228 | 1.3089e-05 | 4.0429e-08 | 2.3144e-07 | 1.1214e-07 | 49.954 | -69.7003 | 3.4028e+38 | 3.4028e+38 | 11499.8047 | 23000.9355 | 19919.7969 | 3.4028e+38 | 3.4028e+38 | 230.0067 | 100.0009 | 11499.7646 | 23000.8867 | 19919.7559 | 50 | -60.002 | 0.016899 | 1.4631e-05 | 1.0309e-07 | 2.4981e-07 | 9.3631e-08 | 50.008 | -40.1237 | 0.018279 | 1.3641e-05 | -5.4104e-08 | 2.5423e-07 | 1.3247e-07 | 49.9513 | -112.2169 | 3.4028e+38 | 3.4028e+38 | 11499.7637 | 23000.8867 | 19919.7559 | 3.4028e+38 | 3.4028e+38 | 230.0068 | 100.0009 | 11499.8066 | 23000.8672 | 19919.7109 | 50 | -60.0018 | 0.017821 | 1.453e-05 | 1.2306e-07 | 2.5722e-07 | 8.641e-08 | 50.9727 | -33.9543 | 0.019014 | 1.3667e-05 | -4.0884e-08 | 2.5985e-07 | 1.5155e-07 | 50.0845 | -105.588 | 3.4028e+38 | 3.4028e+38 | 11499.8057 | 23000.8672 | 19919.7109 | 3.4028e+38 | 3.4028e+38 | 230.007 | 100.001 | 11499.8105 | 23000.9336 | 19919.8047 | 49.9999 | -60.0018 | 0.016491 | 1.5117e-05 | 1.2975e-07 | 2.4596e-07 | -9.3062e-09 | 49.9942 | 4.4874 | 0.017454 | 1.3029e-05 | 4.1679e-08 | 2.2745e-07 | 9.7324e-08 | 50.0382 | -67.6725 | 3.4028e+38 | 3.4028e+38 | 11499.8105 | 23000.9336 | 19919.8027 | 3.4028e+38 | 3.4028e+38 | 230.0074 | 100.0009 | 11499.8184 | 23000.9141 | 19919.7969 | 49.9999 | -60.002 | 0.01569 | 1.4953e-05 | 6.9883e-08 | 2.3584e-07 | 8.5108e-08 | 49.9397 | -50.4092 | 0.016843 | 1.3543e-05 | -5.4932e-08 | 2.2628e-07 | 8.7466e-08 | 49.9678 | -122.0948 | 3.4028e+38 | 3.4028e+38 | 11499.8174 | 23000.9141 | 19919.7969 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0069 | 100.0009 | 11499.8145 | 23000.8711 | 19919.707 | 50 | -60.002 | 0.49997 | 3.4028e+38 | 0 | 230.0051 | 11499.0459 | 23000.7188 | 19919.9766 | -0.029957 | 0.015695 | 1.5065e-05 | 6.9664e-08 | 2.3473e-07 | 8.0353e-08 | 49.9974 | -49.0755 | 0.27361 | 3.4028e+38 | -171.2263 | 0.020521 | 0 | 0 | 0 | 0.0011915 | 0.016217 | 1.3409e-05 | -4.6427e-08 | 2.1743e-07 | 6.8573e-08 | 50.0186 | -122.9279 | -0.20509 | 3.4028e+38 | 95.6281 | 230.1226 | -7.7962e-09 | 0.0030602 | -3.3076e-09 | -0.0016604 | 3.4028e+38 | 3.4028e+38 | 11499.7959 | 23000.9336 | 19919.8477 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 11499.3447 | 11499.3447 | 19920 | -0.00057451 | 230.007 | 100.0009 | 11499.801 | 23000.9074 | 19919.7734 | 50 | -60.0019 | [17 empty] | 44.8046 | 24.6828 | -44.9747 | 69.4383 | 108.8995 |  | -0.0038022
130 | 230.0072 | 100.0008 | 11499.7539 | 23000.8594 | 19919.7832 | 49.9999 | -60.0021 | 0.017817 | 1.472e-05 | 1.5095e-07 | 2.6316e-07 | 4.0684e-08 | 49.9705 | -17.6821 | 0.019346 | 1.3405e-05 | 9.014e-09 | 2.6227e-07 | 1.5851e-07 | 49.9714 | -88.3598 | 3.4028e+38 | 3.4028e+38 | 11499.7539 | 23000.8594 | 19919.7832 | 3.4028e+38 | 3.4028e+38 | 230.0071 | 100.0009 | 11499.8115 | 23000.9473 | 19919.8008 | 49.9999 | -60.0019 | 0.015312 | 1.5054e-05 | 6.0379e-08 | 2.3062e-07 | 7.3328e-08 | 50.0038 | -50.431 | 0.016491 | 1.3425e-05 | -4.58e-08 | 2.1978e-07 | 7.1904e-08 | 50.0052 | -122.8491 | 3.4028e+38 | 3.4028e+38 | 11499.8105 | 23000.9473 | 19919.8008 | 3.4028e+38 | 3.4028e+38 | 230.007 | 100.0008 | 11499.7686 | 23000.8848 | 19919.7949 | 49.9999 | -60.0019 | 0.018006 | 1.4579e-05 | 1.3424e-07 | 2.6051e-07 | 7.9902e-08 | 49.9932 | -32.3689 | 0.019608 | 1.3575e-05 | -3.2934e-08 | 2.6294e-07 | 1.5454e-07 | 49.9828 | -102.1683 | 3.4028e+38 | 3.4028e+38 | 11499.7676 | 23000.8848 | 19919.7949 | 3.4028e+38 | 3.4028e+38 | 230.0073 | 100.0008 | 11499.8223 | 23000.8984 | 19919.8398 | 49.9999 | -60.0021 | 0.017511 | 1.4623e-05 | 1.1684e-07 | 2.5647e-07 | 8.7331e-08 | 50.0032 | -35.8304 | 0.019067 | 1.3638e-05 | -4.6519e-08 | 2.5803e-07 | 1.4651e-07 | 49.9968 | -106.0371 | 3.4028e+38 | 3.4028e+38 | 11499.8213 | 23000.8984 | 19919.8398 | 3.4028e+38 | 3.4028e+38 | 230.0073 | 100.0009 | 11499.7451 | 23000.9141 | 19919.8223 | 50 | -60.0019 | 0.015829 | 1.5024e-05 | 7.1703e-08 | 2.3788e-07 | 8.1402e-08 | 49.9816 | -50.3559 | 0.016915 | 1.3517e-05 | -5.5524e-08 | 2.2296e-07 | 8.6747e-08 | 50.0223 | -120.8591 | 3.4028e+38 | 3.4028e+38 | 11499.7441 | 23000.9141 | 19919.8223 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0072 | 100.001 | 11499.7705 | 23000.9355 | 19919.8555 | 49.9999 | -60.0019 | 0.49997 | 3.4028e+38 | 0 | 230.0059 | 11499.2441 | 23000.8164 | 19920.0078 | -0.01583 | 0.015088 | 1.5202e-05 | 5.6874e-08 | 2.2497e-07 | 6.1534e-08 | 49.9218 | -45.4076 | 0.25686 | 3.4028e+38 | -139.5552 | 0.02026 | 0 | 0 | 0 | 0.0012003 | 0.015522 | 1.3205e-05 | -2.5152e-08 | 2.0498e-07 | 5.6277e-08 | 49.9382 | -112.8392 | -0.11281 | 3.4028e+38 | 126.1983 | 230.0586 | -3.8492e-09 | 0.0030181 | -6.3267e-09 | -0.0016345 | 3.4028e+38 | 3.4028e+38 | 11499.7568 | 23000.9316 | 19919.8184 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 11499.4492 | 11499.4492 | 19920.0977 | 0.015362 | 230.0072 | 100.0008 | 11499.7803 | 23000.9008 | 19919.8082 | 49.9999 | -60.002 | [17 empty] | 45.6858 | 28.665 | -50.8233 | 74.0099 | 118.4762 |  | -0.0041854
131 | 230.0073 | 100.0008 | 11499.7803 | 23000.9238 | 19919.8145 | 49.9999 | -60.002 | 0.015211 | 1.5183e-05 | 5.975e-08 | 2.286e-07 | 6.1362e-08 | 49.9555 | -47.1285 | 0.016157 | 1.3364e-05 | -4.2715e-08 | 2.0878e-07 | 6.287e-08 | 50.0399 | -120.5167 | 3.4028e+38 | 3.4028e+38 | 11499.7793 | 23000.9238 | 19919.8145 | 3.4028e+38 | 3.4028e+38 | 230.0071 | 100.0008 | 11499.7529 | 23000.8711 | 19919.7832 | 49.9999 | -60.002 | 0.017747 | 1.4711e-05 | 1.5056e-07 | 2.609e-07 | 4.6461e-08 | 50.0064 | -16.926 | 0.019528 | 1.3468e-05 | 4.4214e-09 | 2.6535e-07 | 1.581e-07 | 50.0387 | -88.7947 | 3.4028e+38 | 3.4028e+38 | 11499.7529 | 23000.8711 | 19919.7832 | 3.4028e+38 | 3.4028e+38 | 230.0071 | 100.0008 | 11499.7812 | 23000.9141 | 19919.8945 | 49.9999 | -60.0022 | 0.015448 | 1.5148e-05 | 6.1555e-08 | 2.2965e-07 | 7.4459e-08 | 50.0129 | -50.5597 | 0.016472 | 1.3367e-05 | -4.8228e-08 | 2.1888e-07 | 7.3081e-08 | 50.0323 | -121.83 | 3.4028e+38 | 3.4028e+38 | 11499.7803 | 23000.9141 | 19919.8945 | 3.4028e+38 | 3.4028e+38 | 230.0073 | 100.0009 | 11499.75 | 23000.9395 | 19919.8047 | 49.9999 | -60.002 | 0.014659 | 1.5357e-05 | 5.9201e-08 | 2.2513e-07 | 4.4724e-08 | 49.9825 | -38.3557 | 0.015577 | 1.3156e-05 | -2.3212e-08 | 2.0738e-07 | 5.4784e-08 | 49.9831 | -114.4906 | 3.4028e+38 | 3.4028e+38 | 11499.749 | 23000.9395 | 19919.8047 | 3.4028e+38 | 3.4028e+38 | 230.0074 | 100.0009 | 11499.7559 | 23000.8887 | 19919.7988 | 49.9999 | -60.002 | 0.01671 | 1.515e-05 | 1.3471e-07 | 2.532e-07 | -1.4364e-09 | 49.9673 | -0.17449 | 0.017747 | 1.307e-05 | 3.8807e-08 | 2.318e-07 | 1.1502e-07 | 50.014 | -72.2073 | 3.4028e+38 | 3.4028e+38 | 11499.7559 | 23000.8887 | 19919.7969 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0074 | 100.0008 | 11499.833 | 23000.9141 | 19919.7891 | 49.9999 | -60.0019 | 0.49997 | 3.4028e+38 | 0 | 230.0123 | 11499.2832 | 23001.4219 | 19920.709 | -0.025165 | 0.0176 | 1.4852e-05 | 1.5071e-07 | 2.6117e-07 | 2.7465e-08 | 50.0168 | -11.4626 | 0.57778 | 3.4028e+38 | -157.1983 | 0.021059 | 0 | 0 | 0 | 0.0010763 | 0.019342 | 1.3433e-05 | 5.5222e-09 | 2.6424e-07 | 1.609e-07 | 50.0128 | -90.7433 | -0.0079001 | 3.4028e+38 | 133.1013 | 230.0053 | -2.0633e-09 | 0.0031309 | 1.6255e-08 | -0.0016815 | 3.4028e+38 | 3.4028e+38 | 11499.7686 | 23000.8555 | 19919.7793 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 11498.8779 | 11498.8779 | 19920.4492 | -0.05766 | 230.0072 | 100.0009 | 11499.7641 | 23000.9074 | 19919.8191 | 49.9999 | -60.002 | [17 empty] | 46.1499 | 26.8412 | -50.7704 | 72.4213 | 116.0341 |  | -0.0041115
132 | 230 | 100.0007 | 18393.5898 | 23000.1914 | -13809.124 | 50 | 36.8977 | 0.016745 | 1.4716e-05 | 8.0081e-08 | 2.4445e-07 | -9.9147e-08 | 49.9907 | 50.5635 | 0.018078 | 1.3665e-05 | 1.2618e-07 | 2.4623e-07 | 4.8598e-08 | 49.9987 | -21.6133 | 3.4028e+38 | 3.4028e+38 | 18393.5898 | 23000.1914 | -13809.125 | 3.4028e+38 | 3.4028e+38 | 230.0002 | 100.0007 | 18393.5918 | 23000.125 | -13809.1133 | 49.9999 | 36.8977 | 0.015481 | 1.4632e-05 | 7.2202e-08 | 2.2948e-07 | -7.1686e-08 | 49.9201 | 45.1115 | 0.016698 | 1.3805e-05 | 9.3859e-08 | 2.282e-07 | 4.3526e-08 | 50.0285 | -26.3598 | 3.4028e+38 | 3.4028e+38 | 18393.5918 | 23000.125 | -13809.1143 | 3.4028e+38 | 3.4028e+38 | 230.0002 | 100.0007 | 18393.627 | 23000.1719 | -13809.1494 | 50 | 36.8974 | 0.016801 | 1.5496e-05 | 2.4705e-09 | 2.6615e-07 | -1.4504e-07 | 49.8881 | 86.6277 | 0.018526 | 1.3171e-05 | 1.2943e-07 | 2.4543e-07 | -5.6078e-08 | 49.9601 | 19.8284 | 3.4028e+38 | 3.4028e+38 | 18393.627 | 23000.1719 | -13809.1504 | 3.4028e+38 | 3.4028e+38 | 230.0005 | 100.0008 | 18393.6641 | 23000.2227 | -13809.1426 | 50 | 36.8976 | 0.01673 | 1.5452e-05 | -8.1653e-09 | 2.5773e-07 | -1.3548e-07 | 50.0202 | 93.449 | 0.018463 | 1.3127e-05 | 1.238e-07 | 2.3844e-07 | -6.0863e-08 | 50.0674 | 27.4643 | 3.4028e+38 | 3.4028e+38 | 18393.6621 | 23000.2227 | -13809.1436 | 3.4028e+38 | 3.4028e+38 | 230.0007 | 100.0007 | 18393.6582 | 23000.2168 | -13809.1104 | 50 | 36.8975 | 0.014482 | 1.5031e-05 | 5.1264e-09 | 2.2178e-07 | -6.6352e-08 | 50.0084 | 87.8929 | 0.015261 | 1.3681e-05 | 6.1218e-08 | 2.0873e-07 | -2.5253e-08 | 49.9605 | 24.07 | 3.4028e+38 | 3.4028e+38 | 18393.6582 | 23000.2168 | -13809.1113 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0002 | 100.0007 | 18393.6289 | 23000.166 | -13809.0977 | 50 | 36.8975 | 0.79972 | 3.4028e+38 | 0 | 230.0018 | 18394.0078 | 23000.3633 | -13808.9072 | 0.023193 | 0.014345 | 1.4886e-05 | 1.6643e-08 | 2.1207e-07 | -5.8386e-08 | 49.97 | 70.0739 | 0.093333 | 3.4028e+38 | 138.4109 | 0.020125 | 0 | 0 | 0 | -0.0016404 | 0.015673 | 1.3818e-05 | 5.7952e-08 | 2.1654e-07 | -9.4599e-10 | 50.0234 | -3.8532 | 0.27349 | 3.4028e+38 | -153.0764 | 229.9784 | 3.448e-09 | 0.0031921 | -3.9236e-09 | -0.00026916 | 3.4028e+38 | 3.4028e+38 | 18393.6172 | 23000.1777 | -13809.1152 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 18393.875 | 18393.875 | -13808.7275 | 0.011173 | 230.0003 | 100.0007 | 18393.6262 | 23000.1855 | -13809.1279 | 50 | 36.8976 | [17 empty] | 15.1436 | 19.2833 | 49.7116 | 34.0072 | 29.6305 |  | -0.00059888
133 | 230.0006 | 100.0007 | 18393.6797 | 23000.2051 | -13809.1104 | 50 | 36.8975 | 0.01498 | 1.5216e-05 | -9.7626e-09 | 2.2794e-07 | -8.4589e-08 | 49.99 | 97.2634 | 0.016454 | 1.3426e-05 | 7.2766e-08 | 2.1559e-07 | -4.7282e-08 | 50.0203 | 33.2259 | 3.4028e+38 | 3.4028e+38 | 18393.6777 | 23000.2051 | -13809.1113 | 3.4028e+38 | 3.4028e+38 | 230.0002 | 100.0008 | 18393.6562 | 23000.2207 | -13809.0654 | 49.9999 | 36.8974 | 0.016785 | 1.5473e-05 | -6.148e-09 | 2.5479e-07 | -1.3165e-07 | 50.0388 | 93.8983 | 0.01837 | 1.3141e-05 | 1.1746e-07 | 2.4143e-07 | -5.8013e-08 | 50.0196 | 26.9868 | 3.4028e+38 | 3.4028e+38 | 18393.6543 | 23000.2207 | -13809.0664 | 3.4028e+38 | 3.4028e+38 | 230.0003 | 100.0007 | 18393.666 | 23000.2168 | -13809.1064 | 49.9999 | 36.8976 | 0.015871 | 1.4656e-05 | 7.6406e-08 | 2.3203e-07 | -7.4686e-08 | 49.9375 | 44.3999 | 0.017239 | 1.3786e-05 | 1.0334e-07 | 2.3481e-07 | 4.7267e-08 | 49.9751 | -25.1658 | 3.4028e+38 | 3.4028e+38 | 18393.666 | 23000.2168 | -13809.1074 | 3.4028e+38 | 3.4028e+38 | 230.0007 | 100.0007 | 18393.6035 | 23000.1406 | -13809.0742 | 49.9999 | 36.8977 | 0.01483 | 1.4668e-05 | 5.2115e-08 | 2.175e-07 | -5.5047e-08 | 50.0054 | 46.6822 | 0.015615 | 1.3796e-05 | 7.0772e-08 | 2.1983e-07 | 2.5527e-08 | 49.9759 | -20.9653 | 3.4028e+38 | 3.4028e+38 | 18393.6035 | 23000.1406 | -13809.0752 | 3.4028e+38 | 3.4028e+38 | 230.0003 | 100.0006 | 18393.5938 | 23000.1699 | -13809.126 | 49.9999 | 36.8976 | 0.015712 | 1.4666e-05 | 7.433e-08 | 2.3162e-07 | -7.4086e-08 | 49.9622 | 46.6286 | 0.017257 | 1.3803e-05 | 9.8682e-08 | 2.3113e-07 | 4.9303e-08 | 50.0046 | -25.4578 | 3.4028e+38 | 3.4028e+38 | 18393.5938 | 23000.1699 | -13809.127 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0004 | 100.0007 | 18393.6172 | 23000.1836 | -13809.1377 | 50 | 36.8976 | 0.79972 | 3.4028e+38 | 0 | 229.9982 | 18393.7812 | 22999.9434 | -13808.5518 | 0.0081628 | 0.015283 | 1.4692e-05 | 6.2632e-08 | 2.2709e-07 | -6.3515e-08 | 50.0028 | 45.0451 | 0.27578 | 3.4028e+38 | 84.7809 | 0.019698 | 0 | 0 | 0 | -0.0016484 | 0.016342 | 1.3816e-05 | 8.2412e-08 | 2.2743e-07 | 3.8456e-08 | 49.9907 | -25.0743 | 0.35915 | 3.4028e+38 | 90.9889 | 229.9841 | 6.0751e-09 | 0.0031746 | 3.7654e-09 | -0.00030183 | 3.4028e+38 | 3.4028e+38 | 18393.5762 | 23000.1504 | -13809.127 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 18393.7422 | 18393.7422 | -13808.5879 | -0.0027934 | 230.0004 | 100.0007 | 18393.6398 | 23000.1906 | -13809.0965 | 49.9999 | 36.8976 | [17 empty] | 15.7613 | 22.087 | 50.5953 | 37.256 | 35.5043 |  | -0.00038844
134 | 230.0002 | 100.0009 | 18393.6367 | 23000.1895 | -13809.123 | 49.9999 | 36.8975 | 0.017857 | 1.5413e-05 | 2.4282e-08 | 2.7524e-07 | -1.5907e-07 | 50.0156 | 82.0676 | 0.019591 | 1.3073e-05 | 1.5739e-07 | 2.5574e-07 | -3.0898e-08 | 50.0085 | 11.4831 | 3.4028e+38 | 3.4028e+38 | 18393.6367 | 23000.1895 | -13809.124 | 3.4028e+38 | 3.4028e+38 | 230.0002 | 100.0008 | 18393.6465 | 23000.1562 | -13809.0967 | 49.9999 | 36.8976 | 0.014708 | 1.5097e-05 | 4.3894e-09 | 2.1819e-07 | -6.3249e-08 | 49.9785 | 75.6553 | 0.015514 | 1.3667e-05 | 6.1603e-08 | 2.1203e-07 | -2.0876e-08 | 49.9933 | 19.9452 | 3.4028e+38 | 3.4028e+38 | 18393.6465 | 23000.1562 | -13809.0977 | 3.4028e+38 | 3.4028e+38 | 230.0003 | 100.0008 | 18393.6406 | 23000.1973 | -13809.1318 | 50 | 36.8974 | 0.017324 | 1.55e-05 | 1.0914e-08 | 2.7011e-07 | -1.5385e-07 | 50.0653 | 87.5997 | 0.019484 | 1.3092e-05 | 1.4423e-07 | 2.5233e-07 | -4.3798e-08 | 49.9988 | 18.5228 | 3.4028e+38 | 3.4028e+38 | 18393.6406 | 23000.1973 | -13809.1328 | 3.4028e+38 | 3.4028e+38 | 230.0008 | 100.0007 | 18393.6543 | 23000.1934 | -13809.085 | 49.9999 | 36.8975 | 0.014955 | 1.514e-05 | -6.4122e-09 | 2.2607e-07 | -8.3685e-08 | 50.0026 | 97.7619 | 0.0158 | 1.354e-05 | 6.7436e-08 | 2.1758e-07 | -4.4621e-08 | 50.0235 | 31.3881 | 3.4028e+38 | 3.4028e+38 | 18393.6523 | 23000.1934 | -13809.0859 | 3.4028e+38 | 3.4028e+38 | 230.0004 | 100.0006 | 18393.5996 | 23000.1602 | -13809.0684 | 49.9999 | 36.8976 | 0.016206 | 1.4716e-05 | 7.9991e-08 | 2.385e-07 | -9.1087e-08 | 49.9579 | 48.7098 | 0.017248 | 1.368e-05 | 1.0498e-07 | 2.3596e-07 | 4.9377e-08 | 50.0064 | -23.2084 | 3.4028e+38 | 3.4028e+38 | 18393.5996 | 23000.1602 | -13809.0693 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0005 | 100.0007 | 18393.6367 | 23000.2012 | -13809.1006 | 50 | 36.8976 | 0.79972 | 3.4028e+38 | 0 | 229.9991 | 18394.1133 | 23000.1035 | -13808.2529 | 0.068366 | 0.017423 | 1.5041e-05 | 6.6001e-08 | 2.6683e-07 | -1.409e-07 | 49.9543 | 64.9007 | 0.22568 | 3.4028e+38 | 139.7345 | 0.021133 | 0 | 0 | 0 | -0.0017453 | 0.019937 | 1.3153e-05 | 1.6327e-07 | 2.6239e-07 | -9.9853e-09 | 50.0261 | 3.5905 | 0.61083 | 3.4028e+38 | 80.0605 | 229.9215 | 2.0919e-08 | 0.003014 | -1.0505e-08 | -0.0002854 | 3.4028e+38 | 3.4028e+38 | 18393.6582 | 23000.2227 | -13809.1035 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 18394.5176 | 18394.5176 | -13808.6494 | 0.092995 | 230.0004 | 100.0008 | 18393.6355 | 23000.1793 | -13809.101 | 49.9999 | 36.8975 | [17 empty] | 15.9772 | 24.969 | 50.9782 | 39.5245 | 42.3953 |  | -0.00027895
135 | 230.0014 | 0.50002 | 115.0057 | 115.0056 | 0.0085727 | 50 | -0.0042709 | 0.016658 | 3.3863e-06 | 1.478e-09 | 5.7145e-08 | -2.2458e-09 | 74.8155 | 71.4142 | 0.018514 | 3.0248e-06 | 2.7625e-09 | 5.609e-08 | 7.2644e-10 | 74.7623 | -5.7542 | 3.4028e+38 | 3.4028e+38 | 115.0057 | 115.0056 | 0.0085727 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 0.50002 | 115.0059 | 115.0058 | 0.0091785 | 49.9999 | -0.0045727 | 0.015347 | 3.3806e-06 | -4.9159e-11 | 5.3402e-08 | -1.9636e-09 | 74.7582 | 103.5105 | 0.016648 | 3.1003e-06 | 1.3495e-09 | 5.1561e-08 | 2.4181e-10 | 74.7153 | -3.5285 | 3.4028e+38 | 3.4028e+38 | 115.0059 | 115.0058 | 0.0091785 | 3.4028e+38 | 3.4028e+38 | 230.002 | 0.50003 | 115.0068 | 115.0068 | 0.008436 | 49.9999 | -0.0045614 | 0.016065 | 3.3211e-06 | -6.3275e-10 | 5.3879e-08 | -2.7311e-09 | 74.7033 | 113.4874 | 0.017535 | 3.0244e-06 | 1.2594e-09 | 5.3774e-08 | -8.3745e-10 | 74.7476 | 33.6224 | 3.4028e+38 | 3.4028e+38 | 115.0068 | 115.0068 | 0.008436 | 3.4028e+38 | 3.4028e+38 | 230.002 | 0.50002 | 115.0057 | 115.0057 | 0.0089763 | 49.9999 | -0.004472 | 0.015233 | 3.3189e-06 | 8.2887e-10 | 5.0534e-08 | -1.5345e-09 | 74.7639 | 50.8812 | 0.016465 | 3.0122e-06 | 1.3362e-09 | 4.9504e-08 | 8.2819e-11 | 74.8216 | 9.5773 | 3.4028e+38 | 3.4028e+38 | 115.0057 | 115.0057 | 0.0089763 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 0.50002 | 115.0062 | 115.0062 | 0.0081843 | 49.9999 | -0.0043964 | 0.016221 | 3.429e-06 | -9.5599e-10 | 5.4721e-08 | -2.5154e-09 | 74.6627 | 100.6184 | 0.017711 | 3.0738e-06 | 1.0813e-09 | 5.4297e-08 | -1.265e-09 | 74.8088 | 51.3816 | 3.4028e+38 | 3.4028e+38 | 115.0062 | 115.0062 | 0.0081843 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0016 | 0.50002 | 115.0059 | 115.0058 | 0.008409 | 49.9999 | -0.0041894 | 1 | 3.4028e+38 | 0 | 230.0066 | 115.0089 | 115.0088 | 0.011921 | 8.5904e-05 | 0.015518 | 3.3676e-06 | -9.6484e-10 | 5.2198e-08 | -2.5382e-09 | 74.7233 | 113.0674 | -0.019515 | 3.4028e+38 | -18.2768 | 0.019787 | 0 | 0 | 0 | -3.9753e-05 | 0.016267 | 3.0514e-06 | -4.8813e-10 | 4.8611e-08 | -1.2787e-09 | 74.8204 | 108.0736 | -0.010233 | 3.4028e+38 | -11.1183 | 229.4038 | -7.2501e-10 | 0.00071734 | 1.4163e-10 | 2.4191e-05 | 3.4028e+38 | 3.4028e+38 | 115.0059 | 115.0063 | 0.0091929 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 115.0067 | 115.0067 | 0.010801 | -8.8281e-05 | 230.0018 | 0.50002 | 115.0061 | 115.006 | 0.0086696 | 49.9999 | -0.0044547 | [17 empty] | 23.2691 | 7.7351 | 31.0856 | 30.6923 | 1458998.7259 |  | -0.0026982
136 | 230.0019 | 0.50002 | 115.0061 | 115.006 | 0.0093195 | 49.9999 | -0.0043542 | 0.0146 | 3.4347e-06 | -9.1018e-10 | 5.0186e-08 | -1.3459e-09 | 74.6721 | 87.1979 | 0.015905 | 3.1297e-06 | 2.6098e-10 | 4.9626e-08 | -1.2519e-09 | 74.8406 | 65.3926 | 3.4028e+38 | 3.4028e+38 | 115.0061 | 115.006 | 0.0093195 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 0.50002 | 115.0054 | 115.0054 | 0.0082649 | 50 | -0.0041176 | 0.015878 | 3.3597e-06 | 6.5063e-10 | 5.3178e-08 | -1.9695e-09 | 74.7553 | 71.7189 | 0.017395 | 3.0341e-06 | 2.8805e-09 | 5.3306e-08 | 1.0966e-09 | 74.8085 | -20.8417 | 3.4028e+38 | 3.4028e+38 | 115.0054 | 115.0054 | 0.0082649 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 0.50002 | 115.0063 | 115.0062 | 0.0084833 | 49.9999 | -0.0042061 | 0.017736 | 3.3359e-06 | 7.71e-10 | 5.914e-08 | -3.3277e-09 | 74.6596 | 78.3623 | 0.019604 | 2.9798e-06 | 3.6652e-09 | 5.8434e-08 | 2.5362e-10 | 74.7651 | 3.239 | 3.4028e+38 | 3.4028e+38 | 115.0063 | 115.0062 | 0.0084833 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 0.50002 | 115.0061 | 115.0057 | 0.0087338 | 49.9999 | -0.0043512 | 0.014987 | 3.3449e-06 | -1.4444e-09 | 4.9349e-08 | -1.6813e-09 | 74.7491 | 106.061 | 0.016041 | 3.0707e-06 | 8.8266e-11 | 4.8196e-08 | -1.3792e-09 | 74.7796 | 84.0247 | 3.4028e+38 | 3.4028e+38 | 115.0061 | 115.0057 | 0.0087338 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 0.50002 | 115.006 | 115.0059 | 0.0082252 | 49.9999 | -0.0042659 | 0.017847 | 3.3687e-06 | 5.5229e-10 | 5.9062e-08 | -4.2338e-09 | 74.7433 | 66.9691 | 0.020035 | 3.1903e-06 | 2.6908e-09 | 6.257e-08 | 1.9048e-10 | 74.7967 | 29.6948 | 3.4028e+38 | 3.4028e+38 | 115.006 | 115.0059 | 0.0082252 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 0.50002 | 115.0059 | 115.006 | 0.0083602 | 49.9999 | -0.004165 | 1 | 3.4028e+38 | 0 | 230.0048 | 115.0074 | 115.0074 | 0.013701 | 0.0001247 | 0.017806 | 3.3698e-06 | -6.7925e-10 | 6.0081e-08 | -4.2701e-09 | 74.739 | 100.3847 | -0.011297 | 3.4028e+38 | -43.6939 | 0.020624 | 0 | 0 | 0 | -5.3824e-05 | 0.019704 | 2.9809e-06 | 2.6964e-09 | 5.8738e-08 | -1.1874e-09 | 74.762 | 21.3457 | 0.050517 | 3.4028e+38 | -29.6162 | 229.2373 | 3.911e-10 | 0.00068618 | 1.8039e-10 | 5.1615e-06 | 3.4028e+38 | 3.4028e+38 | 115.006 | 115.006 | 0.0086824 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 115.0083 | 115.0083 | 0.013276 | 0.00014106 | 230.0019 | 0.50002 | 115.006 | 115.0058 | 0.0086053 | 49.9999 | -0.004259 | [17 empty] | 22.8519 | 6.3494 | 29.619 | 28.4429 | 1233383.7319 |  | -0.0023394
137 | 230.0018 | 1 | 114.992 | 230.0099 | 199.2024 | 49.9999 | -60.0038 | 0.017398 | 3.4336e-06 | 5.2063e-09 | 5.9747e-08 | -1.8539e-09 | 74.7272 | 19.6001 | 0.0194 | 3.1892e-06 | 5.4955e-09 | 6.0922e-08 | -2.0132e-10 | 74.7577 | -1.7253 | 3.4028e+38 | 3.4028e+38 | 114.992 | 230.0099 | 199.2024 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 1 | 114.9919 | 230.0107 | 199.2029 | 49.9999 | -60.0038 | 0.014939 | 3.4075e-06 | 2.3154e-09 | 5.237e-08 | -1.893e-09 | 74.7585 | 28.4449 | 0.016385 | 2.9871e-06 | 7.3394e-10 | 4.8924e-08 | -7.2264e-10 | 74.8073 | 23.3932 | 3.4028e+38 | 3.4028e+38 | 114.9919 | 230.0107 | 199.2029 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 1 | 114.9911 | 230.0097 | 199.2025 | 50 | -60.0038 | 0.016202 | 3.4408e-06 | 6.4195e-09 | 5.5666e-08 | 2.0637e-09 | 74.7442 | -22.6433 | 0.017703 | 3.1455e-06 | 5.6279e-09 | 5.3994e-08 | 2.0075e-09 | 74.7287 | -20.8661 | 3.4028e+38 | 3.4028e+38 | 114.9911 | 230.0097 | 199.2025 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 1 | 114.9924 | 230.0107 | 199.2033 | 49.9999 | -60.004 | 0.014729 | 3.4497e-06 | 2.9394e-09 | 5.0798e-08 | -1.4281e-09 | 74.6861 | 23.9463 | 0.016144 | 3.0532e-06 | 9.3814e-10 | 5.0398e-08 | -1.1279e-09 | 74.7491 | 50.2483 | 3.4028e+38 | 3.4028e+38 | 114.9924 | 230.0107 | 199.2033 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 1 | 114.9922 | 230.0107 | 199.2031 | 50 | -60.0039 | 0.014425 | 3.4365e-06 | 1.5492e-09 | 5.0996e-08 | -2.2045e-10 | 74.7174 | 4.507 | 0.015728 | 3.0374e-06 | 2.3076e-09 | 4.8363e-08 | -1.1781e-09 | 74.7697 | 47.2749 | 3.4028e+38 | 3.4028e+38 | 114.9922 | 230.0107 | 199.2031 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0014 | 1 | 114.9921 | 230.0105 | 199.2025 | 50 | -60.0038 | 0.49994 | 3.4028e+38 | 0 | 230.0062 | 114.9894 | 230.0148 | 199.2074 | -0.00021294 | 0.015711 | 3.4746e-06 | 4.2797e-09 | 5.5099e-08 | -2.1482e-09 | 74.7312 | 32.117 | 0.078642 | 3.4028e+38 | 25.8782 | 0.020192 | 0 | 0 | 0 | -2.3182e-05 | 0.017633 | 3.011e-06 | 2.0707e-09 | 5.3042e-08 | -9.1007e-10 | 74.7545 | 22.1375 | 0.040003 | 3.4028e+38 | 54.1137 | 229.1186 | -3.4396e-10 | 0.00068846 | -4.5265e-10 | 4.3523e-06 | 3.4028e+38 | 3.4028e+38 | 114.992 | 230.01 | 199.2022 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 114.9883 | 114.9883 | 199.208 | -0.00037048 | 230.0017 | 1 | 114.9919 | 230.0103 | 199.2029 | 49.9999 | -60.0039 | [17 empty] | 21.7782 | 46.2772 | -90.9308 | 68.4872 | 123.6985 |  | -0.0053207
138 | 230.0018 | 1 | 114.9911 | 230.0098 | 199.2023 | 50 | -60.0038 | 0.015537 | 3.4552e-06 | 4.9281e-09 | 5.372e-08 | 1.9172e-09 | 74.6427 | -21.2576 | 0.016832 | 3.1452e-06 | 4.5647e-09 | 5.2915e-08 | 2.3471e-09 | 74.7607 | -33.581 | 3.4028e+38 | 3.4028e+38 | 114.9911 | 230.0098 | 199.2023 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | 114.9923 | 230.0104 | 199.2027 | 49.9999 | -60.0039 | 0.014408 | 3.4174e-06 | 1.7889e-09 | 4.935e-08 | -6.0215e-10 | 74.7318 | -0.029519 | 0.015515 | 3.0372e-06 | 1.5601e-09 | 4.7175e-08 | -4.7274e-10 | 74.7607 | 13.9962 | 3.4028e+38 | 3.4028e+38 | 114.9923 | 230.0104 | 199.2027 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | 114.9918 | 230.0103 | 199.203 | 49.9999 | -60.0039 | 0.01433 | 3.4819e-06 | 3.4908e-09 | 4.9911e-08 | -1.4926e-09 | 74.6933 | 26.3778 | 0.015775 | 3.1139e-06 | 1.488e-09 | 4.7642e-08 | -3.2785e-10 | 74.7267 | 12.4251 | 3.4028e+38 | 3.4028e+38 | 114.9918 | 230.0103 | 199.203 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 1 | 114.9924 | 230.0102 | 199.2026 | 49.9999 | -60.0037 | 0.015956 | 3.4573e-06 | 5.1772e-09 | 5.5358e-08 | -1.7686e-09 | 74.7009 | 21.886 | 0.017602 | 3.151e-06 | 2.4521e-09 | 5.4984e-08 | -5.1322e-10 | 74.6367 | 43.2744 | 3.4028e+38 | 3.4028e+38 | 114.9924 | 230.0102 | 199.2026 | 3.4028e+38 | 3.4028e+38 | 230.0009 | 1 | 114.9921 | 230.0095 | 199.2028 | 49.9999 | -60.0039 | 0.016231 | 3.4064e-06 | 4.6969e-09 | 5.6427e-08 | -3.0553e-09 | 74.6894 | 25.5074 | 0.017994 | 3.0042e-06 | 3.471e-09 | 5.7765e-08 | -7.0468e-11 | 74.7297 | 0.89159 | 3.4028e+38 | 3.4028e+38 | 114.9921 | 230.0095 | 199.2028 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0015 | 1 | 114.9919 | 230.0098 | 199.2024 | 49.9999 | -60.0038 | 0.49995 | 3.4028e+38 | 0 | 230.0057 | 114.9862 | 230.0132 | 199.2057 | -0.00057756 | 0.017553 | 3.4854e-06 | 6.3534e-09 | 6.1176e-08 | -1.3059e-09 | 74.7031 | 11.1745 | 0.10703 | 3.4028e+38 | 63.7322 | 0.02111 | 0 | 0 | 0 | -3.3747e-05 | 0.019795 | 3.1308e-06 | 5.4003e-09 | 6.0269e-08 | 1.4339e-09 | 74.6813 | -14.8703 | 0.089294 | 3.4028e+38 | -53.567 | 229.1714 | 5.2583e-10 | 0.00069738 | 8.1782e-10 | 6.2224e-06 | 3.4028e+38 | 3.4028e+38 | 114.9916 | 230.009 | 199.2018 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 114.9828 | 114.9833 | 199.2014 | -0.0008583 | 230.0016 | 1 | 114.9919 | 230.0101 | 199.2027 | 49.9999 | -60.0038 | [17 empty] | 20.7591 | 44.4687 | -90.1533 | 64.7387 | 119.2669 |  | -0.0052064
139 | 230.0013 | 1 | -0.016222 | 230.0104 | 230.0101 | 50 | -90.0039 | 0.017513 | 3.4065e-06 | 6.5806e-09 | 6.0239e-08 | 2.1077e-09 | 74.7217 | -15.673 | 0.019473 | 3.1162e-06 | 6.6968e-09 | 5.9293e-08 | 2.2452e-09 | 74.7587 | -21.0825 | 3.4028e+38 | 3.4028e+38 | -0.016222 | 230.0104 | 230.0101 | 3.4028e+38 | 3.4028e+38 | 230.0013 | 1 | -0.015789 | 230.0097 | 230.01 | 50 | -90.0039 | 0.0179 | 3.4155e-06 | 6.7268e-09 | 6.0716e-08 | 3.4464e-09 | 74.7427 | -19.3815 | 0.019406 | 3.1564e-06 | 6.3847e-09 | 6.2286e-08 | 4.4025e-09 | 74.7273 | -34.2793 | 3.4028e+38 | 3.4028e+38 | -0.015789 | 230.0097 | 230.01 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 1 | -0.01571 | 230.01 | 230.0102 | 49.9999 | -90.0039 | 0.015786 | 3.5159e-06 | 5.4538e-09 | 5.5531e-08 | 1.0671e-09 | 74.5827 | -10.6501 | 0.016885 | 3.0353e-06 | 3.4017e-09 | 5.1225e-08 | -1.0139e-10 | 74.6957 | 2.1892 | 3.4028e+38 | 3.4028e+38 | -0.01571 | 230.01 | 230.0102 | 3.4028e+38 | 3.4028e+38 | 230.002 | 1 | -0.015897 | 230.01 | 230.01 | 50 | -90.0041 | 0.015521 | 3.377e-06 | 3.3703e-09 | 5.2917e-08 | 3.9997e-09 | 74.7176 | -53.8394 | 0.016671 | 3.1684e-06 | 5.0364e-09 | 5.0004e-08 | 3.0034e-09 | 74.6709 | -37.9389 | 3.4028e+38 | 3.4028e+38 | -0.015897 | 230.01 | 230.01 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | -0.015622 | 230.0105 | 230.0101 | 50 | -90.004 | 0.014213 | 3.5045e-06 | 3.6515e-09 | 4.9e-08 | 7.8757e-10 | 74.7198 | -2.61 | 0.015273 | 3.1149e-06 | 1.1204e-09 | 4.769e-08 | -3.0558e-10 | 74.7506 | 1.12 | 3.4028e+38 | 3.4028e+38 | -0.015622 | 230.0105 | 230.0101 | 3.4028e+38 | 3.4028e+38 | 123 | 230.002 | 1 | -0.015603 | 230.0106 | 230.0107 | 49.9999 | -90.0039 | -6.7131e-05 | 3.4028e+38 | 0 | 230.0055 | -0.020418 | 230.0151 | 230.0116 | -0.00035024 | 0.015841 | 3.4485e-06 | 4.8073e-09 | 5.4595e-08 | -2.7668e-10 | 74.7137 | -0.79214 | 0.086272 | 3.4028e+38 | 40.8562 | 0.020252 | 0 | 0 | 0 | 2.4281e-05 | 0.018122 | 3.0211e-06 | 4.2752e-09 | 5.4698e-08 | 6.8734e-10 | 74.7448 | -3.8758 | 0.069061 | 3.4028e+38 | -68.687 | 229.1719 | -3.4993e-10 | 0.00070272 | 2.083e-10 | -1.8782e-05 | 3.4028e+38 | 3.4028e+38 | -0.015206 | 230.0103 | 230.0103 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -0.025224 | -0.025984 | 230.0077 | -0.00091539 | 230.0016 | 1 | -0.015848 | 230.0101 | 230.0101 | 50 | -90.004 | [17 empty] | 21.0522 | 78.0376 | -317378969.0866 | 97.5877 | 97.3827 |  | -0.0039859
140 | 230.0013 | 1 | -0.015651 | 230.0106 | 230.0106 | 50 | -90.0039 | 0.014589 | 3.4294e-06 | 3.4991e-09 | 4.8885e-08 | 1.4959e-09 | 74.7123 | -26.497 | 0.015654 | 3.0967e-06 | 1.9931e-09 | 4.7251e-08 | 9.6811e-10 | 74.7361 | -25.9073 | 3.4028e+38 | 3.4028e+38 | -0.015651 | 230.0106 | 230.0106 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 1 | -0.01592 | 230.0096 | 230.0094 | 49.9999 | -90.0039 | 0.0171 | 3.4316e-06 | 6.118e-09 | 5.9475e-08 | 8.4246e-10 | 74.7212 | -15.4421 | 0.019017 | 3.1609e-06 | 5.4781e-09 | 6.0198e-08 | 2.0992e-09 | 74.7491 | -20.5095 | 3.4028e+38 | 3.4028e+38 | -0.01592 | 230.0096 | 230.0094 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 1 | -0.01645 | 230.0101 | 230.0106 | 49.9999 | -90.0041 | 0.014451 | 3.4261e-06 | 2.2275e-09 | 4.9504e-08 | 1.9799e-09 | 74.6798 | -38.8513 | 0.015741 | 3.0729e-06 | 1.8917e-09 | 4.8077e-08 | 4.7647e-10 | 74.7736 | -12.788 | 3.4028e+38 | 3.4028e+38 | -0.01645 | 230.0101 | 230.0106 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 1 | -0.015563 | 230.0103 | 230.0105 | 49.9999 | -90.0039 | 0.015713 | 3.4611e-06 | 6.5165e-09 | 5.5452e-08 | 1.5999e-09 | 74.6948 | -14.715 | 0.016979 | 3.1122e-06 | 2.0518e-09 | 5.1684e-08 | 3.6981e-10 | 74.791 | -10.2168 | 3.4028e+38 | 3.4028e+38 | -0.015563 | 230.0103 | 230.0105 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 1 | -0.016075 | 230.0099 | 230.0097 | 50 | -90.004 | 0.015652 | 3.4357e-06 | 5.1475e-09 | 5.3446e-08 | 4.6537e-09 | 74.7087 | -39.2721 | 0.016882 | 3.1421e-06 | 3.4065e-09 | 5.2912e-08 | 3.2047e-09 | 74.7582 | -35.0059 | 3.4028e+38 | 3.4028e+38 | -0.016075 | 230.0099 | 230.0097 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0016 | 1 | -0.01616 | 230.0095 | 230.0094 | 50 | -90.0042 | -7.3211e-05 | 3.4028e+38 | 0 | 230.0003 | -0.022743 | 230.0089 | 230.0061 | -0.00032213 | 0.015058 | 3.443e-06 | 2.8801e-09 | 5.187e-08 | 3.5574e-09 | 74.776 | -36.0383 | 0.075077 | 3.4028e+38 | 11.0085 | 0.020468 | 0 | 0 | 0 | 2.6907e-05 | 0.016039 | 3.0899e-06 | 4.0987e-09 | 4.9546e-08 | 1.6683e-09 | 74.7117 | -28.0144 | 0.061252 | 3.4028e+38 | -23.0822 | 229.3295 | -7.9843e-10 | 0.00071421 | 6.6064e-10 | 1.6105e-07 | 3.4028e+38 | 3.4028e+38 | -0.015913 | 230.0106 | 230.0109 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -0.020833 | -0.020833 | 230.0083 | -5.3593e-05 | 230.0015 | 1 | -0.015932 | 230.0101 | 230.0102 | 49.9999 | -90.004 | [17 empty] | 21.4705 | 87.1303 | -27347127.3873 | 108.2327 | 108.4257 |  | -0.0041073
141 | 230.0017 | 1 | 199.1869 | 230.0102 | 115.0191 | 49.9999 | -30.0038 | 0.017526 | 3.4776e-06 | 3.9647e-09 | 6.0749e-08 | -4.6476e-09 | 74.8152 | 72.9532 | 0.019224 | 2.9718e-06 | 4.0011e-09 | 5.7061e-08 | -1.2824e-09 | 74.7073 | 23 | 3.4028e+38 | 3.4028e+38 | 199.1869 | 230.0102 | 115.0191 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 1 | 199.1865 | 230.0099 | 115.0183 | 49.9999 | -30.0038 | 0.015574 | 3.4078e-06 | 3.8984e-09 | 5.3844e-08 | -4.2248e-10 | 74.6221 | 5.3341 | 0.017331 | 2.9964e-06 | 5.0113e-09 | 5.0829e-08 | 1.5823e-09 | 74.7585 | -9.9845 | 3.4028e+38 | 3.4028e+38 | 199.1865 | 230.0099 | 115.0183 | 3.4028e+38 | 3.4028e+38 | 230.001 | 1 | 199.1869 | 230.0102 | 115.0186 | 49.9999 | -30.0038 | 0.017141 | 3.4326e-06 | 1.8464e-09 | 5.8528e-08 | -5.4683e-09 | 74.7495 | 80.1944 | 0.018842 | 3.1132e-06 | 2.7142e-09 | 5.8497e-08 | -1.7549e-09 | 74.8188 | 43.7348 | 3.4028e+38 | 3.4028e+38 | 199.1869 | 230.0102 | 115.0186 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 1 | 199.1879 | 230.011 | 115.0187 | 49.9999 | -30.0038 | 0.016689 | 3.4808e-06 | 6.2295e-10 | 5.8092e-08 | -5.3479e-09 | 74.7481 | 90.4946 | 0.018442 | 3.0799e-06 | 1.1653e-09 | 5.8003e-08 | -2.1378e-09 | 74.8144 | 54.1686 | 3.4028e+38 | 3.4028e+38 | 199.1879 | 230.011 | 115.0187 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 1 | 199.1872 | 230.0103 | 115.0188 | 49.9999 | -30.0038 | 0.017529 | 3.4686e-06 | 2.2581e-09 | 5.9519e-08 | -6.0102e-09 | 74.697 | 73.0691 | 0.019357 | 3.1565e-06 | 1.694e-09 | 6.0774e-08 | -8.6801e-10 | 74.7761 | 13.9204 | 3.4028e+38 | 3.4028e+38 | 199.1872 | 230.0103 | 115.0188 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0013 | 1 | 199.1872 | 230.0101 | 115.0185 | 50 | -30.0038 | 0.86599 | 3.4028e+38 | 0 | 230.0062 | 199.1878 | 230.015 | 115.0266 | 0.00010967 | 0.01616 | 3.4661e-06 | 2.1861e-10 | 5.5716e-08 | -4.3999e-09 | 74.7768 | 89.3513 | 0.00098103 | 3.4028e+38 | 12.3737 | 0.020121 | 0 | 0 | 0 | -6.2648e-05 | 0.016477 | 3.0714e-06 | 1.2027e-09 | 5.0594e-08 | -1.9064e-09 | 74.8022 | 96.7716 | -0.0044379 | 3.4028e+38 | 54.4789 | 229.0411 | -4.7356e-10 | 0.00069422 | 4.3231e-10 | 3.5963e-05 | 3.4028e+38 | 3.4028e+38 | 199.1874 | 230.0108 | 115.0195 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 199.1861 | 199.1861 | 115.0216 | -0.0002408 | 230.0016 | 1 | 199.1871 | 230.0103 | 115.0187 | 49.9999 | -30.0038 | [17 empty] | 21.7254 | 7.429 | -12.2163 | 29.1062 | 154.8848 |  | -0.0040713
142 | 230.0018 | 1 | 199.1869 | 230.01 | 115.0181 | 50 | -30.0038 | 0.015494 | 3.4251e-06 | 3.3187e-09 | 5.1212e-08 | -2.4831e-10 | 74.7392 | -4.9352 | 0.016597 | 2.996e-06 | 5.0351e-09 | 5.198e-08 | 5.7695e-10 | 74.6849 | -8.8374 | 3.4028e+38 | 3.4028e+38 | 199.1869 | 230.01 | 115.0181 | 3.4028e+38 | 3.4028e+38 | 230.0012 | 1 | 199.1867 | 230.0095 | 115.0179 | 49.9999 | -30.0037 | 0.017877 | 3.4181e-06 | 3.4289e-09 | 6.083e-08 | -3.1013e-09 | 74.8164 | 42.2954 | 0.019701 | 3.1447e-06 | 5.2149e-09 | 6.0718e-08 | -4.2336e-10 | 74.7806 | 5.8287 | 3.4028e+38 | 3.4028e+38 | 199.1867 | 230.0095 | 115.0178 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 1 | 199.1864 | 230.0098 | 115.0193 | 49.9999 | -30.0041 | 0.015136 | 3.4558e-06 | 4.2556e-09 | 5.2357e-08 | -3.9281e-10 | 74.7689 | 5.9377 | 0.016153 | 3.1328e-06 | 5.2485e-09 | 5.2145e-08 | 9.5616e-10 | 74.8867 | -13.9744 | 3.4028e+38 | 3.4028e+38 | 199.1864 | 230.0098 | 115.0193 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 1 | 199.1865 | 230.0094 | 115.0182 | 50 | -30.0038 | 0.015863 | 3.4445e-06 | 3.5224e-09 | 5.3241e-08 | -7.61e-10 | 74.7375 | 12.1911 | 0.016762 | 3.0568e-06 | 4.1425e-09 | 5.2566e-08 | 1.8693e-09 | 74.7633 | -26.5153 | 3.4028e+38 | 3.4028e+38 | 199.1865 | 230.0094 | 115.0182 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 1 | 199.1872 | 230.0105 | 115.0194 | 50 | -30.0036 | 0.016721 | 3.4158e-06 | 1.2769e-09 | 5.6426e-08 | -4.7014e-09 | 74.711 | 75.4937 | 0.018331 | 3.1481e-06 | 1.8188e-09 | 5.4309e-08 | -1.7053e-09 | 74.7168 | 43.155 | 3.4028e+38 | 3.4028e+38 | 199.1872 | 230.0105 | 115.0194 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0016 | 1 | 199.1877 | 230.0104 | 115.0183 | 49.9999 | -30.0037 | 0.86599 | 3.4028e+38 | 0 | 230.0054 | 199.1883 | 230.0144 | 115.024 | -8.9606e-05 | 0.014835 | 3.4598e-06 | 1.2569e-09 | 5.129e-08 | -1.7495e-09 | 74.7728 | 83.99 | 0.0036176 | 3.4028e+38 | 20.3658 | 0.019876 | 0 | 0 | 0 | -5.9882e-05 | 0.015513 | 3.1322e-06 | 1.2725e-09 | 4.8528e-08 | -2.8633e-10 | 74.8799 | 24.2449 | 0.026223 | 3.4028e+38 | 20.5756 | 229.7509 | 4.5155e-10 | 0.00073549 | -2.2852e-10 | 3.8976e-05 | 3.4028e+38 | 3.4028e+38 | 199.1869 | 230.0101 | 115.019 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 199.1817 | 199.1814 | 115.0206 | -0.00053688 | 230.0015 | 1 | 199.1867 | 230.0098 | 115.0186 | 49.9999 | -30.0038 | [17 empty] | 20.5178 | 10.5197 | -14.034 | 29.9134 | 165.5956 |  | -0.0043741
143 | 230.0018 | 2.0001 | -276.0438 | 460.0286 | 368.0077 | 49.9998 | -126.8734 | 0.016228 | 3.3517e-06 | 2.051e-09 | 5.3445e-08 | 3.8703e-09 | 74.7506 | -56.6902 | 0.017305 | 3.0253e-06 | -8.0592e-10 | 5.2298e-08 | 1.0612e-09 | 74.7451 | -138.8039 | 3.4028e+38 | 3.4028e+38 | -276.0438 | 460.0286 | 368.0077 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 2.0001 | -276.0477 | 460.0302 | 368.0064 | 50 | -126.8738 | 0.014189 | 3.372e-06 | -7.8259e-10 | 4.8811e-08 | 1.1242e-09 | 74.6909 | -124.8442 | 0.015301 | 3.0228e-06 | 4.2641e-10 | 4.7784e-08 | -5.7508e-10 | 74.7644 | 1.598 | 3.4028e+38 | 3.4028e+38 | -276.0477 | 460.0302 | 368.0063 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 2.0001 | -276.0489 | 460.0279 | 368.0025 | 50 | -126.8738 | 0.016447 | 3.393e-06 | -2.3601e-09 | 5.5541e-08 | 1.9182e-09 | 74.7675 | -138.5893 | 0.018073 | 3.0985e-06 | 3.9883e-10 | 5.4175e-08 | 9.1684e-10 | 74.8209 | 161.6425 | 3.4028e+38 | 3.4028e+38 | -276.0489 | 460.0279 | 368.0025 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 2.0001 | -276.0483 | 460.031 | 368.0075 | 49.9999 | -126.8739 | 0.014591 | 3.3785e-06 | 9.5193e-10 | 4.9473e-08 | 2.0344e-09 | 74.8074 | -82.4979 | 0.015493 | 3.0505e-06 | -9.3968e-10 | 4.5734e-08 | 9.089e-11 | 74.7018 | 178.0506 | 3.4028e+38 | 3.4028e+38 | -276.0483 | 460.031 | 368.0075 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 2.0001 | -276.0425 | 460.024 | 368.0027 | 50 | -126.8743 | 0.01771 | 3.3815e-06 | -5.5043e-10 | 6.0139e-08 | 3.1943e-09 | 74.7857 | -99.7771 | 0.019757 | 3.0861e-06 | -7.6916e-10 | 5.9859e-08 | 1.7356e-09 | 74.7214 | -101.0942 | 3.4028e+38 | 3.4028e+38 | -276.0426 | 460.024 | 368.0027 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0016 | 2.0001 | -276.0457 | 460.0282 | 368.0051 | 50 | -126.8741 | -0.60006 | 3.4028e+38 | 0 | 230.0062 | -276.0627 | 460.0344 | 367.9919 | -0.0013856 | 0.016879 | 3.3626e-06 | 1.9826e-09 | 5.6751e-08 | 4.067e-09 | 74.688 | -55.1886 | 0.044975 | 3.4028e+38 | -36.8331 | 0.020793 | 0 | 0 | 0 | 5.5624e-05 | 0.017827 | 3.0826e-06 | -3.8726e-10 | 5.4842e-08 | 1.0098e-09 | 74.8008 | -119.1253 | -0.010677 | 3.4028e+38 | 16.6681 | 229.1383 | -5.5767e-11 | 0.00071492 | 5.7955e-10 | -1.0431e-05 | 3.4028e+38 | 3.4028e+38 | -276.046 | 460.0283 | 368.0057 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -276.058 | -276.058 | 367.9995 | -0.00075828 | 230.0018 | 2.0001 | -276.0462 | 460.0284 | 368.0053 | 49.9999 | -126.8738 | [17 empty] | 23.0215 | 70.4577 | 119.752 | 93.1079 | 90.5487 |  | -0.00050811
144 | 230.0021 | 2.0001 | -276.0463 | 460.0309 | 368.0082 | 49.9999 | -126.8737 | 0.015186 | 3.3553e-06 | -1.9326e-09 | 5.0939e-08 | 2.0349e-09 | 74.7503 | -130.6314 | 0.016455 | 3.0628e-06 | 4.2971e-10 | 5.0435e-08 | -9.4239e-11 | 74.7357 | 56.9306 | 3.4028e+38 | 3.4028e+38 | -276.0464 | 460.0309 | 368.0082 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 2.0001 | -276.0438 | 460.0281 | 368.0071 | 49.9999 | -126.8735 | 0.014411 | 3.405e-06 | 2.4911e-10 | 4.8505e-08 | 1.3384e-09 | 74.6868 | -96.8198 | 0.015455 | 3.0723e-06 | -6.5259e-10 | 4.8105e-08 | -4.498e-10 | 74.812 | 134.9275 | 3.4028e+38 | 3.4028e+38 | -276.0439 | 460.0281 | 368.007 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 2.0001 | -276.0449 | 460.0281 | 368.0053 | 49.9999 | -126.8741 | 0.014413 | 3.4315e-06 | 3.8132e-10 | 4.9412e-08 | 2.4694e-09 | 74.8028 | -76.1822 | 0.015789 | 3.1418e-06 | 1.71e-10 | 4.7507e-08 | 1.8593e-10 | 74.7476 | -47.3952 | 3.4028e+38 | 3.4028e+38 | -276.0449 | 460.0281 | 368.0053 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 2.0001 | -276.0461 | 460.0277 | 368.0049 | 50 | -126.8744 | 0.017896 | 3.3691e-06 | -3.5863e-10 | 6.0551e-08 | 3.0057e-09 | 74.7421 | -97.8034 | 0.019564 | 3.0995e-06 | -9.9917e-10 | 6.0705e-08 | 1.3325e-09 | 74.6725 | -121.4217 | 3.4028e+38 | 3.4028e+38 | -276.0461 | 460.0277 | 368.0049 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | -276.0453 | 460.03 | 368.0082 | 49.9999 | -126.874 | 0.014955 | 3.3747e-06 | -2.3906e-09 | 4.9954e-08 | 1.1253e-09 | 74.8086 | -140.6619 | 0.016143 | 3.1354e-06 | 4.2222e-10 | 5.0952e-08 | -5.1972e-10 | 74.7725 | -27.3982 | 3.4028e+38 | 3.4028e+38 | -276.0453 | 460.03 | 368.0082 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 2.0001 | -276.0455 | 460.0288 | 368.0061 | 49.9999 | -126.874 | -0.60006 | 3.4028e+38 | 0 | 230.0004 | -276.0567 | 460.027 | 367.9923 | -3.3783e-05 | 0.015625 | 3.3944e-06 | -1.8605e-09 | 5.3414e-08 | 2.2957e-09 | 74.6043 | -114.9843 | -0.021759 | 3.4028e+38 | 39.5933 | 0.02034 | 0 | 0 | 0 | 5.8731e-05 | 0.017355 | 3.1422e-06 | -3.3187e-10 | 5.1809e-08 | 5.9032e-10 | 74.7832 | -119.3445 | -0.0038619 | 3.4028e+38 | 60.8941 | 229.1646 | -4.3541e-10 | 0.00068915 | -5.6151e-10 | -5.1009e-06 | 3.4028e+38 | 3.4028e+38 | -276.0488 | 460.0303 | 368.0053 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | -276.0564 | -276.0564 | 367.9911 | -0.00010614 | 230.0021 | 2.0001 | -276.0453 | 460.029 | 368.0068 | 49.9999 | -126.8739 | [17 empty] | 23.8612 | 76.5667 | 112.7043 | 97.5419 | 101.2603 |  | -0.00035652
145 | 230.002 | 5.0002 | 995.9536 | 1150.0626 | 575.0878 | 50 | -30.003 | 0.01698 | 3.5211e-06 | 9.1005e-09 | 5.9868e-08 | 7.4395e-09 | 74.6767 | -36.5131 | 0.018443 | 3.1196e-06 | 3.1432e-09 | 6.0766e-08 | 8.6199e-09 | 74.7188 | -63.7662 | 3.4028e+38 | 3.4028e+38 | 995.9536 | 1150.0626 | 575.0878 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 5.0003 | 995.9588 | 1150.0659 | 575.0892 | 50 | -30.003 | 0.016332 | 3.5455e-06 | 8.9158e-09 | 5.6724e-08 | -2.5278e-10 | 74.6879 | 1.6599 | 0.018045 | 3.184e-06 | 3.4431e-09 | 5.5445e-08 | 4.9297e-09 | 74.7466 | -38.6242 | 3.4028e+38 | 3.4028e+38 | 995.9588 | 1150.0659 | 575.0892 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 5.0002 | 995.9579 | 1150.0662 | 575.0842 | 50 | -30.0029 | 0.01435 | 3.5798e-06 | 4.6912e-09 | 5.1349e-08 | 3.5943e-10 | 74.7062 | -3.2041 | 0.015708 | 3.1852e-06 | 3.0293e-09 | 4.7667e-08 | 2.5652e-09 | 74.7718 | -44.3709 | 3.4028e+38 | 3.4028e+38 | 995.9579 | 1150.0662 | 575.0842 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 5.0003 | 995.9563 | 1150.0651 | 575.0822 | 50 | -30.0028 | 0.017896 | 3.5405e-06 | 1.1208e-08 | 6.3402e-08 | 4.57e-09 | 74.7512 | -20.017 | 0.019359 | 3.1668e-06 | 6.5946e-09 | 6.1411e-08 | 9.6174e-09 | 74.7722 | -59.9844 | 3.4028e+38 | 3.4028e+38 | 995.9563 | 1150.0651 | 575.0822 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 5.0002 | 995.9573 | 1150.0648 | 575.087 | 49.9999 | -30.0031 | 0.016311 | 3.5703e-06 | 9.1041e-09 | 5.7594e-08 | 6.2905e-09 | 74.7159 | -36.6171 | 0.018107 | 3.1201e-06 | 3.54e-09 | 5.5933e-08 | 8.2042e-09 | 74.7982 | -71.8186 | 3.4028e+38 | 3.4028e+38 | 995.9573 | 1150.0648 | 575.087 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0022 | 5.0002 | 995.9553 | 1150.0631 | 575.0857 | 50 | -30.0029 | 0.866 | 3.4028e+38 | 0 | 230.0003 | 995.931 | 1150.0568 | 575.0983 | -0.0030547 | 0.016114 | 3.477e-06 | 7.5253e-09 | 5.6502e-08 | 6.122e-09 | 74.8095 | -39.129 | 0.13931 | 3.4028e+38 | 48.6454 | 0.020446 | 0 | 0 | 0 | 2.2549e-05 | 0.017192 | 3.2131e-06 | 2.0696e-09 | 5.5211e-08 | 6.9153e-09 | 74.667 | -75.1219 | 0.034209 | 3.4028e+38 | 26.5885 | 229.4126 | 1.3069e-09 | 0.00073198 | 8.1013e-10 | -8.4716e-05 | 3.4028e+38 | 3.4028e+38 | 995.9548 | 1150.0623 | 575.0832 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 995.9324 | 995.9324 | 575.0944 | -0.0028027 | 230.0019 | 5.0002 | 995.9568 | 1150.0649 | 575.0861 | 50 | -30.0029 | [17 empty] | 21.9872 | 0.2593 | -9.4204 | 22.0872 | 123.5414 |  | -0.0032205
146 | 230.0024 | 5.0003 | 995.9598 | 1150.0693 | 575.0908 | 49.9999 | -30.0032 | 0.014246 | 3.5853e-06 | 5.4734e-09 | 5.1844e-08 | 1.64e-09 | 74.6752 | -17.3768 | 0.015251 | 3.1835e-06 | 3.4103e-09 | 4.8313e-08 | 1.9454e-09 | 74.7678 | -28.6215 | 3.4028e+38 | 3.4028e+38 | 995.9598 | 1150.0693 | 575.0908 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 5.0002 | 995.9519 | 1150.0608 | 575.0875 | 49.9999 | -30.0031 | 0.015102 | 3.5623e-06 | 4.9755e-09 | 5.4065e-08 | 4.847e-09 | 74.6004 | -39.5066 | 0.016178 | 3.1006e-06 | 9.2749e-10 | 5.2366e-08 | 4.873e-09 | 74.8219 | -79.2237 | 3.4028e+38 | 3.4028e+38 | 995.9519 | 1150.0608 | 575.0875 | 3.4028e+38 | 3.4028e+38 | 230.002 | 5.0003 | 995.9623 | 1150.0712 | 575.088 | 49.9999 | -30.0032 | 0.01655 | 3.6082e-06 | 9.2534e-09 | 5.9155e-08 | 1.2886e-09 | 74.6566 | -8.169 | 0.017957 | 3.2354e-06 | 7.3734e-09 | 5.7966e-08 | 5.6575e-09 | 74.7266 | -44.6603 | 3.4028e+38 | 3.4028e+38 | 995.9623 | 1150.0712 | 575.088 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 5.0002 | 995.9564 | 1150.0652 | 575.0842 | 49.9999 | -30.0031 | 0.014684 | 3.5057e-06 | 5.9979e-09 | 5.2628e-08 | 5.1133e-09 | 74.7814 | -41.0612 | 0.016068 | 3.1689e-06 | 2.4168e-09 | 5.0864e-08 | 3.8094e-09 | 74.7669 | -80.9059 | 3.4028e+38 | 3.4028e+38 | 995.9564 | 1150.0652 | 575.0842 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 5.0002 | 995.9573 | 1150.0638 | 575.0843 | 49.9999 | -30.003 | 0.014288 | 3.5457e-06 | 4.444e-09 | 5.1876e-08 | 2.6774e-09 | 74.6896 | -30.1073 | 0.015591 | 3.1153e-06 | 9.7239e-10 | 4.8602e-08 | 2.7273e-09 | 74.7629 | -67.788 | 3.4028e+38 | 3.4028e+38 | 995.9573 | 1150.0638 | 575.0843 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0025 | 5.0002 | 995.9578 | 1150.0613 | 575.086 | 49.9999 | -30.0031 | 0.866 | 3.4028e+38 | 0 | 230.0028 | 995.9532 | 1150.0731 | 575.1005 | -0.00073972 | 0.01468 | 3.5979e-06 | 5.1576e-09 | 5.0943e-08 | 2.4096e-09 | 74.6989 | -25.0418 | 0.096081 | 3.4028e+38 | 9.1693 | 0.020017 | 0 | 0 | 0 | 2.9063e-05 | 0.015416 | 3.1799e-06 | 1.7746e-09 | 4.9149e-08 | 3.4396e-09 | 74.7658 | -62.6436 | 0.036208 | 3.4028e+38 | -59.6071 | 229.5405 | -2.6522e-10 | 0.00072082 | 3.0316e-10 | -7.8669e-05 | 3.4028e+38 | 3.4028e+38 | 995.9551 | 1150.0653 | 575.0842 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 995.9437 | 995.9437 | 575.0992 | -0.0015967 | 230.002 | 5.0002 | 995.9575 | 1150.0661 | 575.0869 | 49.9999 | -30.0031 | [17 empty] | 23.0596 | 1.8438 | -7.4681 | 24.2144 | 125.9839 |  | -0.0033905
147 | 230.0018 | 2.0001 | 460.0275 | 460.0294 | 0.028195 | 49.9999 | -0.0035116 | 0.014832 | 3.4227e-06 | 2.5352e-09 | 4.9547e-08 | 6.8835e-10 | 74.7182 | -15.8279 | 0.0159 | 3.1422e-06 | 2.4402e-09 | 4.9883e-08 | 2.2014e-09 | 74.7174 | -42.4027 | 3.4028e+38 | 3.4028e+38 | 460.0275 | 460.0294 | 0.028195 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | 460.0313 | 460.031 | 0.027659 | 50 | -0.0037192 | 0.017486 | 3.3837e-06 | 2.724e-09 | 5.9212e-08 | -9.0908e-10 | 74.7548 | 18.4553 | 0.019335 | 3.0877e-06 | 4.1565e-09 | 6.0882e-08 | 5.6156e-09 | 74.7478 | -48.1048 | 3.4028e+38 | 3.4028e+38 | 460.0313 | 460.031 | 0.027659 | 3.4028e+38 | 3.4028e+38 | 230.002 | 2.0001 | 460.0259 | 460.0257 | 0.025219 | 49.9999 | -0.0031432 | 0.016681 | 3.3876e-06 | 4.8676e-09 | 5.8041e-08 | -8.1744e-11 | 74.7762 | 1.7021 | 0.018558 | 3.1491e-06 | 3.7689e-09 | 5.6612e-08 | 4.8131e-09 | 74.6182 | -51.9373 | 3.4028e+38 | 3.4028e+38 | 460.0259 | 460.0257 | 0.025219 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 2.0001 | 460.0294 | 460.0292 | 0.02453 | 49.9999 | -0.0035721 | 0.016743 | 3.3734e-06 | 2.9373e-09 | 5.6093e-08 | 5.6829e-10 | 74.8437 | -0.82708 | 0.018269 | 3.0731e-06 | 4.6805e-09 | 5.7439e-08 | 4.6938e-09 | 74.7168 | -43.6043 | 3.4028e+38 | 3.4028e+38 | 460.0294 | 460.0292 | 0.02453 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 2.0001 | 460.0323 | 460.031 | 0.029754 | 49.9999 | -0.0037057 | 0.017312 | 3.3313e-06 | 3.3116e-09 | 5.7698e-08 | -2.8063e-09 | 74.7655 | 39.0804 | 0.018925 | 3.0497e-06 | 5.3618e-09 | 5.8196e-08 | 4.037e-09 | 74.7695 | -37.6211 | 3.4028e+38 | 3.4028e+38 | 460.0323 | 460.031 | 0.029754 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0017 | 2.0001 | 460.0295 | 460.0292 | 0.026568 | 50 | -0.003309 | 1 | 3.4028e+38 | 0 | 230.0059 | 460.0339 | 460.0363 | 0.047362 | 0.00080496 | 0.017444 | 3.4438e-06 | 2.586e-09 | 5.9859e-08 | -1.6225e-09 | 74.7524 | 28.1639 | 0.050629 | 3.4028e+38 | -40.6082 | 0.021045 | 0 | 0 | 0 | -2.9482e-05 | 0.019266 | 3.0859e-06 | 3.5812e-09 | 5.9456e-08 | 3.276e-09 | 74.7744 | -33.4119 | 0.084227 | 3.4028e+38 | -26.1049 | 229.1033 | 8.1229e-10 | 0.00072559 | 1.7685e-09 | -3.845e-05 | 3.4028e+38 | 3.4028e+38 | 460.0318 | 460.0315 | 0.030055 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 460.0385 | 460.0385 | 0.045844 | 0.00029668 | 230.0019 | 2.0001 | 460.0293 | 460.0293 | 0.027071 | 49.9999 | -0.0035304 | [17 empty] | 23.3769 | 7.0647 | 30.2859 | 30.1981 | 679122.0332 |  | -0.0015223
148 | 230.0021 | 2.0001 | 460.032 | 460.0317 | 0.030452 | 50 | -0.0037927 | 0.014502 | 3.4025e-06 | 2.7943e-09 | 4.7942e-08 | 6.1119e-10 | 74.7212 | -12.3377 | 0.015398 | 3.1207e-06 | 2.4294e-09 | 4.7968e-08 | 1.5501e-09 | 74.7661 | -38.1087 | 3.4028e+38 | 3.4028e+38 | 460.032 | 460.0317 | 0.030452 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 2.0001 | 460.025 | 460.0272 | 0.024967 | 50 | -0.0031096 | 0.017347 | 3.4108e-06 | 3.7746e-09 | 5.9494e-08 | 9.1965e-11 | 74.6844 | 2.4132 | 0.019044 | 3.035e-06 | 5.6834e-09 | 6.0145e-08 | 5.3723e-09 | 74.7605 | -43.3882 | 3.4028e+38 | 3.4028e+38 | 460.025 | 460.0272 | 0.024967 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 2.0001 | 460.0266 | 460.0264 | 0.026412 | 49.9999 | -0.0033376 | 0.017466 | 3.4362e-06 | 4.437e-09 | 6.0005e-08 | 8.9186e-12 | 74.6956 | 0.4784 | 0.019273 | 3.1595e-06 | 4.2191e-09 | 6.1242e-08 | 5.09e-09 | 74.7531 | -50.4497 | 3.4028e+38 | 3.4028e+38 | 460.0266 | 460.0264 | 0.026412 | 3.4028e+38 | 3.4028e+38 | 230.0027 | 2.0001 | 460.0296 | 460.0293 | 0.027802 | 49.9999 | -0.0032148 | 0.014929 | 3.3769e-06 | 1.711e-09 | 5.0435e-08 | 1.9719e-10 | 74.7868 | -5.6803 | 0.016193 | 3.0402e-06 | 2.1714e-09 | 4.9214e-08 | 2.2889e-09 | 74.7821 | -56.0781 | 3.4028e+38 | 3.4028e+38 | 460.0296 | 460.0293 | 0.027802 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 2.0001 | 460.0285 | 460.0281 | 0.027396 | 49.9999 | -0.0034122 | 0.015239 | 3.418e-06 | 2.3309e-09 | 5.2063e-08 | 1.3554e-09 | 74.8508 | -11.349 | 0.016481 | 3.1417e-06 | 2.5611e-09 | 5.1117e-08 | 2.7892e-09 | 74.7617 | -54.6419 | 3.4028e+38 | 3.4028e+38 | 460.0285 | 460.0281 | 0.027396 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0021 | 2.0001 | 460.0304 | 460.0302 | 0.027461 | 49.9999 | -0.0035031 | 1 | 3.4028e+38 | 0 | 230.0012 | 460.0235 | 460.0232 | 0.036343 | -0.00069267 | 0.014687 | 3.4333e-06 | 1.5198e-09 | 5.0418e-08 | 6.259e-10 | 74.7523 | -14.9599 | 0.054464 | 3.4028e+38 | -40.7372 | 0.020074 | 0 | 0 | 0 | -1.4013e-05 | 0.015484 | 2.9927e-06 | 1.5366e-09 | 4.9005e-08 | 9.7234e-10 | 74.7369 | -32.3244 | 0.028648 | 3.4028e+38 | 61.6699 | 229.0932 | 6.3843e-10 | 0.00071942 | 7.3257e-10 | -1.7776e-05 | 3.4028e+38 | 3.4028e+38 | 460.0332 | 460.033 | 0.029551 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 460.032 | 460.032 | 0.032967 | -0.00010295 | 230.002 | 2.0001 | 460.0284 | 460.0285 | 0.027406 | 49.9999 | -0.0033734 | [17 empty] | 23.4024 | 12.7104 | 35.9926 | 36.4222 | 429189.1046 |  | -0.00098499
149 | 230.0021 | 2.0001 | 229.9942 | 460.0345 | 398.416 | 50 | -60.0037 | 0.014727 | 3.4099e-06 | 1.7708e-09 | 4.98e-08 | 1.4753e-09 | 74.7143 | -57.5008 | 0.015873 | 3.1466e-06 | 1.3995e-09 | 4.5815e-08 | 4.9523e-10 | 74.7808 | -51.8632 | 3.4028e+38 | 3.4028e+38 | 229.9942 | 460.0345 | 398.416 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 2.0001 | 229.9904 | 460.0291 | 398.4115 | 49.9999 | -60.0034 | 0.017696 | 3.4265e-06 | 4.9642e-09 | 6.0943e-08 | 2.6482e-09 | 74.8201 | -28.3861 | 0.019701 | 3.0946e-06 | 7.1884e-10 | 6.1058e-08 | 4.9726e-09 | 74.7631 | -67.0953 | 3.4028e+38 | 3.4028e+38 | 229.9904 | 460.0291 | 398.4115 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 2.0001 | 229.9903 | 460.0289 | 398.4112 | 49.9999 | -60.0037 | 0.017066 | 3.3569e-06 | 2.5215e-09 | 5.7184e-08 | 3.8872e-09 | 74.8024 | -54.6678 | 0.018701 | 3.068e-06 | 1.4228e-09 | 5.8454e-08 | 3.5212e-09 | 74.785 | -60.3945 | 3.4028e+38 | 3.4028e+38 | 229.9903 | 460.0289 | 398.4112 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 2.0001 | 229.992 | 460.027 | 398.4108 | 50 | -60.0033 | 0.016805 | 3.3636e-06 | 2.3816e-09 | 5.7541e-08 | 2.9604e-09 | 74.8334 | -51.1841 | 0.018718 | 3.1314e-06 | 9.4853e-10 | 5.5495e-08 | 3.4663e-09 | 74.7133 | -73.2599 | 3.4028e+38 | 3.4028e+38 | 229.992 | 460.027 | 398.4108 | 3.4028e+38 | 3.4028e+38 | 230.0025 | 2.0001 | 229.9904 | 460.026 | 398.4084 | 49.9999 | -60.0033 | 0.017607 | 3.3533e-06 | 5.1332e-09 | 5.9468e-08 | 3.5547e-09 | 74.761 | -46.6694 | 0.019431 | 3.1069e-06 | 3.2334e-09 | 6.0532e-08 | 4.9899e-09 | 74.7583 | -70.2096 | 3.4028e+38 | 3.4028e+38 | 229.9904 | 460.026 | 398.4084 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 2.0001 | 229.9915 | 460.0312 | 398.4144 | 49.9999 | -60.0035 | 0.49995 | 3.4028e+38 | 0 | 230.0014 | 229.9729 | 460.0232 | 398.4056 | -0.001618 | 0.016682 | 3.3854e-06 | 3.4216e-09 | 5.6475e-08 | 3.5334e-09 | 74.7331 | -68.6696 | 0.024687 | 3.4028e+38 | 38.7355 | 0.020529 | 0 | 0 | 0 | 3.5825e-05 | 0.017549 | 3.0183e-06 | 8.6177e-10 | 5.2754e-08 | 2.3448e-09 | 74.7476 | -97.513 | -0.0058619 | 3.4028e+38 | -55.4714 | 229.1848 | -7.6201e-10 | 0.00069704 | 1.9697e-10 | -3.0474e-05 | 3.4028e+38 | 3.4028e+38 | 229.9912 | 460.0296 | 398.4124 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 229.9795 | 229.9795 | 398.4092 | -0.0005373 | 230.0021 | 2.0001 | 229.9915 | 460.0291 | 398.4116 | 49.9999 | -60.0035 | [17 empty] | 23.0201 | 51.1986 | -32.8191 | 73.5982 | 112.4943 |  | -0.0036898
150 | 230.0022 | 2.0001 | 229.988 | 460.0297 | 398.4136 | 50 | -60.0034 | 0.015924 | 3.4031e-06 | 2.5282e-09 | 5.4302e-08 | 1.2218e-09 | 74.7985 | -24.7929 | 0.017069 | 3.1261e-06 | 1.9434e-09 | 5.2849e-08 | 2.1429e-09 | 74.837 | -68.8355 | 3.4028e+38 | 3.4028e+38 | 229.988 | 460.0297 | 398.4136 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 2.0001 | 229.9932 | 460.031 | 398.4116 | 49.9999 | -60.0033 | 0.017715 | 3.4593e-06 | 4.0931e-09 | 6.1261e-08 | 2.7925e-09 | 74.6304 | -31.8733 | 0.019643 | 3.0676e-06 | 1.5103e-09 | 6.18e-08 | 5.5779e-09 | 74.8514 | -70.8665 | 3.4028e+38 | 3.4028e+38 | 229.9932 | 460.031 | 398.4116 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 2.0001 | 229.9914 | 460.0284 | 398.4085 | 50 | -60.0031 | 0.017773 | 3.3489e-06 | 3.8863e-09 | 6.0609e-08 | 3.3676e-09 | 74.7701 | -42.757 | 0.020081 | 3.1229e-06 | 2.2697e-09 | 5.9498e-08 | 4.8543e-09 | 74.7634 | -64.9405 | 3.4028e+38 | 3.4028e+38 | 229.9914 | 460.0284 | 398.4085 | 3.4028e+38 | 3.4028e+38 | 230.0022 | 2.0001 | 229.9897 | 460.0291 | 398.4127 | 49.9999 | -60.0037 | 0.014329 | 3.4494e-06 | 1.7322e-09 | 4.8605e-08 | 1.0384e-09 | 74.8352 | -30.9423 | 0.014951 | 3.1408e-06 | 1.1759e-09 | 4.7151e-08 | 9.0507e-10 | 74.7049 | -47.1964 | 3.4028e+38 | 3.4028e+38 | 229.9897 | 460.0291 | 398.4127 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 2.0001 | 229.9908 | 460.0296 | 398.4146 | 49.9999 | -60.004 | 0.014545 | 3.3448e-06 | 2.9243e-09 | 4.8689e-08 | 1.0223e-09 | 74.7813 | -32.2404 | 0.016192 | 3.022e-06 | 1.3688e-09 | 4.8934e-08 | 7.9698e-10 | 74.6802 | -36.769 | 3.4028e+38 | 3.4028e+38 | 229.9908 | 460.0296 | 398.4146 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 2.0001 | 229.9908 | 460.0287 | 398.4081 | 50 | -60.0029 | 0.49996 | 3.4028e+38 | 0 | 230.0056 | 229.9883 | 460.0374 | 398.4211 | -0.000418 | 0.015082 | 3.4042e-06 | 3.3188e-09 | 5.1002e-08 | 1.2924e-09 | 74.8 | -22.0013 | 0.062716 | 3.4028e+38 | -70.747 | 0.020108 | 0 | 0 | 0 | 4.1302e-05 | 0.016126 | 3.1381e-06 | 7.7685e-10 | 4.9062e-08 | 8.4781e-10 | 74.7456 | -47.501 | 0.027094 | 3.4028e+38 | 10.7766 | 229.62 | 5.1635e-10 | 0.00071542 | -3.1355e-11 | -3.0623e-05 | 3.4028e+38 | 3.4028e+38 | 229.992 | 460.0276 | 398.4087 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 229.9876 | 229.9863 | 398.4214 | -0.0001843 | 230.0022 | 2.0001 | 229.9906 | 460.0296 | 398.4122 | 49.9999 | -60.0035 | [17 empty] | 23.1481 | 37.9429 | -53.2113 | 60.6621 | 101.0576 |  | -0.0037949
151 | 230.0017 | 5.0002 | 1150.0625 | 1150.0624 | 0.059763 | 50 | -0.0029774 | 0.014972 | 3.4829e-06 | 5.7273e-09 | 5.214e-08 | 1.1161e-09 | 74.8113 | -12.9728 | 0.016111 | 3.2342e-06 | 3.1677e-09 | 5.1026e-08 | 4.9811e-09 | 74.7399 | -57.5455 | 3.4028e+38 | 3.4028e+38 | 1150.0625 | 1150.0624 | 0.059763 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 5.0002 | 1150.0636 | 1150.061 | 0.062078 | 49.9999 | -0.0030202 | 0.014868 | 3.6297e-06 | 3.5618e-09 | 5.2189e-08 | -2.5875e-09 | 74.6691 | 32.8952 | 0.016217 | 3.1689e-06 | 3.4396e-09 | 5.1214e-08 | 1.4173e-09 | 74.7661 | -27.5872 | 3.4028e+38 | 3.4028e+38 | 1150.0636 | 1150.061 | 0.062078 | 3.4028e+38 | 3.4028e+38 | 230.0015 | 5.0002 | 1150.0605 | 1150.058 | 0.057013 | 49.9999 | -0.0028404 | 0.016554 | 3.5248e-06 | 1.0525e-08 | 5.9543e-08 | 1.8172e-09 | 74.7327 | -6.1416 | 0.018155 | 3.164e-06 | 6.9975e-09 | 5.7516e-08 | 1.0686e-08 | 74.7118 | -52.0743 | 3.4028e+38 | 3.4028e+38 | 1150.0605 | 1150.058 | 0.057013 | 3.4028e+38 | 3.4028e+38 | 230.0019 | 5.0002 | 1150.0627 | 1150.0616 | 0.0596 | 50 | -0.0029693 | 0.017742 | 3.5442e-06 | 9.961e-09 | 6.2876e-08 | -7.6311e-10 | 74.5741 | 2.8405 | 0.019436 | 3.2604e-06 | 1.0074e-08 | 6.2676e-08 | 1.0428e-08 | 74.7648 | -45.9553 | 3.4028e+38 | 3.4028e+38 | 1150.0627 | 1150.0616 | 0.0596 | 3.4028e+38 | 3.4028e+38 | 230.0017 | 5.0002 | 1150.067 | 1150.0646 | 0.056243 | 49.9999 | -0.0030491 | 0.015281 | 3.5617e-06 | 6.5769e-09 | 5.3125e-08 | 2.9835e-09 | 74.6816 | -24.0617 | 0.016315 | 3.1277e-06 | 4.3723e-09 | 5.1886e-08 | 6.0916e-09 | 74.7116 | -54.3305 | 3.4028e+38 | 3.4028e+38 | 1150.067 | 1150.0646 | 0.056243 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0018 | 5.0002 | 1150.0602 | 1150.0577 | 0.059808 | 50 | -0.0030884 | 1 | 3.4028e+38 | 0 | 230.0009 | 1150.0599 | 1150.0575 | 0.080841 | -0.0015276 | 0.014744 | 3.5754e-06 | 7.2163e-09 | 5.2183e-08 | 1.9286e-09 | 74.6742 | -14.9629 | 0.11717 | 3.4028e+38 | -30.4384 | 0.020108 | 0 | 0 | 0 | -4.8876e-05 | 0.015802 | 3.2262e-06 | 2.0628e-09 | 5.1343e-08 | 4.3854e-09 | 74.6783 | -59.3473 | 0.04423 | 3.4028e+38 | -17.421 | 229.6971 | -5.0403e-10 | 0.00074249 | 4.7024e-10 | -5.6859e-05 | 3.4028e+38 | 3.4028e+38 | 1150.0642 | 1150.0618 | 0.057905 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1150.0575 | 1150.0575 | 0.081647 | -0.0020393 | 230.0018 | 5.0002 | 1150.0633 | 1150.0615 | 0.05894 | 49.9999 | -0.0029713 | [17 empty] | 21.1045 | 0.96402 | 23.9699 | 22.4471 | 2006394.4142 |  | -0.0019946
152 | 230.002 | 5.0002 | 1150.0643 | 1150.0619 | 0.055734 | 49.9999 | -0.0027767 | 0.015023 | 3.4719e-06 | 6.47e-09 | 5.2069e-08 | 1.6514e-09 | 74.6997 | -18.5785 | 0.01631 | 3.1397e-06 | 3.6204e-09 | 5.2405e-08 | 5.149e-09 | 74.7897 | -54.8882 | 3.4028e+38 | 3.4028e+38 | 1150.0643 | 1150.0619 | 0.055734 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 5.0002 | 1150.0645 | 1150.062 | 0.060482 | 50 | -0.0030132 | 0.017598 | 3.529e-06 | 9.6861e-09 | 6.3556e-08 | -1.4372e-09 | 74.7402 | 8.4398 | 0.019544 | 3.1823e-06 | 7.9157e-09 | 6.2243e-08 | 9.7736e-09 | 74.7479 | -49.1484 | 3.4028e+38 | 3.4028e+38 | 1150.0645 | 1150.062 | 0.060482 | 3.4028e+38 | 3.4028e+38 | 230.002 | 5.0002 | 1150.0649 | 1150.0626 | 0.059493 | 49.9999 | -0.0031123 | 0.014433 | 3.5793e-06 | 3.5143e-09 | 5.0292e-08 | -1.4961e-09 | 74.7147 | 21.5057 | 0.015312 | 3.2006e-06 | 3.0805e-09 | 4.8629e-08 | 1.6057e-09 | 74.7839 | -29.0029 | 3.4028e+38 | 3.4028e+38 | 1150.0649 | 1150.0626 | 0.059493 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 5.0002 | 1150.0615 | 1150.0588 | 0.06033 | 49.9999 | -0.0030056 | 0.014811 | 3.5356e-06 | 6.4123e-09 | 5.237e-08 | 2.01e-09 | 74.6777 | -15.1754 | 0.016197 | 3.1302e-06 | 2.6533e-09 | 5.0669e-08 | 4.2288e-09 | 74.6584 | -51.567 | 3.4028e+38 | 3.4028e+38 | 1150.0615 | 1150.0588 | 0.06033 | 3.4028e+38 | 3.4028e+38 | 230.002 | 5.0002 | 1150.0631 | 1150.0605 | 0.055732 | 50 | -0.0029667 | 0.016894 | 3.5531e-06 | 1.0674e-08 | 6.0324e-08 | 2.121e-09 | 74.739 | -11.4482 | 0.018672 | 3.2099e-06 | 7.9837e-09 | 5.9118e-08 | 9.952e-09 | 74.8177 | -53.5141 | 3.4028e+38 | 3.4028e+38 | 1150.0631 | 1150.0605 | 0.055732 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0018 | 5.0002 | 1150.0626 | 1150.0602 | 0.058041 | 49.9999 | -0.0030362 | 1 | 3.4028e+38 | 0 | 230.0022 | 1150.0653 | 1150.0629 | 0.1023 | -0.0010004 | 0.017413 | 3.5947e-06 | 9.4151e-09 | 6.294e-08 | -7.6976e-10 | 74.6842 | 4.674 | 0.15392 | 3.4028e+38 | 9.8816 | 0.021226 | 0 | 0 | 0 | -6.5242e-05 | 0.019455 | 3.239e-06 | 8.325e-09 | 6.4063e-08 | 8.8368e-09 | 74.6942 | -48.7313 | 0.13729 | 3.4028e+38 | 25.5177 | 229.7061 | 1.9283e-09 | 0.00074433 | 2.3047e-09 | -7.6787e-05 | 3.4028e+38 | 3.4028e+38 | 1150.0651 | 1150.0625 | 0.059401 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 1150.0867 | 1150.0867 | 0.10872 | 0.0016501 | 230.0019 | 5.0002 | 1150.0637 | 1150.0612 | 0.058354 | 49.9999 | -0.0029749 | [17 empty] | 21.6878 | -0.1759 | 22.6461 | 20.4862 | 1605028.8796 |  | -0.0018589
153 | 230.0019 | 5.0002 | 574.9778 | 1150.062 | 996.0151 | 50 | -60.0031 | 0.015682 | 3.5506e-06 | 4.7539e-09 | 5.4549e-08 | 7.3261e-09 | 74.7478 | -57.0202 | 0.017153 | 3.1593e-06 | -8.2257e-10 | 5.3117e-08 | 4.0932e-09 | 74.7076 | -101.3627 | 3.4028e+38 | 3.4028e+38 | 574.9778 | 1150.062 | 996.0151 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 5.0002 | 574.9785 | 1150.0634 | 996.0186 | 50 | -60.003 | 0.014619 | 3.5876e-06 | 5.4919e-09 | 5.2074e-08 | 2.8647e-09 | 74.6052 | -33.8031 | 0.015143 | 3.1608e-06 | 1.5912e-09 | 4.7774e-08 | 2.7312e-09 | 74.7141 | -72.2675 | 3.4028e+38 | 3.4028e+38 | 574.9785 | 1150.0634 | 996.0186 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 5.0002 | 574.9791 | 1150.0675 | 996.0201 | 49.9999 | -60.0032 | 0.015899 | 3.5531e-06 | 4.67e-09 | 5.6486e-08 | 7.7399e-09 | 74.6254 | -54.4571 | 0.017224 | 3.0864e-06 | -3.7581e-10 | 5.6009e-08 | 4.1683e-09 | 74.8294 | -103.5005 | 3.4028e+38 | 3.4028e+38 | 574.979 | 1150.0675 | 996.0201 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 5.0002 | 574.9807 | 1150.0647 | 996.0181 | 49.9999 | -60.0029 | 0.015704 | 3.5781e-06 | 8.4727e-09 | 5.613e-08 | 3.6506e-09 | 74.7242 | -20.7904 | 0.016827 | 3.168e-06 | 1.3969e-09 | 5.5045e-08 | 5.4984e-09 | 74.7409 | -74.3396 | 3.4028e+38 | 3.4028e+38 | 574.9807 | 1150.0647 | 996.0181 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 5.0003 | 574.9825 | 1150.0688 | 996.0178 | 49.9999 | -60.003 | 0.014318 | 3.5577e-06 | 4.9514e-09 | 5.0749e-08 | 4.1046e-09 | 74.7434 | -39.6581 | 0.015259 | 3.1429e-06 | 4.7184e-10 | 4.7539e-08 | 2.4448e-09 | 74.7229 | -58.5952 | 3.4028e+38 | 3.4028e+38 | 574.9825 | 1150.0688 | 996.0178 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0019 | 5.0002 | 574.9847 | 1150.067 | 996.015 | 50 | -60.0028 | 0.49996 | 3.4028e+38 | 0 | 230.0023 | 574.9648 | 1150.0684 | 996.0192 | -0.0012439 | 0.014389 | 3.5915e-06 | 4.1194e-09 | 5.171e-08 | 5.3032e-09 | 74.7297 | -47.1337 | 0.10106 | 3.4028e+38 | 21.6223 | 0.019957 | 0 | 0 | 0 | 9.7023e-05 | 0.015838 | 3.1047e-06 | -1.4508e-09 | 5.1525e-08 | 2.6999e-09 | 74.8486 | -93.9338 | -0.0036033 | 3.4028e+38 | -70.2432 | 229.208 | -4.358e-10 | 0.00070864 | 4.1251e-10 | -8.2666e-05 | 3.4028e+38 | 3.4028e+38 | 574.978 | 1150.0626 | 996.0142 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 574.9468 | 574.9468 | 996.0078 | -0.0019472 | 230.0019 | 5.0002 | 574.9797 | 1150.0653 | 996.0179 | 49.9999 | -60.0031 | [17 empty] | 21.5608 | 43.8304 | -70.5883 | 65.8042 | 112.887 |  | -0.0045199
154 | 230.0019 | 5.0002 | 574.9803 | 1150.0626 | 996.0148 | 49.9999 | -60.003 | 0.015807 | 3.6041e-06 | 4.3088e-09 | 5.7059e-08 | 7.6266e-09 | 74.7313 | -59.5768 | 0.017293 | 3.1287e-06 | -7.1128e-10 | 5.547e-08 | 4.2613e-09 | 74.782 | -110.0478 | 3.4028e+38 | 3.4028e+38 | 574.9802 | 1150.0626 | 996.0148 | 3.4028e+38 | 3.4028e+38 | 230.0018 | 5.0003 | 574.9833 | 1150.0662 | 996.0146 | 50 | -60.0029 | 0.016688 | 3.533e-06 | 1.1594e-08 | 5.9087e-08 | 5.7378e-09 | 74.6282 | -27.0554 | 0.017974 | 3.1104e-06 | 2.402e-09 | 5.5945e-08 | 7.0644e-09 | 74.8134 | -67.2433 | 3.4028e+38 | 3.4028e+38 | 574.9833 | 1150.0662 | 996.0146 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 5.0002 | 574.9841 | 1150.0653 | 996.0154 | 50 | -60.0029 | 0.014345 | 3.5867e-06 | 5.6486e-09 | 5.1453e-08 | 3.912e-09 | 74.7761 | -37.451 | 0.015255 | 3.1615e-06 | 4.298e-10 | 4.743e-08 | 1.9206e-09 | 74.8192 | -98.5425 | 3.4028e+38 | 3.4028e+38 | 574.9841 | 1150.0653 | 996.0154 | 3.4028e+38 | 3.4028e+38 | 230.0021 | 5.0002 | 574.9794 | 1150.0653 | 996.0149 | 50 | -60.0028 | 0.017037 | 3.5612e-06 | 1.0538e-08 | 6.0654e-08 | 5.3585e-09 | 74.6612 | -27.596 | 0.018456 | 3.1177e-06 | 3.9284e-09 | 5.9339e-08 | 7.6193e-09 | 74.8408 | -69.9549 | 3.4028e+38 | 3.4028e+38 | 574.9794 | 1150.0653 | 996.0149 | 3.4028e+38 | 3.4028e+38 | 230.002 | 5.0002 | 574.981 | 1150.0658 | 996.017 | 50 | -60.003 | 0.017394 | 3.5059e-06 | 6.6759e-09 | 6.1617e-08 | 9.5737e-09 | 74.7453 | -55.1112 | 0.019186 | 3.1737e-06 | 2.8773e-10 | 6.0907e-08 | 7.183e-09 | 74.6752 | -92.9139 | 3.4028e+38 | 3.4028e+38 | 574.981 | 1150.0658 | 996.017 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0022 | 5.0002 | 574.9808 | 1150.0612 | 996.0129 | 49.9999 | -60.003 | 0.49996 | 3.4028e+38 | 0 | 230.0009 | 574.939 | 1150.0529 | 996.0016 | -0.0038677 | 0.016724 | 3.5327e-06 | 6.4297e-09 | 5.8479e-08 | 8.8632e-09 | 74.7476 | -53.832 | 0.10453 | 3.4028e+38 | -49.625 | 0.020988 | 0 | 0 | 0 | 9.0654e-05 | 0.017681 | 3.1043e-06 | 1.0729e-10 | 5.4869e-08 | 4.9038e-09 | 74.8273 | -106.1931 | -0.024878 | 3.4028e+38 | 21.8079 | 229.5366 | 8.9616e-10 | 0.00070963 | -1.413e-09 | -8.7201e-05 | 3.4028e+38 | 3.4028e+38 | 574.981 | 1150.063 | 996.0135 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 574.9412 | 574.9412 | 996.0086 | -0.0030285 | 230.002 | 5.0002 | 574.9816 | 1150.065 | 996.0153 | 50 | -60.0029 | [17 empty] | 22.2053 | 34.277 | -73.811 | 56.7119 | 100.6315 |  | -0.0043142
155 | 230.0023 | 50.0005 | 11500.2031 | 11500.2139 | 0.38245 | 49.9999 | -0.0019054 | 0.01681 | 8.7615e-06 | 7.3816e-08 | 1.4754e-07 | 1.8424e-08 | 71.9752 | -13.6062 | 0.018703 | 8.7495e-06 | 5.0943e-08 | 1.5416e-07 | 8.0673e-08 | 54.8806 | -55.9717 | 3.4028e+38 | 3.4028e+38 | 11500.2031 | 11500.2139 | 0.38245 | 3.4028e+38 | 3.4028e+38 | 230.0025 | 50.0007 | 11500.2695 | 11500.2783 | 0.35006 | 49.9999 | -0.0020176 | 0.017141 | 9.2596e-06 | 5.4909e-08 | 1.5829e-07 | -2.922e-08 | 71.9707 | 22.2425 | 0.019006 | 8.5424e-06 | 7.7608e-08 | 1.6292e-07 | 6.3323e-08 | 71.1894 | -39.4304 | 3.4028e+38 | 3.4028e+38 | 11500.2695 | 11500.2783 | 0.35006 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 50.0005 | 11500.2236 | 11500.251 | 0.39277 | 49.9999 | -0.0019569 | 0.017781 | 9.0395e-06 | 7.7541e-08 | 1.6091e-07 | 3.147e-11 | 68.4515 | -3.9444 | 0.019605 | 8.5835e-06 | 6.8397e-08 | 1.7272e-07 | 9.2017e-08 | 71.4136 | -52.5929 | 3.4028e+38 | 3.4028e+38 | 11500.2236 | 11500.251 | 0.39277 | 3.4028e+38 | 3.4028e+38 | 230.0027 | 50.0006 | 11500.251 | 11500.2578 | 0.38306 | 49.9999 | -0.0019654 | 0.017646 | 9.0976e-06 | 7.2686e-08 | 1.6109e-07 | -7.9945e-09 | 71.0762 | 9.227 | 0.019686 | 8.6007e-06 | 7.4169e-08 | 1.688e-07 | 8.4279e-08 | 72.7951 | -47.8939 | 3.4028e+38 | 3.4028e+38 | 11500.251 | 11500.2578 | 0.38306 | 3.4028e+38 | 3.4028e+38 | 230.0025 | 50.0005 | 11500.2227 | 11500.2295 | 0.39137 | 49.9999 | -0.0019499 | 0.016328 | 8.8544e-06 | 5.958e-08 | 1.374e-07 | 2.2204e-08 | 71.0869 | -23.4338 | 0.017472 | 8.6569e-06 | 2.443e-08 | 1.4463e-07 | 6.1689e-08 | 70.2999 | -68.8095 | 3.4028e+38 | 3.4028e+38 | 11500.2227 | 11500.2295 | 0.39137 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0025 | 50.0006 | 11500.248 | 11500.2549 | 0.40533 | 50 | -0.0021489 | 1 | 3.4028e+38 | 0 | 230.0038 | 11500.3047 | 11500.3135 | 0.60458 | 0.0075757 | 0.014658 | 9.2386e-06 | 2.9067e-08 | 1.3992e-07 | -1.6859e-08 | 71.8124 | 30.2354 | 0.2145 | 3.4028e+38 | -35.4423 | 0.019717 | 0 | 0 | 0 | -0.00030016 | 0.017581 | 8.475e-06 | 6.7247e-08 | 1.494e-07 | 4.3845e-08 | 69.6144 | -34.2456 | 0.47786 | 3.4028e+38 | -43.9859 | 230.0973 | 8.7336e-09 | 0.0019303 | 7.932e-09 | -0.00078724 | 3.4028e+38 | 3.4028e+38 | 11500.2471 | 11500.2686 | 0.36457 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 11500.3535 | 11500.3535 | 0.91161 | 0.0078145 | 230.0024 | 50.0006 | 11500.234 | 11500.2461 | 0.37994 | 49.9999 | -0.001959 | [17 empty] | 26.0439 | -1.6983 | 22.7535 | 23.8121 | 2260253.2165 |  | -0.0013784
156 | 230.0024 | 50.0005 | 11500.251 | 11500.2568 | 0.38524 | 50 | -0.0021387 | 0.016912 | 8.803e-06 | 6.71e-08 | 1.466e-07 | 2.1634e-08 | 72.2269 | -22.0257 | 0.018242 | 8.6822e-06 | 3.9538e-08 | 1.5857e-07 | 7.9514e-08 | 66.9158 | -65.4833 | 3.4028e+38 | 3.4028e+38 | 11500.251 | 11500.2568 | 0.38524 | 3.4028e+38 | 3.4028e+38 | 230.0023 | 50.0005 | 11500.2207 | 11500.2217 | 0.34118 | 50 | -0.0016998 | 0.016862 | 8.7587e-06 | 6.522e-08 | 1.4489e-07 | 2.5945e-08 | 71.8137 | -20.6357 | 0.018134 | 8.909e-06 | 3.6829e-08 | 1.5908e-07 | 7.0988e-08 | 72.8792 | -68.0683 | 3.4028e+38 | 3.4028e+38 | 11500.2207 | 11500.2217 | 0.34118 | 3.4028e+38 | 3.4028e+38 | 230.0026 | 50.0007 | 11500.2695 | 11500.2754 | 0.41368 | 49.9999 | -0.002061 | 0.01436 | 9.1443e-06 | 2.6691e-08 | 1.312e-07 | -7.6409e-09 | 71.9842 | 15.3308 | 0.015097 | 8.6485e-06 | 3.3039e-08 | 1.3167e-07 | 2.7854e-08 | 68.9342 | -37.2566 | 3.4028e+38 | 3.4028e+38 | 11500.2695 | 11500.2754 | 0.41368 | 3.4028e+38 | 3.4028e+38 | 230.0027 | 50.0007 | 11500.2607 | 11500.2676 | 0.39317 | 49.9999 | -0.0019661 | 0.017581 | 9.1908e-06 | 6.7998e-08 | 1.6283e-07 | -1.5583e-08 | 72.1398 | 9.8639 | 0.019445 | 8.601e-06 | 7.6385e-08 | 1.6722e-07 | 7.9897e-08 | 72.0812 | -46.6314 | 3.4028e+38 | 3.4028e+38 | 11500.2607 | 11500.2676 | 0.39317 | 3.4028e+38 | 3.4028e+38 | 230.0024 | 50.0006 | 11500.2441 | 11500.2529 | 0.43453 | 49.9999 | -0.0021649 | 0.016922 | 9.3093e-06 | 5.5955e-08 | 1.581e-07 | -2.3674e-08 | 72.1193 | 23.2266 | 0.018204 | 8.4572e-06 | 7.6234e-08 | 1.5757e-07 | 5.8367e-08 | 70.7581 | -39.3271 | 3.4028e+38 | 3.4028e+38 | 11500.2441 | 11500.2529 | 0.43453 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0027 | 50.0007 | 11500.2754 | 11500.2803 | 0.38168 | 49.9999 | -0.0020559 | 1 | 3.4028e+38 | 0 | 230.0059 | 11500.418 | 11500.4258 | 0.91588 | 0.009885 | 0.01786 | 9.1811e-06 | 7.4557e-08 | 1.6165e-07 | -7.4809e-09 | 72.2041 | 6.3593 | 0.46976 | 3.4028e+38 | 6.7081 | 0.020833 | 0 | 0 | 0 | -0.00033665 | 0.01963 | 8.5797e-06 | 6.9448e-08 | 1.693e-07 | 9.2221e-08 | 71.3007 | -53.6352 | 0.39217 | 3.4028e+38 | -40.9579 | 229.4867 | 5.4605e-09 | 0.0019679 | 8.1636e-09 | -0.00079333 | 3.4028e+38 | 3.4028e+38 | 11500.2354 | 11500.2422 | 0.36985 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 11500.2197 | 11500.2197 | 0.78884 | -0.014646 | 230.0025 | 50.0006 | 11500.2492 | 11500.2549 | 0.39356 | 49.9999 | -0.0020061 | [17 empty] | 25.1514 | -4.7418 | 19.3494 | 19.8481 | 2574637.4383 |  | -0.0014576
157 | 230.004 | 50.0008 | 5749.9126 | 11500.3828 | 9959.7969 | 50 | -60.0017 | 0.015642 | 9.318e-06 | 5.3798e-08 | 1.4387e-07 | 2.8145e-08 | 72.0376 | -27.1452 | 0.016676 | 8.3546e-06 | 1.8691e-08 | 1.3522e-07 | 4.6645e-08 | 70.8423 | -73.5739 | 3.4028e+38 | 3.4028e+38 | 5749.9126 | 11500.3828 | 9959.7969 | 3.4028e+38 | 3.4028e+38 | 230.0041 | 50.0006 | 5749.8936 | 11500.3525 | 9959.791 | 49.9999 | -60.0017 | 0.015997 | 8.9785e-06 | 3.6842e-08 | 1.48e-07 | 6.8249e-08 | 71.9768 | -60.1059 | 0.017434 | 8.5412e-06 | -3.808e-08 | 1.4891e-07 | 3.2125e-08 | 71.9199 | -137.7566 | 3.4028e+38 | 3.4028e+38 | 5749.8931 | 11500.3525 | 9959.791 | 3.4028e+38 | 3.4028e+38 | 230.0037 | 50.0007 | 5749.8979 | 11500.3398 | 9959.7715 | 49.9999 | -60.0017 | 0.016779 | 9.0741e-06 | 7.2179e-08 | 1.5245e-07 | 4.7996e-08 | 71.998 | -31.9308 | 0.018267 | 8.438e-06 | 1.4133e-08 | 1.5398e-07 | 7.3292e-08 | 67.9415 | -80.922 | 3.4028e+38 | 3.4028e+38 | 5749.8979 | 11500.3398 | 9959.7715 | 3.4028e+38 | 3.4028e+38 | 230.0042 | 50.0008 | 5749.9062 | 11500.3818 | 9959.8076 | 50 | -60.0016 | 0.014678 | 9.4293e-06 | 4.351e-08 | 1.378e-07 | 2.6361e-08 | 72.5679 | -27.6031 | 0.015837 | 8.3465e-06 | 9.7055e-09 | 1.3299e-07 | 4.3609e-08 | 72.2933 | -73.7025 | 3.4028e+38 | 3.4028e+38 | 5749.9062 | 11500.3818 | 9959.8076 | 3.4028e+38 | 3.4028e+38 | 230.0039 | 50.0006 | 5749.873 | 11500.3047 | 9959.7129 | 50 | -60.0016 | 0.017047 | 8.8819e-06 | 5.1146e-08 | 1.5234e-07 | 7.3699e-08 | 72.5905 | -54.2589 | 0.018952 | 8.8099e-06 | -3.9276e-08 | 1.6347e-07 | 5.4761e-08 | 72.0931 | -122.3891 | 3.4028e+38 | 3.4028e+38 | 5749.8726 | 11500.3047 | 9959.7129 | 3.4028e+38 | 3.4028e+38 | 123 | 230.004 | 50.0006 | 5749.8682 | 11500.3301 | 9959.7168 | 50 | -60.0016 | 0.49998 | 3.4028e+38 | 0 | 230.0047 | 5749.4233 | 11500.3398 | 9959.7627 | -0.032479 | 0.0175 | 8.7901e-06 | 6.2772e-08 | 1.5523e-07 | 7.5818e-08 | 72.4087 | -49.1385 | 0.41279 | 3.4028e+38 | 22.3834 | 0.021363 | 0 | 0 | 0 | 0.00069768 | 0.019581 | 8.748e-06 | -2.7117e-08 | 1.6917e-07 | 6.7152e-08 | 70.7537 | -111.9892 | -0.13239 | 3.4028e+38 | -56.9315 | 229.5175 | -1.2391e-09 | 0.0020156 | 1.8489e-08 | -0.00094839 | 3.4028e+38 | 3.4028e+38 | 5749.9019 | 11500.3076 | 9959.7148 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 5749.4614 | 5749.4614 | 9959.8438 | -0.033375 | 230.004 | 50.0007 | 5749.8967 | 11500.3523 | 9959.776 | 50 | -60.0017 | [17 empty] | 31.9642 | 33.7131 | -52.3783 | 65.1238 | 106.028 |  | -0.0039345
158 | 230.004 | 50.0006 | 5749.8896 | 11500.3564 | 9959.7881 | 49.9999 | -60.0018 | 0.016585 | 9.0921e-06 | 7.3292e-08 | 1.5114e-07 | 4.1833e-08 | 72.3196 | -29.272 | 0.018035 | 8.3446e-06 | 1.3406e-08 | 1.5535e-07 | 7.621e-08 | 71.5016 | -78.7895 | 3.4028e+38 | 3.4028e+38 | 5749.8896 | 11500.3564 | 9959.7881 | 3.4028e+38 | 3.4028e+38 | 230.0038 | 50.0005 | 5749.8994 | 11500.334 | 9959.7861 | 49.9998 | -60.0018 | 0.016668 | 8.9784e-06 | 3.7813e-08 | 1.4671e-07 | 6.6263e-08 | 72.1558 | -60.9624 | 0.018091 | 8.7362e-06 | -4.3508e-08 | 1.5624e-07 | 3.7963e-08 | 72.1134 | -141.0527 | 3.4028e+38 | 3.4028e+38 | 5749.8989 | 11500.334 | 9959.7861 | 3.4028e+38 | 3.4028e+38 | 230.0042 | 50.0007 | 5749.8862 | 11500.3389 | 9959.7812 | 50 | -60.0016 | 0.015155 | 9.2884e-06 | 5.0222e-08 | 1.3996e-07 | 2.6866e-08 | 71.6914 | -27.7317 | 0.016167 | 8.4339e-06 | 1.2089e-08 | 1.353e-07 | 4.7562e-08 | 71.4378 | -71.4698 | 3.4028e+38 | 3.4028e+38 | 5749.8862 | 11500.3389 | 9959.7812 | 3.4028e+38 | 3.4028e+38 | 230.0038 | 50.0007 | 5749.918 | 11500.373 | 9959.7744 | 50 | -60.0017 | 0.016017 | 9.2299e-06 | 6.7128e-08 | 1.5095e-07 | 3.8472e-08 | 72.2127 | -28.2514 | 0.017348 | 8.2973e-06 | 1.7217e-08 | 1.4559e-07 | 6.6038e-08 | 70.8411 | -74.3837 | 3.4028e+38 | 3.4028e+38 | 5749.918 | 11500.373 | 9959.7744 | 3.4028e+38 | 3.4028e+38 | 230.0043 | 50.0007 | 5749.9053 | 11500.374 | 9959.8047 | 50 | -60.0016 | 0.014795 | 9.2777e-06 | 2.8583e-08 | 1.3651e-07 | 3.7895e-08 | 72.4982 | -50.1451 | 0.01582 | 8.4064e-06 | -1.7153e-08 | 1.328e-07 | 2.1323e-08 | 68.531 | -124.9931 | 3.4028e+38 | 3.4028e+38 | 5749.9048 | 11500.374 | 9959.8047 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0043 | 50.0007 | 5749.9165 | 11500.3701 | 9959.7812 | 49.9999 | -60.0017 | 0.49998 | 3.4028e+38 | 0 | 230.006 | 5749.7397 | 11500.4648 | 9959.9082 | 0.0025982 | 0.014578 | 9.3745e-06 | 3.493e-08 | 1.36e-07 | 2.8439e-08 | 71.7687 | -37.2721 | 0.27476 | 3.4028e+38 | -35.7856 | 0.020009 | 0 | 0 | 0 | 0.00074419 | 0.015385 | 8.3998e-06 | 8.0915e-09 | 1.3082e-07 | 3.6036e-08 | 72.4342 | -78.9184 | 0.053151 | 3.4028e+38 | -97.4493 | 229.1005 | -3.3366e-09 | 0.0019317 | -7.0974e-10 | -0.00091854 | 3.4028e+38 | 3.4028e+38 | 5749.9268 | 11500.4004 | 9959.8135 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 5749.8057 | 5749.8057 | 9959.959 | -0.00050842 | 230.004 | 50.0007 | 5749.8997 | 11500.3553 | 9959.7869 | 49.9999 | -60.0017 | [17 empty] | 31.6807 | 40.8662 | -50.3931 | 72.8209 | 116.5639 |  | -0.0041706
159 | 230.0011 | 50.0005 | 9196.7715 | 11500.2217 | -6904.7754 | 49.9999 | 36.8984 | 0.015837 | 9.3471e-06 | 1.0305e-10 | 1.4841e-07 | -5.7498e-08 | 72.0768 | 92.9973 | 0.01737 | 8.5659e-06 | 6.1699e-08 | 1.4489e-07 | -4.6812e-10 | 68.1864 | 2.8816 | 3.4028e+38 | 3.4028e+38 | 9196.7715 | 11500.2217 | -6904.7759 | 3.4028e+38 | 3.4028e+38 | 230.0007 | 50.0006 | 9196.7646 | 11500.1758 | -6904.772 | 49.9999 | 36.8984 | 0.017467 | 9.3575e-06 | 2.4495e-08 | 1.6234e-07 | -6.2964e-08 | 71.2226 | 68.4758 | 0.019347 | 8.5141e-06 | 8.9755e-08 | 1.641e-07 | 3.6276e-08 | 72.075 | -21.4912 | 3.4028e+38 | 3.4028e+38 | 9196.7646 | 11500.1758 | -6904.7725 | 3.4028e+38 | 3.4028e+38 | 230.0011 | 50.0006 | 9196.79 | 11500.1914 | -6904.7568 | 49.9999 | 36.8984 | 0.016871 | 9.4038e-06 | 9.1428e-09 | 1.5918e-07 | -6.8332e-08 | 71.9462 | 80.658 | 0.018741 | 8.5835e-06 | 7.8414e-08 | 1.5773e-07 | 1.9652e-08 | 70.9148 | -13.2642 | 3.4028e+38 | 3.4028e+38 | 9196.79 | 11500.1914 | -6904.7573 | 3.4028e+38 | 3.4028e+38 | 230.0014 | 50.0005 | 9196.7803 | 11500.1641 | -6904.729 | 49.9999 | 36.8983 | 0.01443 | 8.9247e-06 | 2.0175e-08 | 1.2879e-07 | -1.5921e-08 | 71.6913 | 35.036 | 0.015537 | 8.8741e-06 | 3.5085e-08 | 1.3891e-07 | 2.0201e-08 | 65.4639 | -30.7332 | 3.4028e+38 | 3.4028e+38 | 9196.7803 | 11500.1641 | -6904.7295 | 3.4028e+38 | 3.4028e+38 | 230.0016 | 50.0004 | 9196.7715 | 11500.168 | -6904.7505 | 49.9998 | 36.8985 | 0.01493 | 8.841e-06 | 3.4463e-08 | 1.3082e-07 | -9.2028e-09 | 71.798 | 18.5062 | 0.016145 | 8.7753e-06 | 4.2759e-08 | 1.4182e-07 | 4.1355e-08 | 70.8613 | -43.5173 | 3.4028e+38 | 3.4028e+38 | 9196.7715 | 11500.168 | -6904.751 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0012 | 50.0004 | 9196.75 | 11500.1514 | -6904.7246 | 49.9999 | 36.8986 | 0.79971 | 3.4028e+38 | 0 | 229.9998 | 9196.8867 | 11500.0859 | -6904.4082 | -0.0051521 | 0.015293 | 8.7485e-06 | 3.7078e-08 | 1.3046e-07 | -7.4212e-09 | 71.8041 | 19.0519 | 0.25414 | 3.4028e+38 | -22.9448 | 0.02 | 0 | 0 | 0 | -0.00080932 | 0.015891 | 8.8402e-06 | 4e-08 | 1.4111e-07 | 3.5941e-08 | 69.4116 | -41.94 | 0.27906 | 3.4028e+38 | -84.8597 | 229.8717 | 1.0552e-09 | 0.0020215 | -3.2954e-10 | -0.00027725 | 3.4028e+38 | 3.4028e+38 | 9196.7314 | 11500.1621 | -6904.75 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 9196.8809 | 9196.8799 | -6904.5181 | 0.00043446 | 230.0012 | 50.0005 | 9196.7756 | 11500.1842 | -6904.7567 | 49.9999 | 36.8984 | [17 empty] | 18.526 | 30.762 | 52.7467 | 49.774 | 66.6207 |  | 0.00031249
160 | 230.0008 | 50.0006 | 9196.7764 | 11500.1787 | -6904.7207 | 50 | 36.8982 | 0.015955 | 9.2914e-06 | -7.4171e-10 | 1.4857e-07 | -5.552e-08 | 71.8316 | 88.4721 | 0.017317 | 8.5239e-06 | 6.4664e-08 | 1.484e-07 | 2.26e-09 | 67.9276 | -2.1188 | 3.4028e+38 | 3.4028e+38 | 9196.7754 | 11500.1787 | -6904.7212 | 3.4028e+38 | 3.4028e+38 | 230.0011 | 50.0005 | 9196.748 | 11500.1436 | -6904.7388 | 49.9999 | 36.8986 | 0.014676 | 8.8723e-06 | 2.7394e-08 | 1.2997e-07 | -1.1065e-08 | 72.1057 | 20.5864 | 0.015727 | 8.6296e-06 | 3.8871e-08 | 1.3421e-07 | 2.8605e-08 | 68.5903 | -37.3505 | 3.4028e+38 | 3.4028e+38 | 9196.748 | 11500.1436 | -6904.7393 | 3.4028e+38 | 3.4028e+38 | 230.001 | 50.0006 | 9196.7607 | 11500.1738 | -6904.7559 | 49.9999 | 36.8985 | 0.017595 | 9.2723e-06 | 3.4052e-08 | 1.6176e-07 | -5.608e-08 | 71.7186 | 52.4226 | 0.019797 | 8.4603e-06 | 9.7042e-08 | 1.6991e-07 | 4.875e-08 | 72.3263 | -33.221 | 3.4028e+38 | 3.4028e+38 | 9196.7607 | 11500.1738 | -6904.7563 | 3.4028e+38 | 3.4028e+38 | 230.001 | 50.0006 | 9196.7852 | 11500.1914 | -6904.7485 | 49.9999 | 36.8984 | 0.016648 | 9.3762e-06 | 2.6684e-09 | 1.5141e-07 | -5.8941e-08 | 70.8586 | 88.1325 | 0.018029 | 8.4934e-06 | 6.8503e-08 | 1.5155e-07 | 3.57e-09 | 69.7049 | -3.1078 | 3.4028e+38 | 3.4028e+38 | 9196.7852 | 11500.1914 | -6904.749 | 3.4028e+38 | 3.4028e+38 | 230.0009 | 50.0006 | 9196.7842 | 11500.2021 | -6904.7842 | 49.9999 | 36.8984 | 0.017357 | 9.4561e-06 | 1.5096e-08 | 1.6109e-07 | -6.3768e-08 | 71.5515 | 73.5361 | 0.01926 | 8.4373e-06 | 8.667e-08 | 1.6196e-07 | 2.2406e-08 | 71.6542 | -16.3421 | 3.4028e+38 | 3.4028e+38 | 9196.7842 | 11500.2021 | -6904.7847 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0007 | 50.0005 | 9196.752 | 11500.2061 | -6904.7769 | 49.9999 | 36.8985 | 0.79971 | 3.4028e+38 | 0 | 230.0046 | 9197.1963 | 11500.3447 | -6904.2905 | 0.037757 | 0.017622 | 9.3554e-06 | 3.1295e-08 | 1.6531e-07 | -5.6152e-08 | 73.0382 | 67.618 | 0.15793 | 3.4028e+38 | -14.8518 | 0.019977 | 0 | 0 | 0 | -0.00088648 | 0.019755 | 8.5372e-06 | 9.2649e-08 | 1.6696e-07 | 4.7675e-08 | 71.2219 | -26.6407 | 0.55871 | 3.4028e+38 | -11.6761 | 229.1965 | 9.4048e-09 | 0.0019717 | -6.3354e-09 | -0.00029294 | 3.4028e+38 | 3.4028e+38 | 9196.7842 | 11500.1943 | -6904.7832 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 9197.0752 | 9197.0752 | -6904.2671 | 0.02883 | 230.001 | 50.0006 | 9196.7709 | 11500.1779 | -6904.7496 | 49.9999 | 36.8984 | [17 empty] | 18.2974 | 28.8099 | 52.6948 | 46.8932 | 58.2914 |  | 9.765e-05
161 | 230.0029 | 100.0009 | 23000.4922 | 23000.5039 | 1.1446 | 49.9999 | -0.0027087 | 0.016679 | 1.5453e-05 | 1.2461e-07 | 2.5454e-07 | -2.4006e-08 | 51.0661 | 6.8835 | 0.017932 | 1.5089e-05 | 7.6962e-08 | 2.7052e-07 | 1.0231e-07 | 49.9947 | -60.8303 | 3.4028e+38 | 3.4028e+38 | 23000.4922 | 23000.5039 | 1.1446 | 3.4028e+38 | 3.4028e+38 | 230.0029 | 100.0011 | 23000.5234 | 23000.5371 | 1.0833 | 49.9999 | -0.0026987 | 0.016214 | 1.6453e-05 | 6.6532e-08 | 2.6733e-07 | -1.139e-07 | 49.9984 | 59.7751 | 0.017541 | 1.4442e-05 | 1.2882e-07 | 2.4992e-07 | 1.1858e-08 | 49.9954 | -5.3888 | 3.4028e+38 | 3.4028e+38 | 23000.5234 | 23000.5371 | 1.0833 | 3.4028e+38 | 3.4028e+38 | 230.0029 | 100.0009 | 23000.4805 | 23000.4941 | 1.0731 | 49.9999 | -0.0027275 | 0.018056 | 1.5872e-05 | 1.3786e-07 | 2.8454e-07 | -9.9038e-08 | 49.9308 | 35.3321 | 0.019591 | 1.4718e-05 | 1.4561e-07 | 2.8837e-07 | 1.0051e-07 | 50.0258 | -34.6018 | 3.4028e+38 | 3.4028e+38 | 23000.4805 | 23000.4941 | 1.0731 | 3.4028e+38 | 3.4028e+38 | 230.0035 | 100.0008 | 23000.4844 | 23000.4961 | 1.1584 | 49.9999 | -0.002703 | 0.014361 | 1.5996e-05 | 6.298e-08 | 2.2906e-07 | -3.2062e-08 | 50.0141 | 27.0246 | 0.01543 | 1.5004e-05 | 4.9274e-08 | 2.3545e-07 | 3.8054e-08 | 50.0189 | -37.0658 | 3.4028e+38 | 3.4028e+38 | 23000.4844 | 23000.4961 | 1.1584 | 3.4028e+38 | 3.4028e+38 | 230.0029 | 100.001 | 23000.4941 | 23000.5234 | 1.1028 | 50 | -0.0027472 | 0.017701 | 1.6105e-05 | 1.2084e-07 | 2.8508e-07 | -1.1343e-07 | 49.9887 | 41.4777 | 0.019493 | 1.4533e-05 | 1.5407e-07 | 2.8491e-07 | 7.6781e-08 | 49.9859 | -29.0996 | 3.4028e+38 | 3.4028e+38 | 23000.4941 | 23000.5234 | 1.1028 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0034 | 100.0009 | 23000.498 | 23000.5117 | 1.1174 | 50 | -0.0025743 | 1 | 3.4028e+38 | 0 | 230.0049 | 23000.6934 | 23000.6992 | 2.1708 | 0.037383 | 0.017771 | 1.5796e-05 | 1.4379e-07 | 2.8e-07 | -7.4999e-08 | 49.9338 | 27.3331 | 0.51738 | 3.4028e+38 | 153.7348 | 0.020999 | 0 | 0 | 0 | -0.00086647 | 0.019427 | 1.495e-05 | 1.0953e-07 | 2.8015e-07 | 1.2205e-07 | 49.9014 | -49.7996 | 0.37659 | 3.4028e+38 | 83.8285 | 229.975 | 1.0505e-08 | 0.0034483 | 1.3089e-08 | -0.0014046 | 3.4028e+38 | 3.4028e+38 | 23000.4766 | 23000.4941 | 1.0954 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 23000.3828 | 23000.3438 | 2.0682 | -0.016309 | 230.003 | 100.001 | 23000.4949 | 23000.5109 | 1.1124 | 49.9999 | -0.002717 | [17 empty] | 29.0857 | -9.65 | 18.1139 | 18.8155 | 4944863.2998 |  | -0.0022509
162 | 230.0034 | 100.0008 | 23000.5117 | 23000.5762 | 1.0693 | 50 | -0.0026636 | 0.014502 | 1.6124e-05 | 5.5741e-08 | 2.2926e-07 | -4.2814e-08 | 50.0216 | 33.1897 | 0.0154 | 1.4916e-05 | 5.7846e-08 | 2.2966e-07 | 2.6117e-08 | 49.9939 | -25.4802 | 3.4028e+38 | 3.4028e+38 | 23000.5117 | 23000.5762 | 1.0693 | 3.4028e+38 | 3.4028e+38 | 230.0031 | 100.001 | 23000.4922 | 23000.5078 | 1.0619 | 49.9999 | -0.0027517 | 0.017538 | 1.5573e-05 | 1.4542e-07 | 2.7313e-07 | -7.3277e-08 | 49.9889 | 26.7514 | 0.019286 | 1.4891e-05 | 1.1661e-07 | 2.8788e-07 | 1.2086e-07 | 49.9727 | -44.6824 | 3.4028e+38 | 3.4028e+38 | 23000.4922 | 23000.5078 | 1.0619 | 3.4028e+38 | 3.4028e+38 | 230.0032 | 100.0009 | 23000.5176 | 23000.5254 | 1.1328 | 49.9999 | -0.0027152 | 0.016854 | 1.55e-05 | 1.3737e-07 | 2.613e-07 | -4.1127e-08 | 49.974 | 16.6519 | 0.018708 | 1.508e-05 | 9.0845e-08 | 2.7799e-07 | 1.2346e-07 | 50.0482 | -56.1888 | 3.4028e+38 | 3.4028e+38 | 23000.5176 | 23000.5254 | 1.1328 | 3.4028e+38 | 3.4028e+38 | 230.0034 | 100.001 | 23000.5625 | 23000.5762 | 1.1865 | 50 | -0.0029557 | 0.014707 | 1.6334e-05 | 4.6581e-08 | 2.4167e-07 | -7.401e-08 | 49.9916 | 56.1788 | 0.015989 | 1.4661e-05 | 8.1706e-08 | 2.3272e-07 | 7.0111e-09 | 49.9944 | -4.7706 | 3.4028e+38 | 3.4028e+38 | 23000.5625 | 23000.5762 | 1.1865 | 3.4028e+38 | 3.4028e+38 | 230.0028 | 100.001 | 23000.5137 | 23000.5293 | 1.1124 | 49.9999 | -0.0027577 | 0.017735 | 1.5765e-05 | 1.4132e-07 | 2.8264e-07 | -9.3416e-08 | 50.033 | 35.5122 | 0.019376 | 1.4733e-05 | 1.3899e-07 | 2.8547e-07 | 1.0647e-07 | 50.0037 | -36.5972 | 3.4028e+38 | 3.4028e+38 | 23000.5137 | 23000.5293 | 1.1124 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0029 | 100.001 | 23000.498 | 23000.5352 | 1.079 | 49.9999 | -0.0026878 | 1 | 3.4028e+38 | 0 | 230.0062 | 23000.8633 | 23000.877 | 2.1522 | 0.030171 | 0.017996 | 1.6064e-05 | 1.2453e-07 | 2.8517e-07 | -1.1391e-07 | 50.0507 | 43.7668 | 0.42914 | 3.4028e+38 | 122.9531 | 0.02034 | 0 | 0 | 0 | -0.00090606 | 0.019407 | 1.4463e-05 | 1.5396e-07 | 2.8067e-07 | 6.4698e-08 | 50.0237 | -20.9184 | 0.55518 | 3.4028e+38 | 175.5568 | 229.973 | 1.8967e-08 | 0.0033138 | -1.2798e-09 | -0.0013301 | 3.4028e+38 | 3.4028e+38 | 23000.5664 | 23000.582 | 1.0989 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 23001.0078 | 23001.0195 | 1.8897 | 0.059275 | 230.0032 | 100.001 | 23000.5195 | 23000.543 | 1.1126 | 49.9999 | -0.0027688 | [17 empty] | 28.7154 | -8.6876 | 19.1717 | 20.1966 | 5321423.0351 |  | -0.0023303
163 | 230.0067 | 100.001 | 11499.7061 | 23000.8809 | 19919.8145 | 49.9999 | -60.0022 | 0.01634 | 1.5871e-05 | 8.8801e-08 | 2.5725e-07 | 9.6653e-08 | 50.0322 | -47.4245 | 0.017439 | 1.4951e-05 | -6.2194e-08 | 2.6071e-07 | 1.1494e-07 | 49.9223 | -119.2037 | 3.4028e+38 | 3.4028e+38 | 11499.7051 | 23000.8809 | 19919.8145 | 3.4028e+38 | 3.4028e+38 | 230.0064 | 100.001 | 11499.7314 | 23000.8457 | 19919.7695 | 49.9999 | -60.0021 | 0.016247 | 1.6219e-05 | 1.3541e-07 | 2.6354e-07 | -7.5959e-09 | 49.9907 | 3.3092 | 0.017251 | 1.4366e-05 | 4.9245e-08 | 2.5016e-07 | 1.1429e-07 | 50.0551 | -67.1823 | 3.4028e+38 | 3.4028e+38 | 11499.7314 | 23000.8457 | 19919.7676 | 3.4028e+38 | 3.4028e+38 | 230.0065 | 100.0011 | 11499.7188 | 23000.832 | 19919.7656 | 50 | -60.002 | 0.015266 | 1.6489e-05 | 9.8825e-08 | 2.5174e-07 | -4.9274e-09 | 50.0018 | 1.3856 | 0.016194 | 1.4337e-05 | 3.6341e-08 | 2.3214e-07 | 7.7838e-08 | 49.9933 | -66.7945 | 3.4028e+38 | 3.4028e+38 | 11499.7188 | 23000.832 | 19919.7637 | 3.4028e+38 | 3.4028e+38 | 230.0062 | 100.0008 | 11499.7295 | 23000.8047 | 19919.7285 | 49.9999 | -60.0021 | 0.018083 | 1.5506e-05 | 1.5832e-07 | 2.8041e-07 | 4.6426e-08 | 49.9835 | -16.2283 | 0.019318 | 1.4889e-05 | -6.2339e-09 | 2.8757e-07 | 1.7438e-07 | 50.027 | -87.5543 | 3.4028e+38 | 3.4028e+38 | 11499.7285 | 23000.8047 | 19919.7285 | 3.4028e+38 | 3.4028e+38 | 230.0061 | 100.001 | 11499.7812 | 23000.8203 | 19919.7637 | 49.9999 | -60.002 | 0.017179 | 1.5894e-05 | 1.5638e-07 | 2.7309e-07 | 1.7398e-08 | 50.0484 | -6.2544 | 0.0186 | 1.4546e-05 | 3.6097e-08 | 2.6989e-07 | 1.5083e-07 | 50.0098 | -77.6091 | 3.4028e+38 | 3.4028e+38 | 11499.7812 | 23000.8203 | 19919.7637 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0059 | 100.001 | 11499.7568 | 23000.832 | 19919.7246 | 49.9999 | -60.0021 | 0.49997 | 3.4028e+38 | 0 | 230.01 | 11499.0566 | 23001.1934 | 19920.5215 | -0.040812 | 0.017757 | 1.5597e-05 | 1.6061e-07 | 2.8148e-07 | 5.449e-08 | 50.0246 | -20.4547 | 0.56842 | 3.4028e+38 | 107.594 | 0.021234 | 0 | 0 | 0 | 0.0011594 | 0.019424 | 1.4936e-05 | -2.359e-08 | 2.8798e-07 | 1.7623e-07 | 50.0308 | -100.9069 | -0.11436 | 3.4028e+38 | -178.1252 | 229.9498 | -5.9913e-09 | 0.00347 | 1.4324e-08 | -0.001889 | 3.4028e+38 | 3.4028e+38 | 11499.7295 | 23000.8184 | 19919.7461 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 11498.8096 | 11498.8096 | 19920.25 | -0.060047 | 230.0064 | 100.001 | 11499.7334 | 23000.8367 | 19919.7684 | 49.9999 | -60.0021 | [17 empty] | 41.3204 | 31.8231 | -60.4709 | 72.1827 | 119.611 |  | -0.0044747
164 | 230.0069 | 100.001 | 11499.7236 | 23000.9141 | 19919.752 | 49.9999 | -60.0021 | 0.015269 | 1.6485e-05 | 9.9564e-08 | 2.5172e-07 | -6.9535e-10 | 49.9915 | 0.43111 | 0.016303 | 1.4296e-05 | 3.3055e-08 | 2.2768e-07 | 7.5026e-08 | 50.0555 | -66.9387 | 3.4028e+38 | 3.4028e+38 | 11499.7236 | 23000.9141 | 19919.75 | 3.4028e+38 | 3.4028e+38 | 230.0068 | 100.0009 | 11499.7607 | 23000.8574 | 19919.8027 | 49.9999 | -60.0022 | 0.01446 | 1.6388e-05 | 6.2469e-08 | 2.358e-07 | 4.2797e-08 | 50.0093 | -35.2201 | 0.015341 | 1.455e-05 | -1.7488e-08 | 2.2658e-07 | 5.7739e-08 | 49.9952 | -110.2607 | 3.4028e+38 | 3.4028e+38 | 11499.7598 | 23000.8574 | 19919.8027 | 3.4028e+38 | 3.4028e+38 | 230.0065 | 100.001 | 11499.7324 | 23000.877 | 19919.752 | 50 | -60.002 | 0.015361 | 1.6088e-05 | 6.7924e-08 | 2.4808e-07 | 8.0672e-08 | 50.003 | -49.7975 | 0.016526 | 1.4797e-05 | -5.212e-08 | 2.4452e-07 | 8.5362e-08 | 49.9771 | -123.0147 | 3.4028e+38 | 3.4028e+38 | 11499.7314 | 23000.877 | 19919.752 | 3.4028e+38 | 3.4028e+38 | 230.0067 | 100.0009 | 11499.7891 | 23000.8398 | 19919.7051 | 50 | -60.0019 | 0.015806 | 1.6308e-05 | 1.2321e-07 | 2.5894e-07 | -8.2105e-09 | 49.9769 | 4.5364 | 0.01679 | 1.4283e-05 | 4.5225e-08 | 2.3967e-07 | 1.0018e-07 | 50.0683 | -65.7043 | 3.4028e+38 | 3.4028e+38 | 11499.7891 | 23000.8398 | 19919.7031 | 3.4028e+38 | 3.4028e+38 | 230.0063 | 100.0008 | 11499.7012 | 23000.8145 | 19919.6777 | 49.9999 | -60.0019 | 0.015899 | 1.5905e-05 | 7.8667e-08 | 2.5146e-07 | 9.0622e-08 | 50.0416 | -50.1951 | 0.017301 | 1.4984e-05 | -6.0651e-08 | 2.5505e-07 | 1.0364e-07 | 49.9869 | -120.9988 | 3.4028e+38 | 3.4028e+38 | 11499.7002 | 23000.8145 | 19919.6777 | 3.4028e+38 | 3.4028e+38 | 123 | 230.006 | 100.0007 | 11499.6865 | 23000.7754 | 19919.6699 | 50 | -60.0021 | 0.49997 | 3.4028e+38 | 0 | 230.0067 | 11499.3486 | 23000.8594 | 19919.9844 | 0.011699 | 0.014302 | 1.6497e-05 | 6.9033e-08 | 2.3846e-07 | 3.1282e-08 | 49.9475 | -16.0863 | 0.30151 | 3.4028e+38 | 114.9972 | 0.019785 | 0 | 0 | 0 | 0.0012465 | 0.015367 | 1.4395e-05 | 2.5195e-08 | 2.2358e-07 | 7.0101e-08 | 49.994 | -65.5024 | 0.15313 | 3.4028e+38 | -156.9142 | 229.998 | 4.5874e-09 | 0.0032677 | -2.2835e-09 | -0.0017816 | 3.4028e+38 | 3.4028e+38 | 11499.7607 | 23000.832 | 19919.6562 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 11499.3525 | 11499.3525 | 19920.582 | -0.0050175 | 230.0066 | 100.0009 | 11499.7414 | 23000.8605 | 19919.7379 | 49.9999 | -60.002 | [17 empty] | 43.1264 | 34.0183 | -55.9515 | 76.1711 | 120.7436 |  | -0.0044069
165 | 230.0002 | 100.0007 | 18400.8145 | 23000.1797 | -13799.5791 | 50 | 36.8679 | 0.01734 | 1.5852e-05 | 7.7304e-08 | 2.7228e-07 | -1.3449e-07 | 49.9815 | 57.7118 | 0.019129 | 1.4717e-05 | 1.6315e-07 | 2.8148e-07 | 3.8967e-08 | 50.0156 | -13.6725 | 3.4028e+38 | 3.4028e+38 | 18400.8145 | 23000.1797 | -13799.5801 | 3.4028e+38 | 3.4028e+38 | 229.9998 | 100.0007 | 18400.7734 | 23000.1348 | -13799.5391 | 49.9999 | 36.8678 | 0.017669 | 1.6142e-05 | 5.8396e-08 | 2.8805e-07 | -1.5999e-07 | 49.9501 | 72.7501 | 0.019695 | 1.4562e-05 | 1.7787e-07 | 2.8619e-07 | -4.7004e-10 | 49.9987 | 3.0201 | 3.4028e+38 | 3.4028e+38 | 18400.7734 | 23000.1348 | -13799.54 | 3.4028e+38 | 3.4028e+38 | 229.9998 | 100.0008 | 18400.8125 | 23000.1641 | -13799.5508 | 49.9999 | 36.8677 | 0.017879 | 1.6405e-05 | 2.8409e-08 | 2.9331e-07 | -1.6716e-07 | 50.0344 | 82.1161 | 0.019537 | 1.4404e-05 | 1.6909e-07 | 2.8142e-07 | -4.2435e-08 | 50.0195 | 14.1482 | 3.4028e+38 | 3.4028e+38 | 18400.8125 | 23000.1641 | -13799.5518 | 3.4028e+38 | 3.4028e+38 | 230.0001 | 100.0008 | 18400.8535 | 23000.1992 | -13799.499 | 49.9999 | 36.8678 | 0.017029 | 1.6523e-05 | -5.0912e-09 | 2.8088e-07 | -1.5118e-07 | 50.0351 | 92.2854 | 0.018818 | 1.4508e-05 | 1.3838e-07 | 2.6994e-07 | -6.5187e-08 | 49.9863 | 27.7464 | 3.4028e+38 | 3.4028e+38 | 18400.8516 | 23000.1992 | -13799.5 | 3.4028e+38 | 3.4028e+38 | 230.0005 | 100.0007 | 18400.8145 | 23000.1855 | -13799.5107 | 49.9999 | 36.8677 | 0.015369 | 1.6267e-05 | -1.353e-08 | 2.5002e-07 | -1.0221e-07 | 50.0414 | 97.0546 | 0.016735 | 1.4812e-05 | 8.8361e-08 | 2.4796e-07 | -5.9449e-08 | 50.0067 | 33.7985 | 3.4028e+38 | 3.4028e+38 | 18400.8125 | 23000.1855 | -13799.5117 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0004 | 100.0007 | 18400.8027 | 23000.1582 | -13799.5059 | 50 | 36.8678 | 0.80003 | 3.4028e+38 | 0 | 230.0042 | 18401.4473 | 23000.5703 | -13799.3506 | 0.044317 | 0.015231 | 1.6224e-05 | -1.038e-08 | 2.4242e-07 | -9.3575e-08 | 49.9645 | 97.1134 | -0.037633 | 3.4028e+38 | -108.1816 | 0.020199 | 0 | 0 | 0 | -0.0017794 | 0.016195 | 1.5071e-05 | 7.536e-08 | 2.3787e-07 | -4.5661e-08 | 49.9924 | 30.7367 | 0.30507 | 3.4028e+38 | 99.1044 | 230.0044 | 3.3975e-09 | 0.003464 | -6.091e-09 | -0.00024682 | 3.4028e+38 | 3.4028e+38 | 18400.7988 | 23000.1348 | -13799.4717 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 18401.1543 | 18401.1543 | -13799.248 | 0.012753 | 230.0001 | 100.0007 | 18400.8137 | 23000.1727 | -13799.5357 | 49.9999 | 36.8678 | [17 empty] | 13.8823 | 27.1917 | 50.0455 | 40.9438 | 48.8701 |  | -3.0431e-05
166 | 230.0002 | 100.0007 | 18400.8418 | 23000.1953 | -13799.4902 | 50 | 36.8676 | 0.01773 | 1.6146e-05 | 6.3127e-08 | 2.8404e-07 | -1.5549e-07 | 49.989 | 67.7467 | 0.019409 | 1.4607e-05 | 1.7759e-07 | 2.8568e-07 | 7.862e-09 | 49.9634 | -2.1557 | 3.4028e+38 | 3.4028e+38 | 18400.8418 | 23000.1953 | -13799.4912 | 3.4028e+38 | 3.4028e+38 | 230 | 100.0005 | 18400.75 | 23000.0996 | -13799.5391 | 49.9999 | 36.8679 | 0.016988 | 1.5677e-05 | 8.4627e-08 | 2.6857e-07 | -1.228e-07 | 50.0055 | 58.5221 | 0.01854 | 1.4846e-05 | 1.5078e-07 | 2.7719e-07 | 3.9712e-08 | 50.0058 | -14.21 | 3.4028e+38 | 3.4028e+38 | 18400.75 | 23000.0996 | -13799.54 | 3.4028e+38 | 3.4028e+38 | 230.0006 | 100.0006 | 18400.791 | 23000.1289 | -13799.4961 | 50 | 36.8676 | 0.015118 | 1.6302e-05 | -1.1634e-08 | 2.4437e-07 | -9.4861e-08 | 49.9324 | 96.0182 | 0.016681 | 1.4904e-05 | 8.4751e-08 | 2.4331e-07 | -5.696e-08 | 49.9735 | 32.5842 | 3.4028e+38 | 3.4028e+38 | 18400.7891 | 23000.1289 | -13799.4971 | 3.4028e+38 | 3.4028e+38 | 230.0005 | 100.0005 | 18400.7676 | 23000.1406 | -13799.5303 | 49.9999 | 36.8678 | 0.014702 | 1.5567e-05 | 5.3657e-08 | 2.2833e-07 | -5.8243e-08 | 50.0275 | 55.4563 | 0.01595 | 1.5166e-05 | 7.1046e-08 | 2.3576e-07 | 1.6876e-08 | 50.0241 | 0.50118 | 3.4028e+38 | 3.4028e+38 | 18400.7676 | 23000.1406 | -13799.5312 | 3.4028e+38 | 3.4028e+38 | 230.0006 | 100.0006 | 18400.8496 | 23000.1973 | -13799.5176 | 50 | 36.8675 | 0.014603 | 1.6103e-05 | -5.6728e-09 | 2.3796e-07 | -8.4461e-08 | 49.9692 | 94.9967 | 0.015825 | 1.5024e-05 | 7.7624e-08 | 2.4362e-07 | -4.6429e-08 | 50.0127 | 31.5521 | 3.4028e+38 | 3.4028e+38 | 18400.8477 | 23000.1973 | -13799.5186 | 3.4028e+38 | 3.4028e+38 | 123 | 230.0004 | 100.0007 | 18400.8438 | 23000.1699 | -13799.54 | 49.9999 | 36.8678 | 0.80003 | 3.4028e+38 | 0 | 230.0041 | 18401.3984 | 23000.5449 | -13799.3154 | 0.042429 | 0.014989 | 1.6148e-05 | -8.6262e-09 | 2.3907e-07 | -8.8312e-08 | 49.977 | 95.6241 | -0.026755 | 3.4028e+38 | 131.9306 | 0.020035 | 0 | 0 | 0 | -0.0017905 | 0.015916 | 1.5066e-05 | 7.5016e-08 | 2.4228e-07 | -4.3088e-08 | 49.9938 | 28.1692 | 0.31483 | 3.4028e+38 | -114.3422 | 229.7547 | 8.5201e-09 | 0.0034713 | -5.7197e-09 | -0.00025617 | 3.4028e+38 | 3.4028e+38 | 18400.8516 | 23000.1777 | -13799.4395 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 3.4028e+38 | 18401.166 | 18401.168 | -13799.2646 | 0.0051698 | 230.0004 | 100.0006 | 18400.8 | 23000.1523 | -13799.5146 | 50 | 36.8677 | [17 empty] | 16.2255 | 27.045 | 51.5765 | 42.3626 | 49.6954 |  | -0.00012361

## K2006 Data (hidden)
Reading 1 | [27 empty] | Reading 2 | [27 empty] | Reading 3 | [27 empty] | Reading 4 | [27 empty] | Reading 5 | [27 empty] | Read out of system status (SG command)
Phase Currents |  |  | Phase to Neutral Voltages |  |  | Active Power per Phase |  |  | Reactive Power per Phase |  |  | Apparent Power per Phase |  |  | Active Sum-Power | Reactive Sum-Power | Apparent Sum-power | Phase Angle Current to Voltage |  |  | Signal Freq | Voltage Range |  |  | Current Range |  |  | Phase Currents |  |  | Phase to Neutral Voltages |  |  | Active Power per Phase |  |  | Reactive Power per Phase |  |  | Apparent Power per Phase |  |  | Active Sum-Power | Reactive Sum-Power | Apparent Sum-power | Phase Angle Current to Voltage |  |  | Signal Freq | Voltage Range |  |  | Current Range |  |  | Phase Currents |  |  | Phase to Neutral Voltages |  |  | Active Power per Phase |  |  | Reactive Power per Phase |  |  | Apparent Power per Phase |  |  | Active Sum-Power | Reactive Sum-Power | Apparent Sum-power | Phase Angle Current to Voltage |  |  | Signal Freq | Voltage Range |  |  | Current Range |  |  | Phase Currents |  |  | Phase to Neutral Voltages |  |  | Active Power per Phase |  |  | Reactive Power per Phase |  |  | Apparent Power per Phase |  |  | Active Sum-Power | Reactive Sum-Power | Apparent Sum-power | Phase Angle Current to Voltage |  |  | Signal Freq | Voltage Range |  |  | Current Range |  |  | Phase Currents |  |  | Phase to Neutral Voltages |  |  | Active Power per Phase |  |  | Reactive Power per Phase |  |  | Apparent Power per Phase |  |  | Active Sum-Power | Reactive Sum-Power | Apparent Sum-power | Phase Angle Current to Voltage |  |  | Signal Freq | Voltage Range |  |  | Current Range |  |  | Current Range | Voltage Range | Operating Mode | Meter Constant of output1 | Manual/ Auto Range | Time Base (s) | Time Base Error | DC Measurement (0 = OFF) | Setting of Current Inputs | Setting of Voltage Inputs
I1 | I2 | I3 | U1 | U2 | U3 | P1 | P2 | P3 | Q1 | Q2 | Q3 | S1 | S2 | S3 | ΣP | ΣQ | ΣS | φiu1 | φiu2 | φiu3 |  | U1 | U2 | U3 | I1 | I2 | I3 | I1 | I2 | I3 | U1 | U2 | U3 | P1 | P2 | P3 | Q1 | Q2 | Q3 | S1 | S2 | S3 | ΣP | ΣQ | ΣS | φiu1 | φiu2 | φiu3 |  | U1 | U2 | U3 | I1 | I2 | I3 | I1 | I2 | I3 | U1 | U2 | U3 | P1 | P2 | P3 | Q1 | Q2 | Q3 | S1 | S2 | S3 | ΣP | ΣQ | ΣS | φiu1 | φiu2 | φiu3 |  | U1 | U2 | U3 | I1 | I2 | I3 | I1 | I2 | I3 | U1 | U2 | U3 | P1 | P2 | P3 | Q1 | Q2 | Q3 | S1 | S2 | S3 | ΣP | ΣQ | ΣS | φiu1 | φiu2 | φiu3 |  | U1 | U2 | U3 | I1 | I2 | I3 | I1 | I2 | I3 | U1 | U2 | U3 | P1 | P2 | P3 | Q1 | Q2 | Q3 | S1 | S2 | S3 | ΣP | ΣQ | ΣS | φiu1 | φiu2 | φiu3 |  | U1 | U2 | U3 | I1 | I2 | I3

## PowerCalc
[26 empty] | #REF! | [20 empty] | #REF! | [20 empty] | #REF! | [41 empty] | #REF!
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80 | 81 | 82 | 83 | 84 | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 117 | 118 | 119 | 120 | 121 | 122 | 123 | 124 | 125 | 126 | 127 | 128 | 129 | 130 | 131 | 132 | 133 | 134 | 135 | 136 | 137 | 138 | 139 | 140 | 141 | 142 | 143 | 144 | 145 | 146 | 147 | 148 | 149 | 150 | 151 | 152 | 153 | 154 | 155 | 156 | 157 | 158 | 159 | 160 | 161 | 162 | 163 | 164 | 165 | 166 | 167 | 168 | 169 | 170 | 171 | 172 | 173 | 174 | 175 | 176 | 177 | 178 | 179 | 180 | 181 | 182
|  | 11 | 7 | 6 | 8 | 29 | 30 | 31 | 32 | 28 | 14 | 15 | 35 | [12 empty] | 33 | 34 | [6 empty] | 42 | [12 empty] | 40 | 41 | [6 empty] | 49 | [12 empty] | 47 | 48 | [6 empty] | 56 | [12 empty] | 54 | 55 | [6 empty] | 63 | [12 empty] | 61 | 62 |  |  |  |  | 13 | 12 | 15 | 38 | 45 | 52 | 59 | 66 | 39 | 46 | 53 | 60 | 67 | [41 empty] | 38 | 45 | 52 | 59 | 66 |  |  |  |  | 96
[27 empty] | 37 | [20 empty] | 44 | [20 empty] | 51 | [20 empty] | 58 | [20 empty] | 65
Reading Number |  | Number of Readings | Shunt | Divider Range | Current Transformer Ratio | Wdail | Wsign | Vdail | Vsign | ExactFreq | Divider ratio K | Divider ratio K | Ref Volts V1(cor'd) | Shunt impedance R3 | Shunt impedance R3 | Watts dial alpha | Watts buffer gain | R2/R1 | R2/R1 | Vars dial beta | Vars buffer gain | g_1 conductance | Y1 admittance | R2 impedance | R2 impedance | Detector amp volts | Detector phase | Det Amp Gain G | Det Amp phase Gj | Det amp conductance g_d | Det amp admittance Yd | Divider ratio K | Divider ratio K | Ref Volts V1(cor'd) | Shunt impedance R3 | Shunt impedance R3 | Watts dial alpha | Watts buffer gain | R2/R1 | R2/R1 | Vars dial beta | Vars buffer gain | g_1 conductance | Y1 admittance | R2 impedance | R2 impedance | Detector amp volts | Detector phase | Det Amp Gain G | Det Amp phase Gj | Det amp conductance g_d | Det amp admittance Yd | Divider ratio K | Divider ratio K | Ref Volts V1(cor'd) | Shunt impedance R3 | Shunt impedance R3 | Watts dial alpha | Watts buffer gain | R2/R1 | R2/R1 | Vars dial beta | Vars buffer gain | g_1 conductance | Y1 admittance | R2 impedance | R2 impedance | Detector amp volts | Detector phase | Det Amp Gain G | Det Amp phase Gj | Det amp conductance g_d | Det amp admittance Yd | Divider ratio K | Divider ratio K | Ref Volts V1(cor'd) | Shunt impedance R3 | Shunt impedance R3 | Watts dial alpha | Watts buffer gain | R2/R1 | R2/R1 | Vars dial beta | Vars buffer gain | g_1 conductance | Y1 admittance | R2 impedance | R2 impedance | Detector amp volts | Detector phase | Det Amp Gain G | Det Amp phase Gj | Det amp conductance g_d | Det amp admittance Yd | Divider ratio K | Divider ratio K | Ref Volts V1(cor'd) | Shunt impedance R3 | Shunt impedance R3 | Watts dial alpha | Watts buffer gain | R2/R1 | R2/R1 | Vars dial beta | Vars buffer gain | g_1 conductance | Y1 admittance | R2 impedance | R2 impedance | Detector amp volts | Detector phase | Det Amp Gain G | Det Amp phase Gj | Det amp conductance g_d | Det amp admittance Yd |  |  | Parameter: | Ch1 Imp/Wh | Ch1 Imp/Wh | Ch1 Imp/Wh | Ch1 Imp/Wh | Ch1 Imp/Wh | Watt | Watt | Watt | Watt | Watt | Var | Var | Var | Var | Var | Volts | Volts | Volts | Volts | Volts | Amps | Amps | Amps | Amps | Amps | Degrees (Non-IEC) | Degrees (Non-IEC) | Degrees (Non-IEC) | Degrees (Non-IEC) | Degrees (Non-IEC) | Impulses/w-hr or var-hr |  |  |  |  | Non-IEC | Impulses/w-hr or var-hr | [5 empty] | Error (%) Impulses/w-hr or var-hr | [8 empty] | Ch1 Imp/var-hr | Ch1 Imp/var-hr | Ch1 Imp/var-hr | Ch1 Imp/var-hr | Ch1 Imp/var-hr | Impulses per Var-hr | Impulses per Watt-hr | Var-sec per impulse std dev (%) | Watt-sec per impulse std dev (%) | Pt Resistance (Ohms) | Temperature (°C) | Corrected C1 (uF)
[11 empty] | Re | Im | volts | ohms, Re | Ohms, Im |  |  | Re | Im |  |  | siemens Re | siemens Im | ohms Re | Ohms Im | volts | radians |  | radians | siemens Re | siemens Im | Re | Im | volts | ohms, Re | Ohms, Im |  |  | Re | Im |  |  | siemens Re | siemens Im | ohms Re | Ohms Im | volts | radians |  | radians | siemens Re | siemens Im | Re | Im | volts | ohms, Re | Ohms, Im |  |  | Re | Im |  |  | siemens Re | siemens Im | ohms Re | Ohms Im | volts | radians |  | radians | siemens Re | siemens Im | Re | Im | volts | ohms, Re | Ohms, Im |  |  | Re | Im |  |  | siemens Re | siemens Im | ohms Re | Ohms Im | volts | radians |  | radians | siemens Re | siemens Im | Re | Im | volts | ohms, Re | Ohms, Im |  |  | Re | Im |  |  | siemens Re | siemens Im | ohms Re | Ohms Im | volts | radians |  | radians | siemens Re | siemens Im | watts or vars | imp/w-hr | Reading: | 1 | 2 | 3 | 4 | 5 | 1 | 2 | 3 | 4 | 5 | 1 | 2 | 3 | 4 | 5 | 1 | 2 | 3 | 4 | 5 | 1 | 2 | 3 | 4 | 5 | 1 | 2 | 3 | 4 | 5 | MeanImp/W-hr | MeanWatts | MeanVars | Meanvolts | MeanAmps | MeanDeg | std dev (%) | Watts_stdev | Var_stdev | Volts_stdev | Amps_stdev | Phase_stdev | ErrorPerCent | PF/VF | Mean_ratio | Ratio_stdev
166 |  | 5 | ESI(1) | 230 | 100 | 0.70605 | WP- | 0.52148 | VP- | 50 | 202.9816 | 0.021862 | 1.1331 | 0.99999 | 1.0367e-05 | 0.70605 | -1.0003 | 1 | -1.8848e-07 | 0.52148 | -1.0003 | 1.8311e-08 | 0.00010145 | 10000.0599 | -0.030159 | 0.3135 | -4.3437 | 1135.2724 | -0.045847 | 1e-07 | 6.2832e-08 | 202.9816 | 0.021862 | 1.1331 | 0.99999 | 1.0367e-05 | 0.70605 | -1.0003 | 1 | -1.8848e-07 | 0.52148 | -1.0003 | 1.8311e-08 | 0.00010145 | 10000.0599 | -0.030159 | 0.31258 | -4.3327 | 1135.2724 | -0.045847 | 1e-07 | 6.2832e-08 | 202.9816 | 0.021862 | 1.1331 | 0.99999 | 1.0367e-05 | 0.70605 | -1.0003 | 1 | -1.8848e-07 | 0.52148 | -1.0003 | 1.8311e-08 | 0.00010145 | 10000.0599 | -0.030159 | 0.3136 | -4.3463 | 1135.2724 | -0.045847 | 1e-07 | 6.2832e-08 | 202.9816 | 0.021862 | 1.1331 | 0.99999 | 1.0367e-05 | 0.70605 | -1.0003 | 1 | -1.8848e-07 | 0.52148 | -1.0003 | 1.8311e-08 | 0.00010145 | 10000.0599 | -0.030159 | 0.31248 | -4.3388 | 1135.2724 | -0.045847 | 1e-07 | 6.2832e-08 | 202.9816 | 0.021862 | 1.1331 | 0.99999 | 1.0367e-05 | 0.70605 | -1.0003 | 1 | -1.8848e-07 | 0.52148 | -1.0003 | 1.8311e-08 | 0.00010145 | 10000.0599 | -0.030159 | 0.31357 | -4.3403 | 1135.2724 | -0.045847 | 1e-07 | 6.2832e-08 | watts | 100000 |  | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | 109.115 | 23.5249 | 0.32293
The row below is included so that this spreadsheet can be "livened up". This is necessary if you need to change formula, constants, or uncertainties after making the measurements.
12 |  | 5 | C(.2) | 230 | 1 | 0.35254 | WP- | 0 | VP+ | 50.0002 | 202.9816 | 0.021862 | 1.1331 | 0.19999 | 0 | 0.35254 | -1.0003 | 1 | -1.8848e-07 | 0 | 1 | 1.8311e-08 | 0.00010145 | 10000.0599 | -0.030159 | 0.21171 | -0.42403 | 1135.2724 | -0.045847 | 1e-07 | 6.2832e-08 | 202.9816 | 0.021862 | 1.1331 | 0.19999 | 0 | 0.35254 | -1.0003 | 1 | -1.8848e-07 | 0 | 1 | 1.8311e-08 | 0.00010145 | 10000.0599 | -0.030159 | 0.2138 | -0.4132 | 1135.2724 | -0.045847 | 1e-07 | 6.2832e-08 | 202.9816 | 0.021862 | 1.1331 | 0.19999 | 0 | 0.35254 | -1.0003 | 1 | -1.8848e-07 | 0 | 1 | 1.8311e-08 | 0.00010145 | 10000.0599 | -0.030159 | 0.21233 | -0.42817 | 1135.2724 | -0.045847 | 1e-07 | 6.2832e-08 | 202.9816 | 0.021862 | 1.1331 | 0.19999 | 0 | 0.35254 | -1.0003 | 1 | -1.8848e-07 | 0 | 1 | 1.8311e-08 | 0.00010145 | 10000.0599 | -0.030159 | 0.21147 | -0.41482 | 1135.2724 | -0.045847 | 1e-07 | 6.2832e-08 | 202.9816 | 0.021862 | 1.1331 | 0.19999 | 0 | 0.35254 | -1.0003 | 1 | -1.8848e-07 | 0 | 1 | 1.8311e-08 | 0.00010145 | 10000.0599 | -0.030159 | 0.21563 | -0.41147 | 1135.2724 | -0.045847 | 1e-07 | 6.2832e-08 | watts | 100000 |  | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | 109.147 | 23.6074 | 0.32293
Note on FG9: Industrial Research Limited: In version l and onwards, here we reverse the sign of the angle to allow for the European convention
Note on FX10: Industrial Research Limited: Checked Laurie 6/4/2006
Note on FX12: Industrial Research Limited: Checked Laurie 6/4/2006

## UncCalc
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63
Lookup: | 8 | 11 | 138 | 139 | 12 | 12 | 14 | 15 | 15 | 17 | 18 | 19 | 19 | 21 | 22 | 24 | 24 | 25 | 25 | 27 | 28 | 29 | 30 | 31 | 32 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32
|  |  | Parameter Number: |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
|  | Uncertainty (ppm if not urad): |  |  | 5 | 5 | 4.3566 | 10 | 1 | 2 | 2 | 5 | 1 | 2 | 1 | 15 | 50 | 2 | 2 | 500 | 873 | 500 | 500 | 10000 | 10000
|  |  | check_std_array | 1 | 2: urad; 5: urad; 9: urad; 15: urad; 17: urad; 19: urad | [21 empty] | Watts | [7 empty] | Vars
|  |  | Degrees of Freedom: |  | 6 | 12.5 | 50 | 10 | 50 | 8 | 8 | 8 | 8 | 8 | 8 | 12.5 | 12.5 | 8 | 8 | 12.5 | 8 | 8 | 8 | 8 | 8 | WATTS ARRAY
[5 empty] | 1: Standard Deviations Type B; 3: STANDARD DEVIATION ARRAY | Divider ratio K | Divider ratio K | Ref Volts V1(cor'd) | Shunt impedance R3 | Shunt impedance R3 | Watts dial alpha | Watts buffer gain | R2/R1 | R2/R1 | Vars dial beta | Vars buffer gain | g_1 conductance | Y1 admittance | R2 impedance | R2 impedance | Detector amp volts | Detector phase | Det Amp Gain G | Det Amp phase Gj | Det amp conductance g_d | Det amp admittance Yd | std dev |  |  | degrees of freedom |  |  | IPULSES/WATT-HOUR |  | std dev |  |  | degrees of freedom |  |  | IPULSES/VAR-HOUR
| Current Transformer Ratio | deg of freedom |  |  | Divider ratio K Re | Divider ratio K Im | Ref Volts V1(cor'd) | Shunt impedance R3 Re | Shunt impedance R3 Im | Watts dial alpha | Watts buffer gain | R2/R1 Re | R2/R1 Im | Vars dial beta | Vars buffer gain | g_1 conductance | Y1 admittance | R2 impedance Re | R2 impedance Im | Detector amp volts | Detector phase | Det Amp Gain G | Det Amp phase Gj | Det amp conductance g_d | Det amp admittance Yd | Re | Im | volts | ohms, Re | Ohms, Im |  |  | Re | Im |  |  | siemens Re | siemens Im | ohms Re | Ohms Im | volts | radians |  | radians | siemens Re | siemens Im | Ch 1 Type A watt-sec/imp std unc (%) | Ch 1 Type B std unc % from mfr spec | PB Type B std unc % | Ch 1 Type A watt-sec/imp dfm | Ch 1 Type B dfm | PB Type B watts dfm | imp/watt-hr combined std unc % | imp/w-hr dfm | Ch 1 Type A var-sec/imp std unc % | Ch 1 Type B std unc % from mfr spec | PB Type B std unc % | Ch 1 Type A var-sec/imp dfm | Ch 1 Type B dfm | PB Type B var dfm | imp/watt-hr combined std unc % | imp/w-hr dfm
166 | 100 | 5 |  |  | 0.0010149 | 0.0010149 | 4.9364e-06 | 9.9999e-06 | 9.9999e-07 | 2.9336e-06 | -2.0007e-06 | 5e-06 | 1e-06 | 2.749e-06 | -1.0003e-06 | 1.5218e-09 | 5.0726e-09 | 0.02 | 0.02 | 0.00015675 | 0.000873 | 0.56764 | 0.0005 | 1e-09 | 6.2832e-10 | 202.9816 | 0.021862 | 1.1331 | 0.99999 | 1.0367e-05 | 0.70605 | -1.0003 | 1 | -1.8848e-07 | 0.52148 | -1.0003 | 1.8311e-08 | 0.00010145 | 10000.0599 | -0.030159 | 0.3135 | -4.3437 | 1135.2724 | -0.045847 | 1e-07 | 6.2832e-08 | #NAME? | 0.0037302 | #NAME? | 4 | 6 | #NAME? | #NAME? | #NAME? | #NAME? | 0.0037302 | #NAME? | 4 | 6 | #NAME? | #NAME? | #NAME?
The row below is included so that this spreadsheet can be "livened up". This is necessary if you need to change formula, constants, or uncertainties after making the measurements.
12 | 1 | 5 |  |  | 0.0010149 | 0.0010149 | 4.9364e-06 | 1.9999e-06 | 1.9999e-07 | 1.7051e-06 | -2.0007e-06 | 5e-06 | 1e-06 | 1.3525e-06 | 1e-06 | 1.5218e-09 | 5.0725e-09 | 0.02 | 0.02 | 0.00010585 | 0.000873 | 0.56764 | 0.0005 | 1e-09 | 6.2832e-10 | 202.9816 | 0.021862 | 1.1331 | 0.19999 | 0 | 0.35254 | -1.0003 | 1 | -1.8848e-07 | 0 | 1 | 1.8311e-08 | 0.00010145 | 10000.0599 | -0.030159 | 0.21171 | -0.42403 | 1135.2724 | -0.045847 | 1e-07 | 6.2832e-08 | #NAME? | 0.0006 | #NAME? | 4 | 6 | #NAME? | #NAME? | #NAME? | #NAME? | 0.0006 | #NAME? | 4 | 6 | #NAME? | #NAME? | #NAME?
Note on Q3: Tom Stewart: Changed from 23 to 24. As the uncertainty in phase angle (loss angle) of C1 is 15 urad ie 15 ppm of the susceptance (not conductance).
Note on E6: Laurie Christian: This is for checking to see that the correct PowerCalc column is being referred to.

## constants
TABLE OF CONSTANTS |  |  | !! When copying new constants in, ensure the named cells and lookups from other worksheets stay intact! |  |  |  |  | !! When copying new constants in, ensure the named cells and lookups from other worksheets stay intact! |  |  |  |  | !! When copying new constants in, ensure the named cells and lookups from other worksheets stay intact!
|  |  | Feb-May 2014 Calibrations carried out by T. Stewart |  |  |  |  | Feb-May 2014 Calibrations carried out by T. Stewart |  |  |  |  | Apr/May 2012 Calibrations carried out by T. Stewart |  |  |  |  | Dec 10 Calibrations carried out by T. Stewart |  |  |  |  | Dec 08/Jan 09 Calibrations carried out by T. Stewart |  |  |  |  | April/May 08 Calibrations carried out by B. McLennan and T. Stewart |  |  |  |  | Calibration data from the August/November 2005/February 2006 Calibration (CRM) (post-grounding fixes) |  |  |  | Laurie checked this in June 2004. Recheck by June 30, 2005 or earlier if required. | [9 empty] | values prior | values prior July 2000
|  | Cal Date: |  |  |  |  | Cal Date: |  |  |  |  | Cal Date: |  |  |  |  | Cal Date: | 2010-12-01 |  |  |  | Cal Date: | Dec2008 to Jan 2008 |  |  |  | Cal Date: | April 2008 to May 2008 |  |  |  |  | Aug 2005 to March 2006 |  |  |  | 2004-04-01 | [7 empty] | Phase values for 24/4/97 |  | July 2000 | PhaseOffset | PhOffset,G=1
|  | Cal Ref: | CalPB_Feb14.xls |  | Colour Key |  | Cal Ref: | CalPB_Feb14.xls |  | Colour Key |  | Cal Ref: | CalPB_Apr12.xls |  | Colour Key |  | Cal Ref: | CalPB_Dec10.xls |  | Colour Key |  | Cal Ref: | CalPB_Dec08 updatedLaurieChecked.xls |  | Colour Key |  | Cal Ref: | CalPB_April08 updated Laurie.xls |  |  |  |  | CalPB_August05 updated |  |  |  | CalPBApr04
Comment | Unit | Name |  |  | Tan | TJS updated | Name |  |  | Tan | TJS updated | Name |  |  | L Green | TJS updated | Name |  |  | Cyan | TJS Checked | Name |  |  | Cyan | BEM/TJS Checked | Name | Laurie Christian
53 Hz value, close to 50 Hz val |  | DetAmpGain | 1135.2724 |  | Blue | TJS Checked | DetAmpGain | 1135.2724 |  | Blue | TJS Checked | DetAmpGain | 1135.269 |  | Aqua | Checked | DetAmpGain | 1135.2764 |  | Green | LAC and TJS Checked | DetAmpGain | 1135.3276 |  | Green | LAC Checked | DetAmpGain | 1134.9985 |  |  |  |  | 1134.954 | Supporting Information\CalPB_August05 updated.xls |  |  | 1136.1647 | CalPBApr04 | -785.3344 | 1133.8 |  |  |  |  | -2.4706 | -2.65 | 1134.6 | -2.5384 | -2.5384
|  | HP3458corectn | 1 | -4.3823 | Lavender | needs updating | HP3458corectn | 1 | -4.3823 | Lavender | needs updating | HP3458corectn | 1 |  | Lavender | needs updating | HP3458corectn | 1 |  | Lavender | Will need updating | HP3458corectn | 1 |  | Lavender | Will need updating | HP3458corectn | 1 |  |  |  |  | 1 |  |  |  | 1 |  |  | 1 |  |  |  |  | -2.522 | -2.65 | 0.99999 | -2.589 | -2.589
| ohms | R_1 | 9999.9989 |  | Yellow diag stripes |  | R_1 | 9999.9989 |  | Yellow diag stripes |  | R_1 | 10000.0209 |  | Yellow diag stripes |  | R_1 | 10000.0216 |  | Yellow diag stripes | BEM checked | R_1 | 10000.0062 |  |  |  | R_1 | 9999.9745 |  |  |  |  | 10000.0137 | Supporting Information\CalPB_August05 updated.xls |  |  | 10000.0375 | CalPBApr04 |  | 10001.2928 | from PBCAL.xlsJULY 2000 |  |  |  | -2.5734 | -2.65 | 10001.2017 | -2.6396 | -2.6396
| henries | L_1 | -9e-05 | See WB1 p88-89. |  |  | L_1 | -9e-05 | See WB1 p88-89. |  |  | L_1 | -9e-05 | See WB1 p88-89. |  |  | L_1 | -9e-05 | See WB1 p88-89. |  |  | L_1 | -9e-05 | See WB1 p88-89. |  |  | L_1 | -9e-05 | See WB1 p88-89. |  |  |  | -9e-05 | See WB1 p88-89. |  |  | -3.1e-05 |  |  | -3.1e-05 |  |  |  |  | -2.6248 | -2.65 | -3.1e-05 | -2.6902 | -2.6902
this resistor has large temp co | ohms | R_2 | 10000.0599 |  |  |  | R_2 | 10000.0599 |  |  |  | R_2 | 10000.079 |  |  |  | R_2 | 10000.1001 |  |  |  | R_2 | 10000.092 |  |  |  | R_2 | 10000.0585 |  |  |  |  | 10000.1086 | Supporting Information\CalPB_August05 updated.xls |  |  | 10000.134 | CalPBApr04 |  | 9998.2975 |  |  |  |  | -2.6762 | -2.65 | 9998.2163 | -2.7408 | -2.7408
| henries | L_2 | -9.6e-05 | See WB1 p88-89. |  |  | L_2 | -9.6e-05 | See WB1 p88-89. |  |  | L_2 | -9.6e-05 | See WB1 p88-89. |  |  | L_2 | -9.6e-05 | See WB1 p88-89. |  |  | L_2 | -9.6e-05 | See WB1 p88-89. |  |  | L_2 | -9.6e-05 | See WB1 p88-89. |  |  |  | -9.6e-05 | See WB1 p88-89. |  |  | -2.5e-05 |  |  | -2.5e-05 |  |  |  |  | -2.7276 | -2.65 | -2.5e-05 | -2.7914 | -2.7914
|  | R_2byR_1 | 1 | R_2byR_1 = R_2/R_1 |  |  | R_2byR_1 | 1 | R_2byR_1 = R_2/R_1 |  |  | R_2byR_1 | 1 | R_2byR_1 = R_2/R_1 |  |  | R_2byR_1 | 1 | R_2byR_1 = R_2/R_1 |  |  | R_2byR_1 | 1 | R_2byR_1 = R_2/R_1 |  |  | R_2byR_1 | 1 | R_2byR_1 = R_2/R_1 |  |  |  | 1 |  |  |  | 1 | CalPBApr04 |  | 0.99972 | Rtk2000a_check.xls cell EP351 |  |  |  | -2.779 | -2.65 | 0.9997 | -2.842 | -2.842
| seconds | TCR2_R1 | -5.9995e-10 | TCR2_R1=(L_2*R_1-L_1*R_2)/(R_1*R_1) |  |  | TCR2_R1 | -5.9995e-10 | TCR2_R1=(L_2*R_1-L_1*R_2)/(R_1*R_1) |  |  | TCR2_R1 | -5.9992e-10 | TCR2_R1=(L_2*R_1-L_1*R_2)/(R_1*R_1) |  |  | TCR2_R1 | -5.9993e-10 | TCR2_R1=(L_2*R_1-L_1*R_2)/(R_1*R_1) |  |  | TCR2_R1 | -5.9992e-10 | TCR2_R1=(L_2*R_1-L_1*R_2)/(R_1*R_1) |  |  | TCR2_R1 | -5.9993e-10 | TCR2_R1=(L_2*R_1-L_1*R_2)/(R_1*R_1) |  |  |  | -5.9991e-10 | See Callum McKenzie, Rough Notes and Ideas No1 p22 |  |  | 6e-10 |  | need to add commentary | -6e-10 |  |  |  |  | -2.8304 | -2.65 | 6e-10 | -2.8926 | -2.8926
value at temp of D18 | uF | C_1 | 0.32293 | From "CalPb_C1_Jan2013.xls" |  |  | C_1 | 0.32293 | From "CalPb_C1_Jan2013.xls" |  |  | C_1 | 0.32291 | From "CalPb_C1_Jan2013.xls" |  |  | C_1 | 0.32288 | From "CalPb_C1_Dec2010.xls" |  |  | C_1 | 0.32287 | From "CalPb_C1_May2008.xls" |  |  | C_1 | 0.32285 | From "CalPb_C1_May2008.xls" |  |  |  | 0.32284 | Supporting Information\CalPB_C1_May2006.xls |  |  | 0.32288 | Pb_Auto_V1_cal04.xls |  | 0.32287 | correction of -221 ppm from Rtk2000a_check.xls block EQ366 |  |  |  | -2.8818 | -2.65 | 0.3229 | -2.9432 | -2.9432
| S | g_1 | 1.8311e-08 | From "KJ Copy of CalPB_C1_May2012.xls" |  |  | g_1 | 1.8311e-08 | From "KJ Copy of CalPB_C1_May2012.xls" |  |  | g_1 | 1.8311e-08 | From "KJ Copy of CalPB_C1_May2012.xls" |  |  | g_1 | 1.7508e-08 | Unchanged |  |  | g_1 | 1.7508e-08 | Unchanged |  |  | g_1 | 1.7508e-08 | See New Cal2.xls where it is shown that is not critical |  |  |  | 1.7508e-08 | See New Cal2.xls where it is shown that is not critical |  |  | 1.7508e-08 |  | needs updating | 1.7508e-08 | Rtk2000a_check.xls |  |  |  | -2.9332 | -2.65 | 1.68e-08 | -2.9938 | -2.9938
det amp Cin | pF | C_d | 200 | Unchanged |  |  | C_d | 200 | Unchanged |  |  | C_d | 200 | Unchanged |  |  | C_d | 200 | Unchanged |  |  | C_d | 200 | Unchanged |  |  | C_d | 200 | needs updating |  |  |  | 200 | needs updating |  |  | 200 |  | needs updating | 200 | sheet Calculations! cell ER341 |  |  |  | -2.9846 | -2.65 | 200 | -3.0444 | -3.0444
det amp Gin | S | g_d | 1e-07 | Unchanged |  |  | g_d | 1e-07 | Unchanged |  |  | g_d | 1e-07 | Unchanged |  |  | g_d | 1e-07 | Unchanged |  |  | g_d | 1e-07 | Unchanged |  |  | g_d | 1e-07 | needs updating |  |  |  | 1e-07 | needs updating |  |  | 1e-07 |  | needs updating | 1e-07 | [6 empty] | 1e-07
C1 temp at calibration | deg C | C_1_CalTemp | 23.5113 |  |  |  | C_1_CalTemp | 23.5113 | [8 empty] | Detector Gain Lookup |  |  |  |  | Detector Gain Lookup |  |  |  |  | Detector Gain Lookup
Detector Gain Lookup |  | Min | 0.99981 |  |  |  | Min | 0.99981 |  |  |  | Min | 0.99984 |  |  |  | Min | 0.99983 |  |  |  | Min | 0.99987 |  |  |  | Min | 0.99982 |  |  |  |  | 0.99986 | Supporting Information\CalPB_August05 updated.xls |  |  | 0.99984 | CalPBApr04 |  | 1
|  | Max | 1135.2724 |  |  |  | Max | 1135.2724 |  |  |  | Max | 1135.269 |  |  |  | Max | 1135.2764 |  |  |  | Max | 1135.3276 |  |  |  | Max | 1134.9985 |  |  |  |  | 1134.954 | Supporting Information\CalPB_August05 updated.xls |  |  | 1135.2724 | CalPBApr04 |  | 1135.2724
|  | LOOKUP TABLE OF SHUNT VALUES |  |  |  |  | LOOKUP TABLE OF SHUNT VALUES |  |  |  |  | LOOKUP TABLE OF SHUNT VALUES |  |  |  |  | LOOKUP TABLE OF SHUNT VALUES |  |  |  |  | LOOKUP TABLE OF SHUNT VALUES |  |  |  |  | LOOKUP TABLE OF SHUNT VALUES
ESI 1 ohm box parallel |  | ESI(.1) | 0.099993 | Needs updating but not typically used |  | Not changed | ESI(.1) | 0.099993 | Needs updating but not typically used |  | Not changed | ESI(.1) | 0.099993 | 0 |  | Not changed | ESI(.1) | 0.099993 | 0 |  | Not changed | ESI(.1) | 0.099993 | 0 |  | Not changed | ESI(.1) | 0.099993 | 0 | See ESI_History.xls Nov 2002 | Not changed |  | 0.099994 | 0 | See ESI_History.xls Nov 2002 | Not changed | 0.099993 | 0 | See ESI_History.xls Nov 2002 | resistance | inductance
| ohms, uH | C(.2) | 0.19999 | 0 |  |  | C(.2) | 0.19999 | 0 |  |  | C(.2) | 0.19999 | 0 |  |  | C(.2) | 0.19999 | 0 |  |  | C(.2) | 0.19999 | 0 |  |  | C(.2) | 0.19999 | 0 |  |  |  | 0.19999 | 0 |  |  | 0.19999 | 0 |  | 0.19999 | 0
| ohms, uH | Ti651 | Not used - This need updating | 0.215 |  |  | Ti651 | Not used - This need updating | 0.215 |  |  | Ti651 | 1.7077 | 0.215 |  |  | Ti651 | 1 | 0.215 |  |  | T(51) | 1 | 0.215 |  |  | T(51) | 1 | 0.215 |  |  |  | 1 | 0.215 |  |  | 1 | 0.215 |  | 1 | 0.215
14 ppm added to allow for increase in R at 53 Hz over DC | ohms, uH | Ti(122) | Not used - This need updating | 0.076 | Note: this shunt has large temperature coefficient and self-heating so the 20 C value is current dependent. Probably better to apply current and ambient temperature correction rather than a single oil temperature correction. |  | Ti(122) | Not used - This need updating | 0.076 | Note: this shunt has large temperature coefficient and self-heating so the 20 C value is current dependent. Probably better to apply current and ambient temperature correction rather than a single oil temperature correction. |  | Ti(122) | This need updating | 0.076 | Note: this shunt has large temperature coefficient and self-heating so the 20 C value is current dependent. Probably better to apply current and ambient temperature correction rather than a single oil temperature correction. |  | T(122) | This need updating | 0.076 | Note: this shunt has large temperature coefficient and self-heating so the 20 C value is current dependent. Probably better to apply current and ambient temperature correction rather than a single oil temperature correction. |  | T(122) | This need updating | 0.076 | Note: this shunt has large temperature coefficient and self-heating so the 20 C value is current dependent. Probably better to apply current and ambient temperature correction rather than a single oil temperature correction. |  | T(122) | 0.099998 | 0.076 | Note: this shunt has large temperature coefficient and self-heating so the 20 C value is current dependent. Probably better to apply current and ambient temperature correction rather than a single oil temperature correction. |  |  | 0.1 | 0.076 | Note: this shunt has large temperature coefficient and self-heating so the 20 C value is current dependent. Probably better to apply current and ambient temperature correction rather than a single oil temperature correction. |  | 0.099997 | 0.076 |  | 0.1 | 0.076
| ohms, uH | Ti(707) | Not used - This need updating | 0.96 |  |  | Ti(707) | Not used - This need updating | 0.96 |  |  | Ti(707) | -10.8714 | 0.96 |  |  | T(707) | 9.9999 | 0.96 |  |  | T(707) | 9.9999 | 0.96 |  |  | T(707) | 9.9999 | 0.96 |  |  | T(924) | 10 | 0.96 |  |  | 10 | 0.96 |  | 10 | 0.96
ESI 10 ohm box parallel | ohms, uH | ESI(1) | 0.99999 | 0.033 | 10.35 urad | 0.032945 | ESI(1) | 0.99999 | 0.033 | 10.35 urad | 0.032945 | ESI(1) | 0.99999 | 0.033 |  |  | ESI(1) | 0.99999 | 0.033 |  |  | ESI(1) | 0.99999 | 0.033 |  |  | ESI(1) | 0.99999 | 0.033 |  |  |  | 0.99999 | 0.033 |  |  | 0.99999 | 0.033 |  | 0.99999 | 0.033
ESI 10 ohm box series/parallel | ohms, uH | ESI(10) | 10 | 0.33 | 10.35 urad | 0.32945 | ESI(10) | 10 | 0.33 | 10.35 urad | 0.32945 | ESI(10) | 10 | 0.33 |  |  | ESI(10) | 10 | 0.33 |  |  | ESI(10) | 9.9999 | 0.33 |  |  | ESI(10) | 9.9999 | 0.33 |  |  |  | 9.9999 | 0.33 |  |  | 9.9999 | 0.33 |  | 9.9999 | 0.33
|  | NoShunt | 1 |  |  |  | NoShunt | 1 |  |  |  | NoShunt | 1 |  | 0.19999 |  | NoShunt | 1 |  |  |  | NoShunt | 1 |  |  |  | NoShunt | 1 |  |  |  |  | 1 |  |  |  | 1
|  | LOOKUP TABLE OF DIVIDER BUFFER AMP GAIN |  |  |  |  | LOOKUP TABLE OF DIVIDER BUFFER AMP GAIN |  |  |  |  | LOOKUP TABLE OF DIVIDER BUFFER AMP GAIN |  |  |  |  | LOOKUP TABLE OF DIVIDER BUFFER AMP GAIN |  |  |  |  | LOOKUP TABLE OF DIVIDER BUFFER AMP GAIN |  |  |  |  | LOOKUP TABLE OF DIVIDER BUFFER AMP GAIN
|  | WP+ | 1 |  |  |  | WP+ | 1 |  |  |  | WP+ | 1 |  |  |  | WP+ | 1 |  |  |  | WP+ | 1 |  |  |  | WP+ | 1 |  |  |  |  | 1 | Supporting Information\CalPB_August05 updated.xls |  |  | 1 |  |  | 1
|  | WP- | -1.0003 |  |  |  | WP- | -1.0003 |  |  |  | WP- | -1.0003 |  |  |  | WP- | -1.0003 |  |  |  | WP- | -1.0003 |  |  |  | WP- | -1.0003 |  |  |  |  | -1.0003 | Supporting Information\CalPB_August05 updated.xls |  |  | -1.0003 |  |  | -1.0003 | Rtk2000a_check.xls cell EP353
|  | VP+ | 1 |  |  |  | VP+ | 1 |  |  |  | VP+ | 1 |  |  |  | VP+ | 1 |  |  |  | VP+ | 1 |  |  |  | VP+ | 1 |  |  |  |  | 1 | Supporting Information\CalPB_August05 updated.xls |  |  | 1 |  |  | 1
|  | VP- | -1.0003 |  |  |  | VP- | -1.0003 |  |  |  | VP- | -1.0003 |  |  |  | VP- | -1.0003 |  |  |  | VP- | -1.0003 |  |  |  | VP- | -1.0003 |  |  |  |  | -1.0003 | Supporting Information\CalPB_August05 updated.xls |  |  | -1.0003 |  |  | -1.0003 | Rtk2000a_check.xls cell EP353
|  | LOOKUP TABLE OF DIVIDER RATIOS |  |  |  |  | LOOKUP TABLE OF DIVIDER RATIOS |  |  |  |  | LOOKUP TABLE OF DIVIDER RATIOS |  |  |  |  | LOOKUP TABLE OF DIVIDER RATIOS |  |  |  |  | LOOKUP TABLE OF DIVIDER RATIOS |  |  |  |  | LOOKUP TABLE OF DIVIDER RATIOS | [7 empty] | Supporting Information\CalPB_August05 updated.xls |  |  |  | CalPBApr04
|  | Range | Ratio | Phase_m radians/Hz | Phase_c (radians) |  | Range | Ratio | Phase_m radians/Hz | Phase_c (radians) |  | Range |  |  |  |  | Range | Ratio | Phase_m radians/Hz | Phase_c (radians) |  | Range | Ratio | Phase_m radians | Phase_c |  | Range | Ratio | Phase_m radians | Phase_c |  |  | Ratio | Phase_m radians | Phase_c |  | Ratio | Phase_m radians | Phase_c | Ratio | Phase_m radians
|  | 60 | 67.6706 | 2.3655e-06 | 3.4895e-06 |  | 60 | 67.6706 | 2.3655e-06 | 3.4895e-06 |  | 60 | 67.6706 | 2.6828e-06 | 9.5482e-07 |  | 60 | 67.6701 | 2.2333e-06 | 2.6759e-05 |  | 60 | 67.6704 | 2.481e-06 | 1.8166e-05 |  | 60 | 67.6701 | 2.2303e-06 | 5.0789e-06 |  |  | 67.67 | 1.8413e-06 | -2.2263e-07 |  | 67.6707 | 2.2671e-06 | -1.3409e-05 | 67.6707 | 2.5098e-06
|  | 120 | 101.4975 | 2.4455e-06 | -9.2679e-06 |  | 120 | 101.4975 | 2.4455e-06 | -9.2679e-06 |  | 120 | 101.4975 | 2.6592e-06 | -1.3735e-05 |  | 120 | 101.4968 | 2.6895e-06 | -9.7096e-06 |  | 120 | 101.4972 | 2.6089e-06 | -6.401e-06 |  | 120 | 101.497 | 2.8567e-06 | -1.627e-05 |  |  | 101.4971 | 1.6081e-06 | 1.822e-05 |  | 101.4982 | 2.6957e-06 | -2.9066e-05 | 101.4975 | 2.5098e-06 | note 1
|  | 230 | 202.9816 | 2.6762e-06 | -2.6104e-05 |  | 230 | 202.9816 | 2.6762e-06 | -2.6104e-05 |  | 230 | 202.9816 | 2.6042e-06 | -1.9213e-05 |  | 230 | 202.9811 | 2.5014e-06 | -8.725e-06 |  | 230 | 202.9815 | 2.4463e-06 | -5.248e-06 |  | 230 | 202.9808 | 2.2827e-06 | 2.5598e-05 |  |  | 202.9813 | 1.2108e-06 | 5.7784e-05 |  | 202.9847 | 2.6938e-06 | -2.4016e-05 | 202.9816 | 2.5098e-06
[37 empty] | note 1 |  |  | note 1 | Rtk2000a_check.xls
sheet Calculations! block EQ341
|  | LOOKUP TABLE OF DET AMP PHASE ERRORS |  |  |  |  | LOOKUP TABLE OF DET AMP PHASE ERRORS |  |  |  |  | LOOKUP TABLE OF DET AMP PHASE ERRORS |  |  |  |  | LOOKUP TABLE OF DET AMP PHASE ERRORS |  |  | Supporting Information\CalPb_Det_auto_Dec08.xls |  | LOOKUP TABLE OF DET AMP PHASE ERRORS |  |  | Supporting Information\CalPb_Det_auto_Dec08.xls |  | LOOKUP TABLE OF DET AMP PHASE ERRORS |  |  | Supporting Information\CalPB_April08 updated.xls
| Degrees | NomFreq |  |  |  |  | NomFreq |  |  |  |  | NomFreq |  |  |  |  | NomFreq | Phase, G=1k | Phase,G=1 |  |  | NomFreq | Phase, G=1k | Phase,G=1 |  |  | NomFreq | Phase, G=1k | Phase,G=1 |  |  |  | Phase, G=1k | Phase,G=1 | Supporting Information\CalPB_August05 updated.xls |  | Phase, G=1k | Phase,G=1 | CalPBApr04 | Phase, G=1k | Phase,G=1
|  | 47 | -2.4779 | -2.3525 |  |  | 47 | -2.4779 | -2.3525 |  |  | 47 | -2.4305 | -2.3452 |  |  | 47 | -2.4521 | -2.3662 |  |  | 47 | -2.429 | -2.3374 |  |  | 47 | -2.4511 | -2.327 |  |  |  | -2.4337 | -2.3529 |  |  | -2.5281 | -2.359 |  | -2.3034 | -2.3962 |  | -2.5281 | -2.359
|  | 48 | -2.5275 | -2.4023 |  |  | 48 | -2.5275 | -2.4023 |  |  | 48 | -2.4807 | -2.3956 |  |  | 48 | -2.5026 | -2.416 |  |  | 48 | -2.4774 | -2.3871 |  |  | 48 | -2.4986 | -2.3789 |  |  |  | -2.4804 | -2.4017 |  |  | -2.5758 | -2.4079 |  | -2.3565 | -2.447 |  | -2.5758 | -2.4079
|  | 49 | -2.5772 | -2.452 |  |  | 49 | -2.5772 | -2.452 |  |  | 49 | -2.5308 | -2.446 |  |  | 49 | -2.5531 | -2.4657 |  |  | 49 | -2.5258 | -2.4369 |  |  | 49 | -2.5461 | -2.4309 |  |  |  | -2.5272 | -2.4504 |  |  | -2.6235 | -2.4568 |  | -2.4096 | -2.4978 |  | -2.6235 | -2.4568
|  | 49.5 | -2.602 | -2.4769 |  |  | 49.5 | -2.602 | -2.4769 |  |  | 49.5 | -2.5559 | -2.4712 |  |  | 49.5 | -2.5783 | -2.4906 |  |  | 49.5 | -2.55 | -2.4618 |  |  | 49.5 | -2.5699 | -2.4568 |  |  |  | -2.5505 | -2.4748 |  |  | -2.6474 | -2.4812
|  | 50 | -2.6268 | -2.5018 |  |  | 50 | -2.6268 | -2.5018 |  |  | 50 | -2.5809 | -2.4963 |  |  | 50 | -2.6036 | -2.5155 |  |  | 50 | -2.5742 | -2.4866 |  |  | 50 | -2.5936 | -2.4828 |  |  |  | -2.5739 | -2.4992 |  |  | -2.6712 | -2.5056 |  | -2.4627 | -2.5486 |  | -2.6712 | -2.5056
|  | 50.5 | -2.6516 | -2.5267 |  |  | 50.5 | -2.6516 | -2.5267 |  |  | 50.5 | -2.606 | -2.5215 |  |  | 50.5 | -2.6288 | -2.5404 |  |  | 50.5 | -2.5984 | -2.5115 |  |  | 50.5 | -2.6174 | -2.5088 |  |  |  | -2.5973 | -2.5236 |  |  | -2.6951 | -2.53
|  | 51 | -2.6765 | -2.5515 |  |  | 51 | -2.6765 | -2.5515 |  |  | 51 | -2.6311 | -2.5467 |  |  | 51 | -2.6541 | -2.5653 |  |  | 51 | -2.6226 | -2.5364 |  |  | 51 | -2.6411 | -2.5347 |  |  |  | -2.6206 | -2.5479 |  |  | -2.7189 | -2.5545 |  | -2.5158 | -2.5994 |  | -2.7189 | -2.5545
|  | 52 | -2.7261 | -2.6013 |  |  | 52 | -2.7261 | -2.6013 |  |  | 52 | -2.6812 | -2.5971 |  |  | 52 | -2.7046 | -2.6151 |  |  | 52 | -2.671 | -2.5861 |  |  | 52 | -2.6886 | -2.5867 |  |  |  | -2.6674 | -2.5967 |  |  | -2.7667 | -2.6033 |  | -2.5689 | -2.6502 |  | -2.7667 | -2.6033
|  | 53 | -2.7758 | -2.651 |  |  | 53 | -2.7758 | -2.651 |  |  | 53 | -2.7313 | -2.6474 |  |  | 53 | -2.7551 | -2.6648 |  |  | 53 | -2.7194 | -2.6359 |  |  | 53 | -2.7361 | -2.6386 |  |  |  | -2.7141 | -2.6455 |  |  | -2.8144 | -2.6522 |  | -2.622 | -2.701 |  | -2.8144 | -2.6522
|  | 54 | -2.8254 | -2.7008 |  |  | 54 | -2.8254 | -2.7008 |  |  | 54 | -2.7815 | -2.6978 |  |  | 54 | -2.8056 | -2.7146 |  |  | 54 | -2.7678 | -2.6856 |  |  | 54 | -2.7836 | -2.6906 |  |  |  | -2.7609 | -2.6942 |  |  | -2.8621 | -2.701 |  | -2.6751 | -2.7518 |  | -2.8621 | -2.701
|  | 55 | -2.8751 | -2.7505 |  |  | 55 | -2.8751 | -2.7505 |  |  | 55 | -2.8316 | -2.7482 |  |  | 55 | -2.8561 | -2.7644 |  |  | 55 | -2.8162 | -2.7354 |  |  | 55 | -2.8311 | -2.7425 |  |  |  | -2.8076 | -2.743 |  |  | -2.9098 | -2.7499 |  | -2.7282 | -2.8026 |  | -2.9098 | -2.7499
|  | 56 | -2.9247 | -2.8002 |  |  | 56 | -2.9247 | -2.8002 |  |  | 56 | -2.8817 | -2.7986 |  |  | 56 | -2.9066 | -2.8142 |  |  | 56 | -2.8646 | -2.7851 |  |  | 56 | -2.8786 | -2.7945 |  |  |  | -2.8543 | -2.7917 |  |  | -2.9575 | -2.7988 |  | -2.7813 | -2.8534 |  | -2.9575 | -2.7988
|  | 57 | -2.9744 | -2.85 |  |  | 57 | -2.9744 | -2.85 |  |  | 57 | -2.9319 | -2.8489 |  |  | 57 | -2.9571 | -2.8639 |  |  | 57 | -2.913 | -2.8348 |  |  | 57 | -2.9261 | -2.8464 |  |  |  | -2.9011 | -2.8405 |  |  | -3.0052 | -2.8476 |  | -2.8344 | -2.9042 |  | -3.0052 | -2.8476
|  | 58 | -3.024 | -2.8997 |  |  | 58 | -3.024 | -2.8997 |  |  | 58 | -2.982 | -2.8993 |  |  | 58 | -3.0076 | -2.9137 |  |  | 58 | -2.9613 | -2.8846 |  |  | 58 | -2.9736 | -2.8984 |  |  |  | -2.9478 | -2.8893 |  |  | -3.0529 | -2.8965
|  | 59 | -3.0737 | -2.9495 |  |  | 59 | -3.0737 | -2.9495 |  |  | 59 | -3.0321 | -2.9497 |  |  | 59 | -3.0581 | -2.9635 |  |  | 59 | -3.0097 | -2.9343 |  |  | 59 | -3.0211 | -2.9503 |  |  |  | -2.9945 | -2.938 |  |  | -3.1007 | -2.9453
|  | 60 | -3.1233 | -2.9992 |  |  | 60 | -3.1233 | -2.9992 |  |  | 60 | -3.0823 | -3 |  |  | 60 | -3.1086 | -3.0133 |  |  | 60 | -3.0581 | -2.9841 |  |  | 60 | -3.0686 | -3.0023 |  |  |  | -3.0413 | -2.9868 |  |  | -3.1484 | -2.9942
|  | 61 | -3.173 | -3.049 |  |  | 61 | -3.173 | -3.049 |  |  | 61 | -3.1324 | -3.0504 |  |  | 61 | -3.1591 | -3.063 |  |  | 61 | -3.1065 | -3.0338 |  |  | 61 | -3.1161 | -3.0542 |  |  |  | -3.088 | -3.0355 |  |  | -3.1961 | -3.043
|  | 62 | -3.2226 | -3.0987 |  |  | 62 | -3.2226 | -3.0987 |  |  | 62 | -3.1825 | -3.1008 |  |  | 62 | -3.2096 | -3.1128 |  |  | 62 | -3.1549 | -3.0836 |  |  | 62 | -3.1636 | -3.1062 |  |  |  | -3.1348 | -3.0843 |  |  | -3.2438 | -3.0919
|  | 63 | -3.2723 | -3.1485 |  |  | 63 | -3.2723 | -3.1485 |  |  | 63 | -3.2327 | -3.1511 |  |  | 63 | -3.2601 | -3.1626 |  |  | 63 | -3.2033 | -3.1333 |  |  | 63 | -3.2111 | -3.1581 |  |  |  | -3.1815 | -3.1331 |  |  | -3.2915 | -3.1408
|  | 75 | -3.8681 | -3.7454 |  |  | 75 | -3.8681 | -3.7454 |  |  | 75 | -3.8342 | -3.7556 |  |  | 75 | -3.8661 | -3.7599 |  |  | 75 | -3.784 | -3.7303 |  |  | 75 | -3.7812 | -3.7814
WG2 PRT coefficients | Ohms | R0 | 99.9625 |  |  |  | R0 | 99.9636 |  |  |  | R0 | 99.9617 |  |  |  | R0 | 99.961 |  |  |  | R0 | 99.9612 |  |  |  | R0 | 99.9573 |  |  |  |  | 99.9668
| 1/K | Alpha | 0.0039125 |  |  |  | Alpha | 0.0039125 |  |  |  | Alpha | 0.003908 |  |  |  | Alpha | 0.003908 |  |  |  | Alpha | 0.003908 |  |  |  | Alpha | 0.003908 |  |  |  |  | 0.0039114
| 1/K^2 | Beta | -6.279e-07 |  |  |  | Beta | -6.279e-07 |  |  |  | Beta | -5.143e-07 |  |  |  | Beta | -5.143e-07 |  |  |  | Beta | -5.143e-07 |  |  |  | Beta | -5.143e-07 |  |  |  |  | -5.964e-07
| ppm | 3478A Error | -124.2698 |  |  |  | 3478A Error | -113.0657 |  |  |  | 3478A Error | -88.9827 |  |  |  | 3478A Error | 0 |  |  |  | 3478A Error | 0 |  |  |  | 3478A Error | 0 |  |  |  |  | 0
Counter correction | ppm | CountCorr | 0 | For Agilent 53230A |  |  | CountCorr | -1.79 | For HP53132A (s# KR91200497) |  |  | CountCorr | -1.58 | For HP53132A (s# KR91200497)
Note on D6: t.stewart: Cell M49 of "Detector Feb 2014" worksheet of "CalPb_Det_auto_Feb 14.xls"
Note on P6: t.stewart: Cell M49 of "Detector May 2012" worksheet of "CalPb_Det_auto_May12.xls"
Note on V6: t.stewart: Cell M49 of "Detector Dec 2010" worksheet of "CalPb_Det_auto_Dec10.xls"
Note on AB6: t.stewart: Cell M49 of "Detector Jan 2009" worksheet of "CalPb_Det_auto_Dec08.xls"
Note on D7: Industrial Research Limited: Cell BS195 of "Date" worksheet HP245dcv2014#9ver1.21.xls (May 14 calibration). The DMM reads high on positive volts
Note on P7: Industrial Research Limited: Cell BP189 of "23Feb11" worksheet HP245dcv2011#9ver1.13corrected.xls ( Feb 11 calibration). The DMM reads high on positive volts
Note on V7: Industrial Research Limited: Cell BP189 of "23Feb11" worksheet HP245dcv2011#9ver1.13.xls ( Feb 11 calibration). The DMM reads high on positive volts
Note on AB7: Industrial Research Limited: Feb 08 calibration. The DMM reads low on positive volts
Note on AH7: Industrial Research Limited: Feb 08 calibration. The DMM reads low on positive volts
Note on AM7: Industrial Research Limited: See HP3458CalibrationHistory.xls
Note on AQ7: Laurie Christian: From dcV calibration report for this DVM- May 2004
Note on D8: t.stewart: Cell G13 of "DC Resistance Feb 2014" worksheet of CalPb_Feb14.xls"
Note on P8: t.stewart: Cell G13 of "DC Resistance Apr 2012" worksheet of CalPb_Apr12.xls"
Note on V8: t.stewart: Cell G99 of "DC Resistance Dec 2010" worksheet of CalPb_Dec10.xls"
Note on AM8: Laurie Christian: 'DC Resistance Dec 2005' G26
Note on AQ8: Laurie Christian: 'DC Resistance' G37
Note on B9: Andrew Corney: L=-CR^2 where C is parallel capacitance
Note on AQ9: Laurie Christian: Historic- measured on universal bridge- minor term at 50 Hz
Note on D10: t.stewart: Cell G14 of "DC Resistance Feb 2014" worksheet of CalPb_Feb14.xls"
Note on P10: t.stewart: Cell G14 of "DC Resistance Apr 2012" worksheet of CalPb_Apr12.xls"
Note on V10: t.stewart: Cell G100 of "DC Resistance Dec 2010" worksheet of CalPb_Dec10.xls"
Note on B11: Andrew Corney: L=-CR^2 where C is parallel capacitance
Note on AQ11: Laurie Christian: Historic- measured on universal bridge- minor term at 50 Hz
Note on V12: Callum McKenzie: Calculated from R1, R2, Ignores terms second-order in L.
Note on AB12: Callum McKenzie: Calculated from R1, R2, Ignores terms second-order in L.
Note on AH12: Callum McKenzie: Calculated from R1, R2, Ignores terms second-order in L.
Note on AM12: Callum McKenzie: Calculated from R1, R2, Ignores terms second-order in L. Previoulsy this was calculated off-line and inserted as a constant.
Note on B13: Andrew Corney: Time consant of ratio. Multiply by omega to get imaginary part of ratio
Note on D13: This is now calculated from R1, R2, L1 and L2 template values. It ignores terms second-order in L. See page 23 of Callum McKenzie Rough Notes and Ideas 1.
Note on P13: This is now calculated from R1, R2, L1 and L2 template values. It ignores terms second-order in L. See page 23 of Callum McKenzie Rough Notes and Ideas 1.
Note on V13: This is now calculated from R1, R2, L1 and L2 template values. It ignores terms second-order in L. See page 23 of Callum McKenzie Rough Notes and Ideas 1.
Note on AB13: This is now calculated from R1, R2, L1 and L2 template values. It ignores terms second-order in L. See page 23 of Callum McKenzie Rough Notes and Ideas 1.
Note on AH13: This is now calculated from R1, R2, L1 and L2 template values. It ignores terms second-order in L. See page 23 of Callum McKenzie Rough Notes and Ideas 1.
Note on AM13: Industrial Research Limited: Calculated from R1, R2, L1 and L2. Ignores terms second-order in L. Previoulsy this was calculated off-line and inserted as a constant.
Note on V14: t.stewart: From cell BT25 of "CalPb_C1_Dec2010.xls"
Note on AB14: t.stewart: From cell BT77 of From "CalPb_C1_Dec2008.xls"
Note on AM14: Callum McKenzie: C1 at 24.10C with Pt internal thermomter
Note on AQ14: Laurie Christian: Measured at 23.6 C true diecast box temperature
Note on CM16: A 14 ppm correction is applied here to allow for the increase in resistance at 53 Hz over the dc value.
Note on D19: t.stewart: Cell M48 of "Detector Feb 2014" worksheet of "CalPb_Det_auto_Feb 14.xls"
Note on P19: t.stewart: Cell M48 of "Detector May 2012" worksheet of "CalPb_Det_auto_May 12.xls"
Note on V19: t.stewart: Cell M48 of "Detector Dec 2010" worksheet of "CalPb_Det_auto_Dec10.xls"
Note on AB19: t.stewart: Cell M48 of "Detector Jan 2009" worksheet of "CalPb_Det_auto_Dec08.xls"
Note on AM19: Industrial Research Limited: 'Detector 2005(4)' page
Note on AN19: Callum McKenzie: These values are after the grounding problesm were fixed.
Note on D20: t.stewart: Cell M49 of "Detector Feb 2014" worksheet of "CalPb_Det_auto_Feb 14.xls"
Note on P20: t.stewart: Cell M49 of "Detector May 2012" worksheet of "CalPb_Det_auto_May12.xls"
Note on V20: t.stewart: Cell M49 of "Detector Dec 2010" worksheet of "CalPb_Det_auto_Dec10.xls"
Note on AB20: t.stewart: Cell M49 of "Detector Jan 2009" worksheet of "CalPb_Det_auto_Dec08.xls"
Note on AM20: Industrial Research Limited: 'Detector 2005(4)' page
Note on AN20: Callum McKenzie: These values are after the grounding problesm were fixed.
Note on C22: Laurie Christian: ESI 10*1 ohm in parallel
Note on D22: Industrial Research Limited: -73.610 ppm From BuildupESIboxes2014.xls located "I:\MSL\Private\Electricity\Ongoing\OHM\ESI_Boxes" worksheet "ResultsFromT16_T81", cell E54. This is a corrected to 20° value dated at 30/1/14. TJS
Note on O22: Laurie Christian: ESI 10*1 ohm in parallel
Note on P22: Industrial Research Limited: -70.877 ppm from cell E10 of "ValuesUsed" worksheet in "BuildESIboxes2012 2012 Restraint.xls" , located "I:\MSL\Private\Electricity\Ongoing\OHM\ESI_Boxes". This is a corrected to 20°C number and is dated 16-3-2012. TJS
Note on U22: Laurie Christian: ESI 10*1 ohm in parallel
Note on V22: Industrial Research Limited: -70.223 ppm from cell AE777 of "Calculations." worksheet in "BuildESIboxes2006a.xls" August 2006, located "G:\Msl\estds\ONGOING\OHM\ESI_Boxes". This is a corrected to 20°C number and is dated 14-2-2007. May be closer to -75ppm in 2009. BEM.
Note on AA22: Laurie Christian: ESI 10*1 ohm in parallel
Note on AB22: Industrial Research Limited: -70.223 ppm from cell AE777 of "Calculations." worksheet in "BuildESIboxes2006a.xls" August 2006, located "G:\Msl\estds\ONGOING\OHM\ESI_Boxes". As of 06/03/09 there is an indication that this has moved to -64ppm. BEM.
Note on AG22: Laurie Christian: ESI 10*1 ohm in parallel
Note on AH22: Industrial Research Limited: -70.223 ppm from cell AE777 of "Calculations." worksheet in "BuildESIboxes2006a.xls" August 2006, located "G:\Msl\estds\ONGOING\OHM\ESI_Boxes".
Note on AM22: Bruce McLennan: From ESI_History.xls in Ongoing/ESI_Boxes. August 2006 value for 0.1 ohm configuration
Note on AQ22: Laurie Christian: Cal temp = 19.76 C
Note on D23: Industrial Research Limited: -26 ppm. From BuildupESIboxes2014.xls located in "I:\MSL\Private\Electricity\Ongoing\OHM\ESI_Boxes" worksheet "ResultsFromT16_T81", average of cells L9:L13. This is a corrected to 20° value dated at 31/1/14. See also report S21895 for supporting information including phase. TJS
Note on P23: Industrial Research Limited: From cell L32 of "ValuesUsed" worksheet of "Buildup2010.xls" in "G:\Msl\estds\ONGOING\OHM\" and using date of interest as 12 Dec 2011. BM & TJS 18/4/11
Note on V23: Industrial Research Limited: From cell L32 of "ValuesUsed" worksheet of "Buildup2010.xls" in "G:\Msl\estds\ONGOING\OHM\" and using date of interest as 12 Dec 2011. BM & TJS 18/4/11
Note on AB23: Industrial Research Limited: From Buildup2008.xls, cell Y983 of Calculations page. 20 C value using 1.43 ppm/C temp co.)
Note on AH23: Industrial Research Limited: From Buildup2008.xls, cell Y983 of "Calculations." worksheet. 20 C value using 1.43 ppm/C temp co.)
Note on AM23: Bruce McLennan: Value from Estds\Ongoing\Shunts\Shunt_hist.xls. 8-Dec-06 value extrapolated to 20 C. Drift with time is linear with slope of approx 2.2 ppm/year. Heating with cooling from PaddleWheelShunt.xls in Calibrat/Power&Energy/PaddleWheel. (0.368 ppm/Watt). Temp Co in Watt0.2.xls in Ongoing/Ohm/Shunts. Approx 1.4 ppm/C
Note on AQ23: Bruce McLennan: Value from MI_Values in G:\estds\Reports\2004\S21537. At 20 C. Heating with cooling from PaddleWheelShunt.xls in Calibrat/Power&Energy/PaddleWheel. (0.368 ppm/Watt). Temp Co in Watt0.2.xls in Ongoing/Ohm/Shunts. Approx 1.4 ppm/C
Note on D24: Industrial Research Limited: From cell F17 of "ValuesUsed" worksheet of "Buildup2012 2012 Restraint.xls" in "\\Electricity\Ongoing\OHM\" and using Date of interest as 19/04/2012. TJS
Note on P24: Industrial Research Limited: From cell F17 of "ValuesUsed" worksheet of "Buildup2012 2012 Restraint.xls" in "\\Electricity\Ongoing\OHM\" and using Date of interest as 19/04/2012. TJS
Note on V24: Industrial Research Limited: From cell L33 of "ValuesUsed" worksheet of "Buildup2010.xls" in "G:\Msl\estds\ONGOING\OHM\" and using Date of interest as 14/12/2010. BM
Note on AB24: Industrial Research Limited: From cell D19 of "Current Value" worksheet of "1ohm_data.xls" in "G:\Msl\estds\ONGOING\OHM\NZ ohm June 2005 with Least squares equations" and using Date of interest as 23/01/2009. Consistent with Buildup2008.xls
Note on AH24: Industrial Research Limited: From cell D19 of "Current Value" worksheet of "1ohm_data.xls" in "G:\Msl\estds\ONGOING\OHM\NZ ohm June 2005 with Least squares equations" and using Date of interest as 13/11/2008. Consistent with Buildup2008.xls
Note on AM24: Laurie Christian: 8-Dec-2006 value at 20 C. I n 1ohm_data.xls in G:\estds\ONGOING\OHM\NZ ohm June 2005 with Least squares equations
Note on AQ24: Laurie Christian: At 20 C. I n Buildup2003a.xls in Ongoing/Ohm/NZ Ohm Feb 2004/
Note on D25: Industrial Research Limited: Value at 1 A and 20 C from LoShunt2008.xls. Heating at 8 A from 1 A is about 150 ppm. Can apply temp correction and measured resistor temp for other currents (see 8 A value)
Note on P25: Industrial Research Limited: Value at 1 A and 20 C from LoShunt2008.xls. Heating at 8 A from 1 A is about 150 ppm. Can apply temp correction and measured resistor temp for other currents (see 8 A value)
Note on V25: Industrial Research Limited: Value at 1 A and 20 C from LoShunt2008.xls. Heating at 8 A from 1 A is about 150 ppm. Can apply temp correction and measured resistor temp for other currents (see 8 A value)
Note on AB25: Industrial Research Limited: Value at 1 A and 20 C from LoShunt2008.xls. Heating at 8 A from 1 A is about 150 ppm. Can apply temp correction and measured resistor temp for other currents (see 8 A value)
Note on AH25: Industrial Research Limited: Value at 1 A and 20 C from LoShunt2008.xls. Heating at 8 A from 1 A is about 150 ppm. Can apply temp correction and measured resistor temp for other currents (see 8 A value)
Note on AM25: Bruce McLennan: This is the 10 A value extrapolated to 20 C using temperature coefficients of 26.4 ppm/C and -0.6 ppm/C^2 (from Analysis1_Ti122 in G:\estds\Ongoing\Shunts\Lo Shunt 2006.xls). At present there is no allowance made for the temperature or current dependence. Between 5 A and 10 A at a fixed ambient temperature the resistor changes by 130 ppm and can't be ignored.
Note on AQ25: Bruce McLennan: From MI_Values in G:\estds\Reports\2004\S21537. At 20.00 C
Note on D26: Industrial Research Limited: From cell F18 of "ValuesUsed" worksheet of "Buildup2012 2012 Restraint.xls" in "\\Electricity\Ongoing\OHM\" and using Date of interest as 07/05/2012. TJS
Note on P26: Industrial Research Limited: From cell F18 of "ValuesUsed" worksheet of "Buildup2012 2012 Restraint.xls" in "\\Electricity\Ongoing\OHM\" and using Date of interest as 07/05/2012. TJS
Note on U26: Industrial Research Limited: TI(924) should not be taken from oil bath
Note on V26: Industrial Research Limited: From cell D19 of "Current Value" worksheet of "10ohm_data.xls" in "G:\Msl\estds\ONGOING\OHM\NZ ohm June 2005 with Least squares equations" and using Date of interest as 23/01/2009.
Note on AA26: Industrial Research Limited: TI(924) should not be taken from oil bath
Note on AB26: Industrial Research Limited: From cell D19 of "Current Value" worksheet of "10ohm_data.xls" in "G:\Msl\estds\ONGOING\OHM\NZ ohm June 2005 with Least squares equations" and using Date of interest as 23/01/2009.
Note on AG26: Industrial Research Limited: TI(924) should not be taken from oil bath
Note on AH26: Industrial Research Limited: From cell D19 of "Current Value" worksheet of "10ohm_data.xls" in "G:\Msl\estds\ONGOING\OHM\NZ ohm June 2005 with Least squares equations" and using Date of interest as 13/11/2008.
Note on AM26: Laurie Christian: 8-Dec-2006 value at 20 C. I n 10ohm_data.xls in G:\estds\ONGOING\OHM\NZ ohm June 2005 with Least squares equations
Note on AQ26: Laurie Christian: At 20 C. In Buildup2003a in Ongoing.
Note on D27: Industrial Research Limited: -6.439 ppm From BuildupESIboxes2014.xls located "I:\MSL\Private\Electricity\Ongoing\OHM\ESI_Boxes" worksheet "ResultsFromT16_T81", cell L54. This is a corrected to 20° value dated at 31/1/14. See also report S21895 for supporting information including phase. TJS
Note on P27: Industrial Research Limited: -5.878ppm from cell E11 of "ValuesUsed" worksheet in "BuildESIboxes2012 2012 Restraint.xls" , located "I:\MSL\Private\Electricity\Ongoing\OHM\ESI_Boxes". This is a corrected to 20°C number and is dated 11-3-2012. TJS
Note on V27: Industrial Research Limited: -6.562 ppm from cell AG39 of "Calculations" worksheet in "BuildESIboxes2009.xls" . The test was done in March 2009, the spreadsheet is located in: "G:\Msl\estds\ONGOING\OHM\ESI_Boxes".
Note on AB27: Industrial Research Limited: -7.6007 ppm from cell O46 of "ResultsFromT16_T81" worksheet in "BuildESIboxes2006a.xls" August 2006 located "G:\Msl\estds\ONGOING\OHM\ESI_Boxes".
Note on AH27: Industrial Research Limited: -7.6007 ppm from cell O46 of "ResultsFromT16_T81" worksheet in "BuildESIboxes2006a.xls" August 2006 located "G:\Msl\estds\ONGOING\OHM\ESI_Boxes".
Note on AM27: Bruce McLennan: From ESI_History.xls in Ongoing/ESI_Boxes. August 2006 value for 1 ohm configuration
Note on AQ27: Bruce McLennan: From MI_Values in G:\estds\Reports\2004\S21537. At 20.514 C
Note on D28: Industrial Research Limited: -4.582 ppm From BuildupESIboxes2014.xls located "I:\MSL\Private\Electricity\Ongoing\OHM\ESI_Boxes" worksheet "ResultsFromT16_T81", cell M54. This is a corrected to 20° value dated at 1/2/14. See also report S21895 for supporting information including phase. TJS
Note on P28: Industrial Research Limited: -4.283 ppm from cell E12 of "ValuesUsed" worksheet in "BuildESIboxes2012 2012 Restraint.xls" , located "I:\MSL\Private\Electricity\Ongoing\OHM\ESI_Boxes". This is a corrected to 20°C number and is dated 6-3-2012. TJS
Note on V28: Industrial Research Limited: -4.729 ppm from cell AL39 of "Calculations" worksheet in "BuildESIboxes2009.xls". The test was done in March 2009, the spreadsheet is located in: "G:\Msl\estds\ONGOING\OHM\ESI_Boxes".
Note on AB28: Industrial Research Limited: -6.112 ppm from cell P46 of "ResultsFromT16_T81" worksheet in "BuildESIboxes2006a.xls" August 2006, located "G:\Msl\estds\ONGOING\OHM\ESI_Boxes".
Note on AH28: Industrial Research Limited: -6.112 ppm from cell P46 of "ResultsFromT16_T81" worksheet in "BuildESIboxes2006a.xls" August 2006, located "G:\Msl\estds\ONGOING\OHM\ESI_Boxes".
Note on AM28: Bruce McLennan: From ESI_History.xls in Ongoing/ESI_Boxes. August 2006 value for 10 ohm configuration
Note on AQ28: Bruce McLennan: From ESI_History.xls in Ongoing
Note on D32: t.stewart: From cells G29, G30 of "AC gains Feb 2014" worksheet of "CalPb_Feb12.xls".
Note on P32: t.stewart: From cells G29, G30 of "AC gains May 2012" worksheet of "CalPb_Apr12.xls".
Note on V32: t.stewart: From cells G29, G30 of "AC gains Dec 2010" worksheet of "CalPb_Dec10.xls".
Note on AU33: Andrew Corney: see Rtk2000a_check.xls cell EP353
Note on AU35: Andrew Corney: see Rtk2000a_check.xls cell EP353
Note on D39: t.stewart: From cells E137-139 of "HV divider Feb 2014" worksheet of "CalPb_Feb14.xls".
Note on E39: t.stewart: From cells R89:S91 of "Phase v3 May 2014" worksheet of "CalPb_Feb14.xls".
Note on P39: t.stewart: From cells E137-139 of "HV divider May 2012" worksheet of "CalPb_Apr12.xls".
Note on Q39: t.stewart: From cells R82:S84 of "Phase v2 May 2012" worksheet of "CalPb_Apr12.xls".
Note on V39: t.stewart: From cells E137-139 of "HV divider Dec 2010" worksheet of "CalPb_Dec10.xls".
Note on W39: t.stewart: From cells R82:S84 of "Phase v2 Dec 2010" worksheet of "CalPb_Dec10.xls".
Note on AM39: Callum McKenzie: Taken from the HV Divider (3) worksheet.
Note on AN39: Industrial Research Limited: From Phase Dec 2005 page
Note on AV40: Andrew Corney: from Rtk2000a_check.xls sheet Calculations! block EQ341
Note on AM45: Laurie Christian: Found columns reversed on 16 June 2004
Note on AO45: Callum McKenzie: These values are after the grounding problesm were fixed.
Note on AQ45: Laurie Christian: Found columns reversed on 16 June 2004
Note on D46: t.stewart: Cells U53:V72 of "Detector Feb 2014" worksheet of "CalPb_Det_auto_Feb14.xls"
Note on P46: t.stewart: Cells U53:V72 of "Detector May 2012" worksheet of "CalPb_Det_auto_May12.xls"
Note on V46: t.stewart: Cells U53:V72 of "Detector Dec 2010" worksheet of "CalPb_Det_auto_Dec10.xls"
Note on AB46: t.stewart: Cells U53:V72 of "Detector Jan 2009" worksheet of "CalPb_Det_auto_Dec08.xls"
Note on V65: l.christian: Is significantly different to the last two times- keep an eye on this.
Note on D66: Industrial Research Limited: Jan 2015 Calibration. From Cell M8 of "Results" worksheet of "WG2.xls" located in "I:\MSL\Private\Electricity\Ongoing\OHM\Temperature_PRTs\"
Note on P66: Industrial Research Limited: Dec 11 Calibration. From Cell M8 of "Results" worksheet of "WG2.xls" located in "I:\MSL\Private\Electricity\Ongoing\OHM\Temperature_PRTs\"
Note on V66: Industrial Research Limited: Nov 10 Calibration. From Cell M8 of "Results" worksheet of "WG2.xls" located in "G:\Msl\estds\ONGOING\OHM\Temperature_PRTs\"
Note on AB66: Industrial Research Limited: Nov 08 Calibration. From Cell M8 of "Results" worksheet of "WG2.xls" located in "G:\Msl\estds\ONGOING\OHM\Temperature_PRTs\"
Note on AH66: Industrial Research Limited: From Temperature/2008/910, 31 October 2008. Probe dried out prior to these values because it had significant leakage.
Note on AM66: Callum McKenzie: These three from the WG2 spreadsheet.
Note on D67: Industrial Research Limited: Dec 13 Calibration. From Cell H8 of "Results" worksheet of "WG2.xls" located in "I:\MSL\Private\Electricity\Ongoing\OHM\Temperature_PRTs\"
Note on P67: Industrial Research Limited: Dec 11 Calibration. From Cell H8 of "Results" worksheet of "WG2.xls" located in "I:\MSL\Private\Electricity\Ongoing\OHM\Temperature_PRTs\"
Note on V67: t.stewart: November 2010 Calibration. From Cell H8 of "Results" worksheet of "WG2.xls" located in "G:\Msl\estds\ONGOING\OHM\Temperature_PRTs\"
Note on AB67: t.stewart: October 2008 Calibration. From Cell H8 of "Results" worksheet of "WG2.xls" located in "G:\Msl\estds\ONGOING\OHM\Temperature_PRTs\"
Note on D68: Industrial Research Limited: Dec 13 Calibration. From Cell I8 of "Results" worksheet of "WG2.xls" located in "I:\MSL\Private\Electricity\Ongoing\OHM\Temperature_PRTs\"
Note on P68: Industrial Research Limited: Dec 11 Calibration. From Cell I8 of "Results" worksheet of "WG2.xls" located in "I:\MSL\Private\Electricity\Ongoing\OHM\Temperature_PRTs\"
Note on V68: t.stewart: November 2010 Calibration. From Cell I8 of "Results" worksheet of "WG2.xls" located in "G:\Msl\estds\ONGOING\OHM\Temperature_PRTs\"
Note on AB68: t.stewart: October 2008 Calibration. From Cell I8 of "Results" worksheet of "WG2.xls" located in "G:\Msl\estds\ONGOING\OHM\Temperature_PRTs\"
Note on D69: t.stewart: Calibrated 20 Jan 2015 (HP3478A Maintenance Book), From cell D20 of Summary page of Ongoing\DMMs\HP3478A\HP3478A.xls
Note on P69: t.stewart: Calibrated 19 June 2012 (HP3478A Maintenance Book), From cell D18 of Summary page of Ongoing\DMMs\HP3478A\HP3478A.xls
Note on V69: t.stewart: Calibrated in June 2010 (HP3478A Maintenance Book), error is effectively 0 for this instrument as it only has 4.5 digits.
Note on AB69: t.stewart: Calibrated in April 08 (HP3478A Maintenance Book), and has an error of -2 ppm. This is effectively 0 for this instrument as it only has 4.5 digits.
Note on AH69: Callum McKenzie: Nulled out by Bruce. See 10 Feb 2006 cals and 3478 maintenance book.
Note on AM69: Callum McKenzie: Nulled out by Bruce. See 10 Feb 2006 cals and 3478 maintenance book.
Note on E71: t.stewart: Measured 28/5/14 against HP3325A locked to D2 of time lab.
Note on Q71: t.stewart: Measured 25/6/12 against HP53132A counter locked to D2 of clock room.

## Uncertainties
|  | Uncertainties updated based on APMP K5.1 power comparison (S21885) |  |  |  | LAC checked 19 May 2011 |  | This uncertainty modified 18/5/11 because of new linearity measurements done | [6 empty] | Laurie checked this data in June 2004 |  | This uncertainty modified 2/12/04 because of new linearity measurements done- see Pb_Auto_V1d_cal04A.xls | [6 empty] | Laurie checked this data in April 2006
[15 empty] | Need rechecking by June 30 2005 or before if required |  |  | Have gone back to Uncertainties from Pb_Auto_V1g (08/02/05) | [5 empty] | Need rechecking by June 30 2007 or before if required
|  | Standard Uncertainties | Accuracy of | Degrees of |  | Standard Uncertainties | Accuracy of | Degrees of | [6 empty] | Standard Uncertainties | Accuracy of | Degrees of | as these were the last approved uncertainties (TJS 14/5/09) | [5 empty] | Standard Uncertainties | Accuracy of | Degrees of |  | These Uncertainties are yet to be approved (14/05/09)
|  | u (ppm) | uncertainty | freedom, v | Comment | u (ppm) | uncertainty | freedom, v | Comment | [5 empty] | u (ppm) | uncertainty | freedom, v | Comment | [5 empty] | u (ppm) | uncertainty | freedom, v | Comment
1 | Divider Ratio | 5 |  | 6 |  | 5 | 40 | 3.125 |  | 5 | 5 |  |  |  | 5 | 40 | 3.125 |  | 5 | 5 |  |  |  | 7 | 40 | 3.125 |  | 5 | 5
2 | Divider Phase | 5 | 20 | 12.5 |  | 15 | 30 | 5.5556 |  | 15 | 15 | V1_calsdv | 15 | 40 | 15 | 30 | 5.5556 |  | 15 | 15 | V1_calsdv | 15 | 40 | 15 | 30 | 5.5556 |  | 15 | 15 | V1_calsdv | 15 | 40
3 | Ref Volts DMM Cal | 4.3566 |  | 50 |  | 6.8 | 30 | 5.5556 |  | 15 | 15 | Va_calsdv | 500 | 20 | 6.8 | 30 | 5.5556 |  | 15 | 15 | Va_calsdv | 500 | 20 | 5.2202 | 30 | 5.5556 |  | 15 | 15 | Va_calsdv | 500 | 20
4 | Shunt Resistance | 10 |  | 10 | #NAME? | 9.2 | 30 | 5.5556 |  | 9.2 | 9.2 | Phi_calsdv | 872.6646 | 30 | 9.2 | 30 | 5.5556 |  | 9.2 | 9.2 | Phi_calsdv | 872.6646 | 30 | 9.2 | 30 | 5.5556 |  | 9.2 | 9.2 | Phi_calsdv | 872.6646 | 30
5 | Shunt Phase | 1 |  | 50 | #NAME? | 7 | 30 | 5.5556 |  | 7 | 7 | K_sdv | 5 | 40 | 7 | 30 | 5.5556 |  | 7 | 7 | K_sdv | 5 | 40 | 7 | 30 | 5.5556 |  | 7 | 7 | K_sdv | 5 | 40
6 | Watts Dial | 2 | 25 | 8 |  | 2 | 30 | 5.5556 |  | 1 | 0 | Kj_sdv | 15 | 30 | 5 | 30 | 5.5556 | unexplained dial non-linearity | 1 | 0 | Kj_sdv | 15 | 30 | 5 | 30 | 5.5556 |  | 1 | 0 | Kj_sdv | 15 | 30
7 | Watts Buffer | 2 | 25 | 8 |  | 2 | 30 | 5.5556 |  | 1 | 0 | S_sdv | 9.1897 | 30 | 2 | 30 | 5.5556 |  | 1 | 0 | S_sdv | 9.1897 | 30 | 2 | 30 | 5.5556 |  | 1 | 0 | S_sdv | 9.1897 | 30
8 | R2_R1 ratio | 5 | 25 | 8 |  | 5 | 30 | 5.5556 |  | 2 | 0 | Sj_sdv | 7 | 30 | 5 | 30 | 5.5556 |  | 2 | 0 | Sj_sdv | 7 | 30 | 3 | 30 | 5.5556 |  | 2 | 0 | Sj_sdv | 7 | 30
9 | R2_R1 Phase | 1 | 25 | 8 |  | 1 | 30 | 5.5556 |  | 5 | 0 | Cj_sdv | 15 | 30 | 1 | 30 | 5.5556 |  | 5 | 0 | Cj_sdv | 15 | 30 | 3.0625 | 30 | 5.5556 |  | 5 | 0 | Cj_sdv | 15 | 30
10 | Vars Dial | 2 | 25 | 8 |  | 2 | 30 | 5.5556 |  | 1 | 0 | C_sdv | 50 | 30 | 15 | 30 | 5.5556 | unexplained dial non-linearity | 1 | 0 | C_sdv | 50 | 30 | 5 | 30 | 5.5556 |  | 1 | 0 | C_sdv | 50 | 30
11 | Vars Buffer | 1 | 25 | 8 |  | 1 | 30 | 5.5556 |  | 1 | 0 |  |  |  | 1 | 30 | 5.5556 |  | 1 | 0 |  |  |  | 2 | 30 | 5.5556 |  | 1 | 0
12 | C1 phase | 15 | 20 | 12.5 |  | 15 | 30 | 5.5556 |  | 50 | 15 |  |  |  | 15 | 30 | 5.5556 |  | 50 | 15 |  |  |  | 15 | 30 | 5.5556 |  | 50 | 15
13 | C1 capacitance | 50 | 20 | 12.5 |  | 50 | 30 | 5.5556 |  | 15 | 50 |  |  |  | 50 | 30 | 5.5556 |  | 15 | 50 |  |  |  | 28.4431 | 30 | 5.5556 |  | 15 | 50
14 | R2 resistance | 2 | 25 | 8 |  | 2 | 30 | 5.5556 |  | 2 | 0 |  |  |  | 2 | 30 | 5.5556 |  | 2 | 0 |  |  |  | 3 | 30 | 5.5556 |  | 2 | 0
15 | R2 phase (urad) | 2 | 25 | 8 |  | 2 | 30 | 5.5556 |  | 2 | 0 |  |  |  | 2 | 30 | 5.5556 |  | 2 | 0 |  |  |  | 3.0625 | 30 | 5.5556 |  | 2 | 0
16 | Det volts | 500 | 20 | 12.5 |  | 500 | 20 | 12.5 |  | 500 | 500 |  |  |  | 500 | 20 | 12.5 |  | 500 | 500 |  |  |  | 500 | 20 | 12.5 |  | 500 | 500
17 | Det phase (urad) | 873 | 25 | 8 |  | 873 | 30 | 5.5556 |  | 873 | 873 |  |  |  | 873 | 30 | 5.5556 |  | 873 | 873 |  |  |  | 873 | 30 | 5.5556 |  | 873 | 873
18 | Det Amp Gain | 500 | 25 | 8 |  | 500 | 30 | 5.5556 |  | 500 | 0 |  |  |  | 500 | 30 | 5.5556 |  | 500 | 0 |  |  |  | 500 | 30 | 5.5556 |  | 500 | 0
19 | Det Amp Phase (urad) | 500 | 25 | 8 |  | 500 | 30 | 5.5556 |  | 500 | 0 |  |  |  | 500 | 30 | 5.5556 |  | 500 | 0 |  |  |  | 1000 | 30 | 5.5556 |  | 500 | 0
20 | Det Amp conductance | 10000 | 25 | 8 |  | 10000 | 30 | 5.5556 |  | 10000 | 0 |  |  |  | 10000 | 30 | 5.5556 |  | 10000 | 0 |  |  |  | 10000 | 30 | 5.5556 |  | 10000 | 0
21 | Det Amp Admittance | 10000 | 25 | 8 |  | 10000 | 30 | 5.5556 |  | 10000 | 0 |  |  |  | 10000 | 30 | 5.5556 |  | 10000 | 0 |  |  |  | 10000 | 30 | 5.5556 |  | 10000 | 0
22 | Watts dial offset | 1 | 10 | 50 |  | 1 | 10 | 50
23 | Vars dial offset | 1 | 10 | 50 |  | 1 | 10 | 50
24 | Watts dial phase shift | 1 | 10 | 50 |  | 1 | 10 | 50
25 | Vars dial phaseshift | 1 | 10 | 50 |  | 1 | 10 | 50
Note on E5: Tom Stewart: Changed to 6 dof based on variation of calibrated history. See history graph on HV divider May 2012 worksheet in CalPB_Apr12.xls
Note on AD5: Industrial Research Limited: Between 10 and 20 ppm change since April 2004- need to keep an eye on this but increase uncertainty in the meantime
Note on C6: Tom Stewart: See GTC calculation WB_divider_phase.py for uncertainty determination of calibration method (<2 urad). The change over time shows a drift of ~1.5 urad per year. Also need to account for inverter phase of ~ 1 urad. 5 urad is therefore conservative.
Note on C7: Tom Stewart: 3.7 ppm is the AC-DC diff std unc for 1 V on 10 V range as measured in S21559. 2.3 ppm is the DC gain std unc as reported in S21557.
Note on H7: Laurie Christian: S21559 and S21557 = (3.6+4+6)/2
Note on S7: Laurie Christian: S21559 and S21557 = (3.6+4+6)/2
Note on AD7: Laurie Christian: From S21559 and HP3458CalibrationHistory.xls
Note on C8: Tom Stewart: From S21895 10 ppm is std unc for MSL 0.2 ohm shunt.
Note on C9: Tom Stewart: From S21895
Note on C10: Tom Stewart: 18/5/11 Most of the dial non-linarity is based on the variation of the dial offsets. This is now included as an absolute dial uncertainty (#22 below). The dial linearity is then taken as 2 ppm of the dial setting. This is based on measurements in "IVD Gain Check May11.xls" in E044\Supporting info\2011. 2 ppm 1 sigma is a conservative estimate based on both the difference in gain measured at the top of the range and also the step change seen between 511 and 512. TJS.
Note on H10: Tom Stewart: 18/5/11 Most of the dial non-linarity is based on the variation of the dial offsets. This is now included as an absolute dial uncertainty (#22 below). The dial linearity is then taken as 2 ppm of the dial setting. This is based on measurements in "IVD Gain Check May11.xls" in E044\Supporting info\2011. 2 ppm 1 sigma is a conservative estimate based on both the difference in gain measured at the top of the range and also the step change seen between 511 and 512. TJS.
Note on S10: Laurie Christian: Take 5ppm one sigma down to dial readings of 0.4 and flat 15 ppm one sigma below. Actual curve more complex but this is a conservative estimate. 6/5/11 There is evidence to suggest that the dial linearity issue is complicated by the variation of the IVD offsets. "IVD offset SR830 readings Mar11.xls" shows that there is a variation in the offset, incuding relaxation, depending on what the previous dial setting was. "Dial Scan May11.xls" shows that for the Vars dial, a second measure of linearity after the dials had settled for a time at 0000, shows lineartiy within 5 ppm down to 0029 dial setting as opposed to up to 100 ppm. Current uncertainties are left as the last approved CMC uncertainties, but new uncertainties will be similar to allow for the offset variation affecting low dial settings. TJS.
Note on AD10: Laurie Christian: Take 5ppm one sigma down to dial readings of 0.2 and flat 15 ppm one sigma below. Actual curve more complex but this is a conservative estimate.
Note on AD12: Industrial Research Limited: Seems generous - make it 3 ppm - Andy may have assumed a higher drift rate for new resistors- Temp co is specified as 2 ppm/C so if want a significantly lower uncertainty would have to track temperature
Note on AD13: Industrial Research Limited: Phase offset observed for zero degree measurements is not consistent with measured L1 and L2. Origins unknown so incorporate as an uncertainty at this stage (zero-degree tests.xls) Phase offset is 0.35 mdegree.
Note on C14: Tom Stewart: 18/5/11 Most of the dial non-linarity is based on the variation of the dial offsets. This is now included as an absolute dial uncertainty (#23 below). The dial linearity is then taken as 2 ppm of the dial setting. This is based on measurements in "IVD Gain Check May11.xls" in E044\Supporting info\2011. 2 ppm 1 sigma is a conservative estimate based on both the difference in gain measured at the top of the range and also the step change seen between 511 and 512. TJS.
Note on H14: Tom Stewart: 18/5/11 Most of the dial non-linarity is based on the variation of the dial offsets. This is now included as an absolute dial uncertainty (#23 below). The dial linearity is then taken as 2 ppm of the dial setting. This is based on measurements in "IVD Gain Check May11.xls" in E044\Supporting info\2011. 2 ppm 1 sigma is a conservative estimate based on both the difference in gain measured at the top of the range and also the step change seen between 511 and 512. TJS.
Note on S14: Andrew Corney: increase to 15 ppm until offset problem resolved 6/5/11 See comment on watts dial uncertainty (C10). This uncertainty left as the last approved CMC uncertainty, however there is the possibility to reduce this for high dial settings. TJS.
Note on AD14: Laurie Christian: Take 5ppm one sigma down to dial readings of 0.2 and flat 15 ppm one sigma below. Actual curve more complex but this is a conservative estimate.
Note on D16: Tom Stewart: Change from 30 to 20. Confidence gained from re-evaluating g_1 as part of APMP.EM-K5.1 comparison.
Note on AD16: Industrial Research Limited: No measurements made- leave as is.
Note on C17: Laurie Christian: This includes temperature coefficient of capacitor
Note on H17: Laurie Christian: This includes temperature coefficient of capacitor
Note on S17: Laurie Christian: This includes temperature coefficient of capacitor
Note on AD17: Laurie Christian: This includes 0.1 C and 126 ppm/C temperature coefficient of capacitor. Also includes an additional 12ppm (added in quadrature) due to frequency dependence of C1 (see standard deviation of May results). Also include 45 ppm change suspected to be due to humidity and treat as 2 sigma.
Note on AD18: Industrial Research Limited: Haven't tracked temperature and temp co is 2ppm/C
Note on AD19: Industrial Research Limited: Phase offset observed for zero degree measurements is not consistent with measured L1 and L2. Origins unknown so incorporate as an uncertainty at this stage (zero-degree tests.xls) Phase offset is 0.35 mdegree.
Note on AD20: Industrial Research Limited: Probably wrong since should include a fixed noise term. Is about the right amplitude for 1 V judging by repeatability of RefVolts FFT values
Note on AD21: Industrial Research Limited: Almost certainly an overestimate judging by repeatability of FFT values of Ref Phase
Note on AD23: Industrial Research Limited: Increased because of the change since April 2004
Note on C26: t.stewart: 18/5/11 This is a conservative estimate of the absolute variation of dial setting based on measurements in cell CH22 or "31 March 11" page in "IVD offset SR830 readings.xls" located in E044\Supporting info\IVD offsets Mar11. The variation of the offset depends on the previous setting of the other dial. Added in the UncCalc page as an absolute uncertainty this starts to have impact at low dial settings. TJS
Note on H26: t.stewart: 18/5/11 This is a conservative estimate of the absolute variation of dial setting based on measurements in cell CH22 or "31 March 11" page in "IVD offset SR830 readings.xls" located in E044\Supporting info\IVD offsets Mar11. The variation of the offset depends on the previous setting of the other dial. Added in the UncCalc page as an absolute uncertainty this starts to have impact at low dial settings. TJS
Note on C27: t.stewart: 18/5/11 This is a conservative estimate of the absolute variation of dial setting based on measurements in cell CI22 or "31 March 11" page in "IVD offset SR830 readings.xls" located in E044\Supporting info\IVD offsets Mar11. The variation of the offset depends on the previous setting of the other dial. Added in the UncCalc page as an absolute uncertainty this starts to have impact at low dial settings. TJS
Note on H27: t.stewart: 18/5/11 This is a conservative estimate of the absolute variation of dial setting based on measurements in cell CI22 or "31 March 11" page in "IVD offset SR830 readings.xls" located in E044\Supporting info\IVD offsets Mar11. The variation of the offset depends on the previous setting of the other dial. Added in the UncCalc page as an absolute uncertainty this starts to have impact at low dial settings. TJS
Note on C28: Tom Stewart: This is an estimate of the IVD phase angle uncertainty at full dial setting. Based on phase measurements at the input then output carried out in Jul2013
Note on H28: Tom Stewart: This is an estimate of the IVD phase angle uncertainty at full dial setting. Based on phase measurements at the input then output carried out in Jul2013
Note on C29: Tom Stewart: This is an estimate of the IVD phase angle uncertainty at full dial setting. Based on phase measurements at the input then output carried out in Jul2013
Note on H29: Tom Stewart: This is an estimate of the IVD phase angle uncertainty at full dial setting. Based on phase measurements at the input then output carried out in Jul2013

## TABLE RM11
| Watts | Radian RM-11
row select | Voltage U (V) | Current I (A) | Phase Difference φ (°) | Power Factor | Energised Current Input(s) | Measured Impulses (Imp/(W.h)) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k | watts/vars | IEC phase angle | Comments |  | Power factor rounding | Uncert/CMC
48 | 229.9964 | 0.50002 | 0.0016197 | 1 | A | 100002.022 | 0.002022 | 0.0050581 | 2.0614 | watts | 0.0016197 | 0 |  | 1 | 1.2645
52 | 229.9968 | 0.5 | 59.9992 | 0.5 Lag | A | 99993.6857 | -0.0063143 | 0.0088894 | 2.0406 | watts | 59.9992 | 0 | Lag | 1 | 1.0337
54 | 229.9968 | 0.49999 | -36.8966 | 0.8 Lead | A | 100005.1806 | 0.0051806 | 0.0057979 | 2.0488 | watts | -36.8966 | 0 | Lead | 1 | 1.1254
59 | 229.9969 | 1 | 0.0012966 | 1 | A | 100002.5232 | 0.0025232 | 0.0047128 | 2.0834 | watts | 0.0012966 | 0 |  | 1 | 1.1782
60 | 229.9969 | 0.99999 | 59.9985 | 0.5 Lag | A | 99994.2698 | -0.0057302 | 0.0086367 | 2.0474 | watts | 59.9985 | 0 | Lag | 1 | 1.0043
64 | 229.9969 | 0.99999 | -36.8972 | 0.8 Lead | A | 100006.2852 | 0.0062852 | 0.0057025 | 2.0544 | watts | -36.8972 | 0 | Lead | 1 | 1.1068
13 | 229.9968 | 2.0001 | 0.0025417 | 1 | A | 100003.2208 | 0.0032208 | 0.0047755 | 2.0756 | watts | 0.0025417 | 0 |  | 1 | 1.1939
17 | 229.9967 | 2 | 60.0004 | 0.5 Lag | A | 99996.7223 | -0.0032777 | 0.0086925 | 2.0447 | watts | 60.0004 | 0 | Lag | 1 | 1.0107
19 | 229.9966 | 2 | -36.8664 | 0.8 Lead | A | 100003.6358 | 0.0036358 | 0.0057159 | 2.0531 | watts | -36.8664 | 0 | Lead | 1 | 1.1099
22 | 229.9967 | 5.0002 | 0.001244 | 1 | A | 100002.9853 | 0.0029853 | 0.0047237 | 2.0819 | watts | 0.001244 | 0 |  | 1 | 1.1809
25 | 229.9968 | 5.0001 | 59.9988 | 0.5 Lag | A | 99995.7944 | -0.0042056 | 0.0086098 | 2.0483 | watts | 59.9988 | 0 | Lag | 1 | 1.0012
29 | 229.9967 | 5 | -36.8673 | 0.8 Lead | A | 100006.4108 | 0.0064108 | 0.0056157 | 2.0607 | watts | -36.8673 | 0 | Lead | 1 | 1.0905
85 | 229.9966 | 6.0003 | 0.0020049 | 1 | A | 100002.8067 | 0.0028067 | 0.0047269 | 2.0815 | watts | 0.0020049 | 0 |  | 1 | 1.1817
88 | 229.9968 | 6.0002 | 59.9998 | 0.5 Lag | A | 99998.3939 | -0.0016061 | 0.0086501 | 2.0466 | watts | 59.9998 | 0 | Lag | 1 | 1.0058
91 | 229.9966 | 6 | -36.8669 | 0.8 Lead | A | 100008.0184 | 0.0080184 | 0.005707 | 2.0537 | watts | -36.8669 | 0 | Lead | 1 | 1.1082
94 | 229.9969 | 10.0004 | 0.00095927 | 1 | A | 100001.6436 | 0.0016436 | 0.004711 | 2.0836 | watts | 0.00095927 | 0 |  | 1 | 1.1777
98 | 229.9969 | 10.0001 | 59.9987 | 0.5 Lag | A | 99993.6506 | -0.0063493 | 0.0085107 | 2.0525 | watts | 59.9987 | 0 | Lag | 1 | 0.98966
99 | 229.9968 | 10.0001 | -36.8678 | 0.8 Lead | A | 100005.5057 | 0.0055057 | 0.0056644 | 2.0569 | watts | -36.8678 | 0 | Lead | 1 | 1.0999
115 | 229.9967 | 50.0008 | 0.0007548 | 1 | A | 100003.9288 | 0.0039288 | 0.004763 | 2.0771 | watts | 0.0007548 | 0 |  | 1 | 1.1907
118 | 229.9968 | 49.9992 | 59.9981 | 0.5 Lag | A | 99997.8756 | -0.0021244 | 0.0086206 | 2.0475 | watts | 59.9981 | 0 | Lag | 1 | 1.0025
121 | 229.9969 | 49.9993 | -36.8985 | 0.8 Lead | A | 100006.9411 | 0.0069411 | 0.0057865 | 2.0493 | watts | -36.8985 | 0 | Lead | 1 | 1.1231
123 | 229.9965 | 100.0017 | 0.0004366 | 1 | ABC | 100005.0558 | 0.0050558 | 0.0047358 | 2.0805 | watts | 0.0004366 | 0 |  | 1 | 1.1839
124 | 229.9966 | 99.9978 | 59.9978 | 0.5 Lag | ABC | 99998.2526 | -0.0017474 | 0.0086924 | 2.0458 | watts | 59.9978 | 0 | Lag | 1 | 1.0108
125 | 229.9966 | 99.9985 | -36.8979 | 0.8 Lead | ABC | 100007.4543 | 0.0074543 | 0.0056326 | 2.0593 | watts | -36.8979 | 0 | Lead | 1 | 1.0933
| Vars | Radian RM-11
| Voltage U (V) | Current I (A) | Phase Difference φ (°) | Reactive Power Factor | Energised Current Input(s) | Measured Impulses (Imp/(var.h)) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k |  | IEC phase angle | Comments |  | Power factor rounding
76 | 229.9968 | 0.99996 | 90 | 1 | A | 100002.4671 | 0.0024671 | 0.013795 | 2.083 | vars | 90 | 0 |  | 1 | 1.0611
79 | 229.9968 | 1 | 30.0004 | 0.5 Lag | A | 100000.5493 | 0.00054935 | 0.014776 | 2.0849 | vars | 30.0004 | 0 | Lag | 1 | 0.8692
83 | 229.9967 | 0.99996 | 126.8725 | 0.8 Lead | A | 100000.3486 | 0.00034864 | 0.014006 | 2.0832 | vars | 126.8725 | 0 | Lead | 1 | 1.0004
32 | 229.9969 | 1.9999 | 89.9998 | 1 | A | 100001.4318 | 0.0014318 | 0.013811 | 2.0819 | vars | 89.9998 | 0 |  | 1 | 1.0624
35 | 229.9968 | 2.0001 | 30 | 0.5 Lag | A | 99997.2717 | -0.0027283 | 0.014872 | 2.0807 | vars | 30 | 0 | Lag | 1 | 0.87481
42 | 229.9969 | 2 | 126.8733 | 0.8 Lead | A | 99998.5614 | -0.0014387 | 0.014047 | 2.081 | vars | 126.8733 | 0 | Lead | 1 | 1.0034
38 | 229.9969 | 4.9999 | 89.9998 | 1 | A | 100000.3522 | 0.00035219 | 0.013797 | 2.0829 | vars | 89.9998 | 0 |  | 1 | 1.0613
40 | 229.9968 | 5.0002 | 29.9999 | 0.5 Lag | A | 99999.3589 | -0.00064106 | 0.014758 | 2.0858 | vars | 29.9999 | 0 | Lag | 1 | 0.86811
46 | 229.9966 | 4.9999 | 126.8724 | 0.8 Lead | A | 99999.7967 | -0.00020332 | 0.014003 | 2.0834 | vars | 126.8724 | 0 | Lead | 1 | 1.0002
103 | 229.9968 | 5.9999 | 90.0004 | 1 | A | 100002.1203 | 0.0021203 | 0.013799 | 2.0825 | vars | 90.0004 | 0 |  | 1 | 1.0615
105 | 229.9968 | 6.0003 | 30.0002 | 0.5 Lag | A | 100000.2278 | 0.00022777 | 0.014768 | 2.0854 | vars | 30.0002 | 0 | Lag | 1 | 0.8687
110 | 229.9969 | 5.9999 | 126.8727 | 0.8 Lead | A | 100000.5954 | 0.00059537 | 0.014017 | 2.0829 | vars | 126.8727 | 0 | Lead | 1 | 1.0012
Note on P4: l.christian: First part picks out the 0, 90,180, 270 points and the second part the quadrant- 0 to 90 is EVEN, 90 to 180 is ODD, 180 to 270 is EVEN and 270 to 0 is ODD. Sign of PF handled in col F

## Metrix RM11 History
Watts | Radian RM-11 | [12 empty] | Watts | Radian RM-11
[14 empty] | S21864 | 2013 | [11 empty] | S21785 | 2011 | [10 empty] | S21722 | 2009 | [6 empty] | S21659 | 2007 | [17 empty] | S21588 | 2005 | [12 empty] | Change (%) | Change/EU | New EU/ | S21493 | 2003 | [10 empty] | S21411 | 2001
Voltage U (V) | Current I (A) | Phase Difference φ (°) | Power Factor | Energised Current Input(s) | Measured Impulses (Imp/(W.h)) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k | Ratio of uncertainties | En vs 2013 | En vs 2011 | En vs 2009 | En vs 2007 | Voltage U (V) | Current I (A) | Reactive Power Factor | Energised Current Input(s) | Measured Impulses (Imp/var.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k |  | En vs 2011 | Ratio of uncertainties | En vs 2009 | En vs 2007 | Voltage U (V) | Current I (A) | Power Factor | Energised Current Input(s) | Measured Impulses (Imp/Watt.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k |  | En vs 2009 | Ratio of uncertainties | En vs 2007 | Voltage U (V) | Current I (A) | Power Factor | Energised Current Input(s) | Measured Impulses (Imp/Watt.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k | Voltage U (V) | Current I (A) | Power Factor | Measured Impulses (Imp/Watt.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k |  |  | Watts measurements | [6 empty] | Change % | Change / unc |  | Voltage V (V) | Current I (A) | Power Factor | Energised Phase(s) | Measured Impulses (Imp/Watt.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k |  | Power Bridge phase angle | Comments |  | Power factor rounding |  |  |  | Old EU | Voltage V (V) | Current I (A) | Power Factor | Energised Phase(s) | Measured Impulses (Imp/W.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k |  | 23-2-2001 S21411 | Change | Change / unc | Voltage V (V) | Current I (A) | Power Factor l | Energised Phases | Measured Impulses (Imp/W.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k | Current |  | 2015 | 2013 | 2011 | 2009 | 2007 | 2005 | 2003 | 2001
229.9964 | 0.50002 | 0.0016197 | 1 | A | 100002.022 | 0.002022 | 0.0050581 | 2.0614 | 0.979 | 0.6999 |  | 0.39975 | 0.39975 | 229.9157 | 0.50017 | 1 | A | 99998.4818 | -0.0015182 | 0.0051666 | 2.0893
229.9968 | 0.5 | 59.9992 | 0.5 Lag | A | 99993.6857 | -0.0063143 | 0.0088894 | 2.0406 | 0.98705 | 0.36855 |  | -0.71032 | -0.71032 | 229.9112 | 0.50017 | 0.5 Lag | A | 99990.4095 | -0.0095905 | 0.009006 | 2.0756
229.9968 | 0.49999 | -36.8966 | 0.8 Lead | A | 100005.1806 | 0.0051806 | 0.0057979 | 2.0488 | 0.9552 | 0.53298 |  | 0.89353 | 0.89353 | 229.9052 | 0.50016 | 0.8 Lead | A | 100002.0904 | 0.0020904 | 0.0060698 | 2.0608
229.9969 | 1 | 0.0012966 | 1 | A | 100002.5232 | 0.0025232 | 0.0047128 | 2.0834 | 0.91724 | 0.72092 | 0.73019 | 0.73019 | -48813.4547 | 229.9065 | 0.99968 | 1 | A | 99999.1257 | -0.00087428 | 0.005138 | 2.0931 |  | 0.0085045 | 0.96156 | 0.22486 | 0.071888 | 229.9941 | 0.99976 | 1 | A | 99999.082 | -0.00091797 | 0.0053434 | 2.0722 |  | 0.20804 | 0.92682 | 0.060947 | 230.0489 | 0.99983 | 1 | A | 99997.9704 | -0.0020296 | 0.0057653 | 2.0612 | 230.0062 | 0.99983 | 1 | 99998.7564 | -0.0012436 | 0.0047438 | 2.0834 |  |  |  | -60 |  | 0 |  | 36.9 |  | -0.00024359 | -0.051348 | 12 | 109.9966 | 0.99994 | 1 | A | 100002.1799 | 0.0021799 | 0.0082084 | 2.133 | watts | 0.08846 | ACal All done. 20.90°C 59.26%. Tc = 24.2°C |  | 1 |  | -0.0039316 | -0.47897 | 0.88793 | 110.0127 | 0.99984 | 1 | A | 100006.1115 | 0.0061115 | 0.0092444 | 2.1833 | watts | -0.004 | 0.0061799 | 0.75288 | 110 | 1 | 1 | A | 99995.8846 | -0.0041154 | 0.0086228 | 2.5864 | 0.99976 | 1 | 0.0025232 | -0.00087428 | -0.00091797 | -0.0020296 | -0.0012436 | -0.0010001 | 0.0027779 | -0.006582
229.9969 | 0.99999 | 59.9985 | 0.5 Lag | A | 99994.2698 | -0.0057302 | 0.0086367 | 2.0474 | 0.96735 | 0.18913 | -0.31562 | -0.31562 | -26636.5031 | 229.9053 | 0.99967 | 0.5 Lag | A | 99992.6363 | -0.0073637 | 0.0089282 | 2.0802 |  | -0.48827 | 0.95751 | -0.51955 | -0.15201 | 229.9849 | 0.99973 | 0.5 Lag | A | 99996.9957 | -0.0030043 | 0.0093244 | 2.0612 |  | -0.029951 | 1.0031 | 0.32197 | 230.0452 | 0.99983 | 0.5 Lag | A | 99997.2749 | -0.0027251 | 0.0092954 | 2.0588 | 230.0171 | 0.99977 | 0.5 Lag | 99993.9935 | -0.0060065 | 0.0077705 | 2.086 |  |  | 1 | -0.0057302 | 0.0086367 | 0.0025232 | 0.0047128 | 0.0062852 | 0.0057025 | -0.0010821 | -0.13926 | 13 | 109.9957 | 1 | 0.5 Lag | A | 99990.4612 | -0.0095388 | 0.0099354 | 2.0687 | watts | -59.9115 | 0 | Lag | 1 |  | -0.014539 | -1.4634 | 0.92206 | 110.0128 | 0.9998 | 0.5 lag | A | 100005.0003 | 0.0050003 | 0.010775 | 2.0862 | watts | -0.015 | 0.0054612 | 0.54967 | 110 | 1 | 0.5 lag | A | 99984.8051 | -0.015195 | 0.011739 | 2.1413 | 0.99973 | 0.5 Lag | -0.0057302 | -0.0073637 | -0.0030043 | -0.0027251 | -0.0060065 | -0.0049244 | 0.0072227 | -0.016501
229.9969 | 0.99999 | -36.8972 | 0.8 Lead | A | 100006.2852 | 0.0062852 | 0.0057025 | 2.0544 | 0.94997 | 0.39622 | 1.1514 | 1.1514 | -40339.978 | 229.9052 | 0.99966 | 0.8 Lead | A | 100004.0257 | 0.0040257 | 0.0060028 | 2.0664 |  | 0.71744 | 0.96304 | 1.0544 | 1.1867 | 229.9847 | 0.99975 | 0.8 Lead | A | 99999.719 | -0.00028097 | 0.0062332 | 2.052 |  | 0.32449 | 0.9544 | 0.45192 | 230.0458 | 0.99987 | 0.8 Lead | A | 99997.6964 | -0.0023036 | 0.006531 | 2.0418 | 230.0185 | 0.99984 | 0.8 Lead | 99996.9021 | -0.0030979 | 0.005308 | 2.0483 |  |  | 5 | -0.0042056 | 0.0086098 | 0.0029853 | 0.0047237 | 0.0064108 | 0.0056157 | 0.0016556 | 0.3119 | 14 | 109.9954 | 0.99993 | 0.8 Lead | A | 100004.9413 | 0.0049413 | 0.0084708 | 2.1066 | watts | 36.9911 | 0 | Lead | 1 |  | -0.0039484 | -0.46611 | 0.89325 | 110.013 | 0.99984 | 0.8 lead | A | 100008.8897 | 0.0088897 | 0.0094832 | 2.1482 | watts | 0.002 | 0.0029413 | 0.34723 | 110 | 1 | 0.8 lead | A | 100002.3481 | 0.0023481 | 0.0085927 | 2.4259 | 0.99975 | 0.8 Lead | 0.0062852 | 0.0040257 | -0.00028097 | -0.0023036 | -0.0030979 | -0.0047535 | 1.4552e-14 | -0.0038654
229.9968 | 2.0001 | 0.0025417 | 1 | A | 100003.2208 | 0.0032208 | 0.0047755 | 2.0756 | 0.91984 | 1.1563 |  | 0.67445 | 0.67445 | 229.9199 | 2.0013 | 1 | A | 99997.6988 | -0.0023012 | 0.0051916 | 2.0863 | [34 empty] | 50 | -0.0021244 | 0.0086206 | 0.0039288 | 0.004763 | 0.0069411 | 0.0057865 | [40 empty] | 0 | 0 | 0.0032208 | -0.0023012
229.9967 | 2 | 60.0004 | 0.5 Lag | A | 99996.7223 | -0.0032777 | 0.0086925 | 2.0447 | 0.95973 | 0.61953 |  | -0.37707 | -0.37707 | 229.911 | 2.0013 | 0.5 Lag | A | 99991.3371 | -0.0086629 | 0.0090573 | 2.0746 | [34 empty] | 100 | -0.0017474 | 0.0086924 | 0.0050558 | 0.0047358 | 0.0074543 | 0.0056326 | [40 empty] | 0 | 0 | -0.0032777 | -0.0086629
229.9966 | 2 | -36.8664 | 0.8 Lead | A | 100003.6358 | 0.0036358 | 0.0057159 | 2.0531 | 0.94301 | 0.73844 |  | 0.63608 | 0.63608 | 229.9078 | 2.0014 | 0.8 Lead | A | 99999.415 | -0.00058505 | 0.0060613 | 2.0614 | [81 empty] | 0 | 0 | 0.0036358 | -0.00058505
229.9967 | 5.0002 | 0.001244 | 1 | A | 100002.9853 | 0.0029853 | 0.0047237 | 2.0819 | 0.91986 | 0.89112 | 0.70558 | 0.70558 | -48700.7291 | 229.9031 | 4.9997 | 1 | A | 99998.7759 | -0.0012241 | 0.0051352 | 2.0935 |  | -0.17067 | 1.0004 | 0.012076 | 0.26112 | 229.9823 | 5.0001 | 1 | A | 99999.6524 | -0.00034762 | 0.0051329 | 2.0938 |  | 0.18283 | 1.0256 | 0.43198 | 230.0487 | 5.0007 | 1 | A | 99998.7139 | -0.0012861 | 0.005005 | 2.0792 | 230.0199 | 5.0009 | 1 | 99997.4351 | -0.0025649 | 0.0036531 | 2.0724 | [9 empty] | -0.0008717 | -0.23862 | 15 | 240.0203 | 0.99992 | 1 | A | 99998.9999 | -0.0010001 | 0.0081707 | 2.1378 | watts | 0.08448 | 0 |  | 1 |  | -0.0037779 | -0.46237 | 0.90135 | 240.0555 | 0.99989 | 1 | A | 100002.7779 | 0.0027779 | 0.0090649 | 2.2144 | watts | -0.007 | 0.0059999 | 0.73433 | 240 | 1 | 1 | A | 99993.418 | -0.006582 | 0.0086183 | 2.5878 | 5.0001 | 1 | 0.0029853 | -0.0012241 | -0.00034762 | -0.0012861 | -0.0025649 | -0.0016932 | 0.0027779 | -0.0034127
229.9968 | 5.0001 | 59.9988 | 0.5 Lag | A | 99995.7944 | -0.0042056 | 0.0086098 | 2.0483 | 0.96142 | 0.35828 | -0.25503 | -0.25503 | -26719.9683 | 229.9027 | 4.9998 | 0.5 Lag | A | 99992.7097 | -0.0072903 | 0.0089552 | 2.0795 |  | -0.58965 | 1.0032 | -0.80497 | -0.38235 | 229.9805 | 5.0001 | 0.5 Lag | A | 99997.9901 | -0.0020099 | 0.0089263 | 2.0808 |  | -0.21602 | 1.0049 | 0.20797 | 230.0483 | 5.0006 | 0.5 Lag | A | 99999.9184 | -8.1633e-05 | 0.0088829 | 2.0778 | 230.0213 | 5.0006 | 0.5 Lag | 99996.1337 | -0.0038663 | 0.0072788 | 2.1221 | [9 empty] | -0.0013305 | -0.18279 | 16 | 240.0199 | 1 | 0.5 Lag | A | 99995.0756 | -0.0049244 | 0.010016 | 2.0648 | watts | -59.9173 | 0 | Lag | 1 |  | -0.012147 | -1.2128 | 0.92503 | 240.0536 | 0.99982 | 0.5 lag | A | 100007.2227 | 0.0072227 | 0.010827 | 2.086 | watts | -0.017 | 0.012076 | 1.2057 | 240 | 1 | 0.5 lag | A | 99983.4993 | -0.016501 | 0.0096659 | 2.1978 | 5.0001 | 0.5 Lag | -0.0042056 | -0.0072903 | -0.0020099 | -8.1633e-05 | -0.0038663 | -0.0025359 | 0.0044446 | -0.011834
229.9967 | 5 | -36.8673 | 0.8 Lead | A | 100006.4108 | 0.0064108 | 0.0056157 | 2.0607 | 0.93646 | 0.82974 | 0.96058 | 0.96058 | -40963.4925 | 229.9025 | 4.9997 | 0.8 Lead | A | 100001.7512 | 0.0017512 | 0.0059967 | 2.0668 |  | 0.12252 | 0.99999 | 0.60933 | 0.57674 | 229.982 | 5.0001 | 0.8 Lead | A | 100001.0165 | 0.0010165 | 0.0059968 | 2.0664 |  | 0.4868 | 1.0199 | 0.45421 | 230.0456 | 5.0007 | 0.8 Lead | A | 99998.0972 | -0.0019028 | 0.00588 | 2.0563 | 230.0215 | 5.0008 | 0.8 Lead | 99998.2927 | -0.0017073 | 0.0044263 | 2.0422 | [9 empty] | -0.00086212 | -0.19477 | 17 | 240.0201 | 0.99998 | 0.8 Lead | A | 99995.2465 | -0.0047535 | 0.0084891 | 2.1046 | watts | 36.9865 | 0 | Lead | 1 |  | -0.0047535 | -0.55995 | 0.91266 | 240.054 | 0.99991 | 0.8 lead | A | 100000 | 1.4552e-14 | 0.0093015 | 2.1698 | watts | -0.004 | -0.00075346 | -0.088757 | 240 | 1 | 0.8 lead | A | 99996.1346 | -0.0038654 | 0.0086095 | 2.4324 | 5.0001 | 0.8 Lead | 0.0064108 | 0.0017512 | 0.0010165 | -0.0019028 | -0.0017073 | -0.0008452 | 1.4552e-14 | -0.0025602
229.9966 | 6.0003 | 0.0020049 | 1 | A | 100002.8067 | 0.0028067 | 0.0047269 | 2.0815 | 0.91498 | 0.90449 |  | 0.59377 | 0.59377 | 229.9239 | 5.9981 | 1 | A | 99998.5313 | -0.0014687 | 0.0051661 | 2.0893
229.9968 | 6.0002 | 59.9998 | 0.5 Lag | A | 99998.3939 | -0.0016061 | 0.0086501 | 2.0466 | 0.96395 | 0.57125 |  | -0.18567 | -0.18567 | 229.9119 | 5.9981 | 0.5 Lag | A | 99993.4526 | -0.0065474 | 0.0089736 | 2.0775
229.9966 | 6 | -36.8669 | 0.8 Lead | A | 100008.0184 | 0.0080184 | 0.005707 | 2.0537 | 0.94511 | 0.90829 |  | 1.405 | 1.405 | 229.9116 | 5.998 | 0.8 Lead | A | 100002.8347 | 0.0028347 | 0.0060385 | 2.0633
229.9969 | 10.0004 | 0.00095927 | 1 | A | 100001.6436 | 0.0016436 | 0.004711 | 2.0836 | 0.91564 | 0.85766 |  | 0.34889 | 0.34889 | 229.9046 | 9.993 | 1 | A | 99997.6032 | -0.0023968 | 0.005145 | 2.092
229.9969 | 10.0001 | 59.9987 | 0.5 Lag | A | 99993.6506 | -0.0063493 | 0.0085107 | 2.0525 | 0.94694 | 0.41683 |  | -0.74604 | -0.74604 | 229.9038 | 9.993 | 0.5 Lag | A | 99990.1031 | -0.0098969 | 0.0089876 | 2.0763
229.9968 | 10.0001 | -36.8678 | 0.8 Lead | A | 100005.5057 | 0.0055057 | 0.0056644 | 2.0569 | 0.93393 | 0.73074 |  | 0.972 | 0.972 | 229.8996 | 9.9928 | 0.8 Lead | A | 100001.3666 | 0.0013666 | 0.0060651 | 2.0614
[27 empty] | 229.9775 | 24.9977 | 1 | A | 100000.2208 | 0.00022076 | 0.0054719 | 2.0643 |  | 0.23972 | 0.96084 | 0.25345 | 230.0438 | 24.9616 | 1 | A | 99998.909 | -0.001091 | 0.0056949 | 2.0655 | 230.0158 | 24.9924 | 1 | 99998.8339 | -0.0011661 | 0.0038349 | 2.0547 | [9 empty] | 0.0018284 | 0.47678 | 18 | 109.9984 | 5.0011 | 1 | A | 100000.1416 | 0.00014162 | 0.0077771 | 2.1875 | watts | 0.038557 | 0 |  | 1 |  | -0.0026362 | -0.33898 | 0.85908 | 110.0244 | 5.0009 | 1 | A | 100002.7779 | 0.0027779 | 0.0090528 | 2.2186 | watts | -0.002 | 0.0021416 | 0.27538 | 110 | 5 | 1 | A | 99998.2127 | -0.0017873 | 0.0085986 | 2.5402 | 24.9977 | 1 |  |  | 0.00022076 | -0.001091 | -0.0011661 | -0.0029945 | 0.0033334 | -0.0015441
[27 empty] | 229.9961 | 24.9967 | 0.5 Lag | A | 99998.9361 | -0.0010639 | 0.0093828 | 2.0586 |  | -0.22237 | 1.0327 | 0.36802 | 230.0451 | 24.9605 | 0.5 Lag | A | 100001.0226 | 0.0010226 | 0.0090859 | 2.0657 | 230.0199 | 24.9746 | 0.5 Lag | 99995.4831 | -0.0045169 | 0.0079534 | 2.0916 | [9 empty] | 0.00058422 | 0.073455 | 19 | 109.9986 | 5.0015 | 0.5 Lag | A | 99996.4456 | -0.0035544 | 0.0095451 | 2.0887 | watts | -59.9632 | 0 | Lag | 1 |  | -0.0057767 | -0.6052 | 0.84926 | 110.0221 | 5.0005 | 0.5 lag | A | 100002.2223 | 0.0022223 | 0.011239 | 2.0772 | watts | -0.009 | 0.0054456 | 0.57051 | 110 | 5 | 0.5 lag | A | 99990.7156 | -0.0092844 | 0.0098093 | 2.1753 | 24.9967 | 0.5 Lag |  |  | -0.0010639 | 0.0010226 | -0.0045169 | -0.0051011 | 0.0011111 | -0.013653
[27 empty] | 229.9799 | 24.9977 | 0.8 Lead | A | 99998.4417 | -0.0015583 | 0.0069886 | 2.0488 |  | -0.037112 | 1.1374 | -0.19685 | 230.0446 | 24.9616 | 0.8 Lead | A | 99998.7011 | -0.0012989 | 0.0061442 | 2.045 | 230.0193 | 24.9763 | 0.8 Lead | 99999.8175 | -0.00018253 | 0.0045813 | 2.0334 | [9 empty] | -0.0018203 | -0.39733 | 20 | 109.9983 | 5.0013 | 0.8 Lead | A | 100003.5317 | 0.0035317 | 0.0080464 | 2.1443 | watts | 36.941 | 0 | Lead | 1 |  | 0.00019824 | 0.024638 | 0.8483 | 110.0197 | 5.0008 | 0.8 lead | A | 100003.3334 | 0.0033334 | 0.0094853 | 2.1484 | watts | 0.001 | 0.0025317 | 0.31464 | 110 | 5 | 0.8 lead | A | 100001.3601 | 0.0013601 | 0.0085966 | 2.4351 | 24.9977 | 0.8 Lead |  |  | -0.0015583 | -0.0012989 | -0.00018253 | 0.0016378 | 0.00055556 | -0.00013426
229.9967 | 50.0008 | 0.0007548 | 1 | A | 100003.9288 | 0.0039288 | 0.004763 | 2.0771 | 0.91638 | 1.6058 | 1.371 | 1.371 | -48296.2563 | 229.9059 | 50.0511 | 1 | A | 99996.2802 | -0.0037198 | 0.0051976 | 2.0855 |  | -0.21523 | 0.62279 | -0.50672 | 0.54259 | 229.9817 | 50.001 | 1 | A | 99997.3989 | -0.0026011 | 0.0083457 | 2.2077 |  | -0.18154 | 1.0115 | 0.47196 | 230.0379 | 50.0042 | 1 | A | 99998.914 | -0.001086 | 0.0082507 | 2.1276 | 230.0263 | 50.0488 | 1 | 99993.46 | -0.00654 | 0.0048288 | 2.0714 | [9 empty] | 0.015795 | 3.2711 | 21 | 240.0225 | 5.0013 | 1 | A | 99998.3068 | -0.0016932 | 0.0077666 | 2.1898 | watts | 0.03393 | 0 |  | 1 |  | -0.0044711 | -0.57568 | 0.85779 | 240.0697 | 5.0008 | 1 | A | 100002.7779 | 0.0027779 | 0.0090542 | 2.2181 | watts | -0.003 | 0.0013068 | 0.16826 | 240.1 | 5 | 1 | A | 99996.5873 | -0.0034127 | 0.0086219 | 2.587 | 50.001 | 1 |  | -0.0037198 | -0.0026011 | -0.001086 | -0.00654 | -0.022335 | -0.012221 | -0.011087
229.9968 | 49.9992 | 59.9981 | 0.5 Lag | A | 99997.8756 | -0.0021244 | 0.0086206 | 2.0475 | 0.9379 | 0.56243 | 0.16252 | 0.16252 | -26684.8059 | 229.8978 | 50.0492 | 0.5 Lag | A | 99993.0271 | -0.0069729 | 0.0091914 | 2.067 |  | -0.37507 | 0.67301 | -0.9856 | -0.54367 | 229.9795 | 49.9999 | 0.5 Lag | A | 99996.4745 | -0.0035255 | 0.013657 | 2.1668 |  | -0.41089 | 1.0132 | -0.11347 | 230.0375 | 50.0035 | 0.5 Lag | A | 100002.0861 | 0.0020861 | 0.01348 | 2.1426 | 230.0237 | 50.0396 | 0.5 Lag | 99998.0242 | -0.0019758 | 0.0079125 | 2.0769 | [9 empty] | 0.0080785 | 1.021 | 22 | 240.0234 | 5.0015 | 0.5 Lag | A | 99997.4641 | -0.0025359 | 0.0095394 | 2.0891 | watts | -59.9674 | 0 | Lag | 1 |  | -0.0069805 | -0.73175 | 0.88257 | 240.0666 | 5.0005 | 0.5 lag | A | 100004.4446 | 0.0044446 | 0.010809 | 2.0873 | watts | -0.012 | 0.0094641 | 0.99211 | 240 | 5 | 0.5 lag | A | 99988.1665 | -0.011834 | 0.0096765 | 2.1938 | 49.9999 | 0.5 Lag |  | -0.0069729 | -0.0035255 | 0.0020861 | -0.0019758 | -0.010054 | 0.0055559 | -0.010134
229.9969 | 49.9993 | -36.8985 | 0.8 Lead | A | 100006.9411 | 0.0069411 | 0.0057865 | 2.0493 | 0.93613 | 0.50923 | 0.88074 | 0.88074 | -39752.8017 | 229.8926 | 50.0482 | 0.8 Lead | A | 100003.9944 | 0.0039944 | 0.0061814 | 2.0541 |  | 0.34778 | 0.63846 | 1.3642 | 1.0787 | 229.9714 | 50.0001 | 0.8 Lead | A | 100001.8446 | 0.0018446 | 0.0096817 | 2.1775 |  | 0.64896 | 1.0219 | 0.46669 | 230.0383 | 50.0082 | 0.8 Lead | A | 99995.5615 | -0.0044385 | 0.0094743 | 2.1304 | 230.0217 | 50.0458 | 0.8 Lead | 99997.3263 | -0.0026737 | 0.0056786 | 2.0271 | [9 empty] | 0.010371 | 1.8263 | 23 | 240.0222 | 5.0014 | 0.8 Lead | A | 99999.1548 | -0.0008452 | 0.0080473 | 2.1442 | watts | 36.936 | 0 | Lead | 1 |  | -0.0008452 | -0.10503 | 0.86619 | 240.0612 | 5.0007 | 0.8 lead | A | 100000 | 1.4552e-14 | 0.0092904 | 2.1732 | watts | -0.003 | 0.0021548 | 0.26777 | 240 | 5 | 0.8 lead | A | 99997.4398 | -0.0025602 | 0.0086183 | 2.3974 | 50.0001 | 0.8 Lead |  | 0.0039944 | 0.0018446 | -0.0044385 | -0.0026737 | -0.013045 | -0.01111 | -0.00042117
229.9965 | 100.0017 | 0.0004366 | 1 | ABC | 100005.0558 | 0.0050558 | 0.0047358 | 2.0805 | 0.91519 | 0.84595 | 0.45987 | 0.45987 | -48573.0873 | 229.8573 | 100.0317 | 1 | A B C | 100001.0496 | 0.0010496 | 0.0051746 | 2.0883 |  | -0.35333 | 0.5144 | 0.59773 | 0.19495 | 230.0809 | 99.8925 | 1 | A B C | 100002.8779 | 0.0028779 | 0.01006 | 2.1447 |  | 0.48923 | 1.2309 | 0.28204 | 230.0365 | 99.9764 | 1 | A B C | 99997.9565 | -0.0020435 | 0.0081726 | 2.1378 | 230.0138 | 100.0185 | 1 | 100000.0408 | 4.0752e-05 | 0.00473 | 2.0852 | [9 empty] | 0.0015992 | 0.3381 | 38 | 109.9987 | 24.9933 | 1 | A | 100001.1409 | 0.0011409 | 0.0088138 | 2.1001 | watts | -0.089372 | Repeat row 32 |  | 1 |  | -0.001637 | -0.18573 | 0.92033 | 110.0134 | 24.9689 | 1 | A | 100002.7779 | 0.0027779 | 0.0095768 | 2.1551 | watts | -0.002 | 0.0031409 | 0.35636 | 110 | 24.97 | 1 | A | 99997.6678 | -0.0023322 | 0.0085948 | 2.4827 | 99.8925 | 1 |  |  | 0.0028779 | -0.0020435 | 4.0752e-05 | -0.0015584 | 0.0016667 | -0.006127
229.9966 | 99.9978 | 59.9978 | 0.5 Lag | ABC | 99998.2526 | -0.0017474 | 0.0086924 | 2.0458 | 0.93979 | 0.99865 | 0.95778 | 0.95778 | -26464.269 | 229.8548 | 100.0331 | 0.5 Lag | A B C | 99989.5719 | -0.010428 | 0.0092493 | 2.0663 |  | -0.038416 | 0.66635 | -0.35874 | -0.55682 | 229.9958 | 100.1061 | 0.5 Lag | A B C | 99989.9272 | -0.010073 | 0.013881 | 2.1478 |  | -0.21345 | 1.0313 | -0.34544 | 230.0367 | 99.9689 | 0.5 Lag | A B C | 99992.89 | -0.00711 | 0.01346 | 2.1444 | 230.0201 | 100.006 | 0.5 Lag | 99994.7221 | -0.0052779 | 0.0077718 | 2.0863 | [9 empty] | 0.0074211 | 0.95488 | 33 | 109.9979 | 24.9913 | 0.5 Lag | A | 99990.7424 | -0.0092576 | 0.010529 | 2.0539 | watts | -60.1068 | 0 | Lag | 1 |  | -0.012035 | -1.1431 | 1.0012 | 110.0144 | 24.966 | 0.5 lag | A | 100002.7779 | 0.0027779 | 0.010517 | 2.1039 | watts | -0.009 | -0.00025761 | -0.024468 | 110 | 24.98 | 0.5 lag | A | 99991.4367 | -0.0085633 | 0.01057 | 2.1347 | 100.1061 | 0.5 Lag |  |  | -0.010073 | -0.00711 | -0.0052779 | -0.012699 | -0.0022222 | -0.019034
229.9966 | 99.9985 | -36.8979 | 0.8 Lead | ABC | 100007.4543 | 0.0074543 | 0.0056326 | 2.0593 | 0.92928 | 0.81675 | 0.96545 | 0.96545 | -40837.9402 | 229.8525 | 100.0312 | 0.8 Lead | A B C | 100002.8539 | 0.0028539 | 0.0060612 | 2.0616 |  | 0.13819 | 0.52246 | 0.76924 | 0.9417 | 229.9978 | 100.0692 | 0.8 Lead | A B C | 100002.0163 | 0.0020163 | 0.011601 | 2.1373 |  | 0.3297 | 1.229 | 0.4198 | 230.0295 | 99.9721 | 0.8 Lead | A B C | 99998.1914 | -0.0018086 | 0.0094394 | 2.1374 | 230.0162 | 100.0132 | 0.8 Lead | 99997.1461 | -0.0028539 | 0.005342 | 2.0457 | [9 empty] | 0.0011377 | 0.21297 | 34 | 109.9974 | 24.9898 | 0.8 Lead | A | 100005.1587 | 0.0051587 | 0.0088789 | 2.0832 | watts | 36.7954 | 0 | Lead | 1 |  | 0.00071405 | 0.08042 | 0.92559 | 110.0129 | 24.9681 | 0.8 lead | A | 100004.4446 | 0.0044446 | 0.0095927 | 2.1402 | watts | 0 | 0.0051587 | 0.581 | 110 | 24.97 | 0.8 lead | A | 99999.946 | -5.3983e-05 | 0.008606 | 2.4148 | 100.0692 | 0.8 Lead |  |  | 0.0020163 | -0.0018086 | -0.0028539 | -0.0039916 | 0.0066671 | -0.0047038
[65 empty] | 35 | 240.0221 | 24.9895 | 1 | A | 99997.0055 | -0.0029945 | 0.0082541 | 2.1277 | watts | -0.10988 | 0 |  | 1 |  | -0.0063279 | -0.76664 | 0.89382 | 240.0756 | 24.9751 | 1 | A | 100003.3334 | 0.0033334 | 0.0092346 | 2.184 | watts | -0.002 | -0.00099449 | -0.12049 | 240 | 24.97 | 1 | A | 99998.4559 | -0.0015441 | 0.0086038 | 2.561
[65 empty] | 36 | 240.0227 | 24.9896 | 0.5 Lag | A | 99994.8989 | -0.0051011 | 0.010271 | 2.0578 | watts | -60.1152 | 0 | Lag | 1 |  | -0.0062123 | -0.60486 | 0.94933 | 240.075 | 24.9685 | 0.5 lag | A | 100001.1111 | 0.0011111 | 0.010819 | 2.0869 | watts | -0.014 | 0.0088989 | 0.86644 | 240 | 24.98 | 0.5 lag | A | 99986.3467 | -0.013653 | 0.0097079 | 2.187
[65 empty] | 37 | 240.0219 | 24.989 | 0.8 Lead | A | 100001.6378 | 0.0016378 | 0.0085113 | 2.1026 | watts | 36.7854 | 0 | Lead | 1 |  | 0.0010822 | 0.12715 | 0.8975 | 240.0856 | 24.9823 | 0.8 lead | A | 100000.5556 | 0.00055556 | 0.0094834 | 2.1481 | watts | 0 | 0.0016378 | 0.19242 | 240 | 24.97 | 0.8 lead | A | 99999.8657 | -0.00013426 | 0.0087975 | 2.3031
[27 empty] | 50 A 0.8 Lead Watts | 50 A Unity Watts |  | 100 A Unity Vars | 100 A 0.5 Lag Vars | 5 A 0.5 Lag Watts | 5 A 0.5 Lag Vars | [6 empty] | 50 A 0.8 Lead Watts | 50 A Unity Watts | 100 A Unity Vars | 100 A 0.5 Lag Vars | 25 A 0.5 Lag Watts | 50 A 0.5 Lag Vars | [19 empty] | 43 | 109.9984 | 50.0438 | 1 | A | 100001.3892 | 0.0013892 | 0.0083334 | 2.1199 | watts | 0.0023179 | 0 |  | 1 |  | -0.0019442 | -0.2333 | 0.79399 | 110.0138 | 50.0099 | 1 | A | 100003.3334 | 0.0033334 | 0.010496 | 2.1392 | watts | 0 | 0.0013892 | 0.16671 | 110 | 50 | 1 | A | 100000.4877 | 0.00048774 | 0.0085868 | 2.4831
[26 empty] | 2001 | -0.00042117 | -0.011087 | 0.008602 | -0.0362 | -0.040269 | -0.016501 | -0.012837 | [5 empty] | 2001 | -0.00042117 | -0.011087 | -0.0362 | -0.040269 | -0.013653 | -0.026048 | [19 empty] | 44 | 109.9967 | 50.0473 | 0.5 Lag | A | 99992.6421 | -0.0073579 | 0.010136 | 2.0595 | watts | -59.9989 | 0 | Lag | 1 |  | -0.011802 | -1.1644 | 0.89189 | 110.0145 | 49.9967 | 0.5 lag | A | 100004.4446 | 0.0044446 | 0.011365 | 2.0755 | watts | -0.007 | -0.00035786 | -0.035305 | 110 | 50.01 | 0.5 lag | A | 99993.0468 | -0.0069532 | 0.0097746 | 2.1773
[23 empty] | -60 | 0 | 36.9 | 2003 | -0.01111 | -0.012221 | 0.0099577 | 0.021116 | 0.011112 | 0.0072227 | 0.011112 |  |  | -60 | 0 | 36.9 | 2003 | -0.01111 | -0.012221 | 0.021116 | 0.011112 | 0.0011111 | 0.010001 | [19 empty] | 45 | 109.9962 | 50.0497 | 0.8 Lead | A | 100007.8726 | 0.0078726 | 0.0085811 | 2.097 | watts | 36.8862 | 0 | Lead | 1 |  | 0.0023168 | 0.26999 | 0.92438 | 110.0119 | 50.0003 | 0.8 lead | A | 100005.5559 | 0.0055559 | 0.009283 | 2.1741 | watts | 0.001 | 0.0068726 | 0.80091 | 110 | 50.01 | 0.8 lead | A | 100001.4304 | 0.0014304 | 0.008636 | 2.3788
[22 empty] | 1 | -0.0073637 | -0.00087428 | 0.0040257 | 2005 | -0.013045 | -0.022335 | 0.0082148 | -0.0217 | -0.028025 | -0.0049244 | -0.011383 |  | 1 | -0.0073637 | -0.00087428 | 0.0040257 | 2005 | -0.013045 | -0.022335 | -0.0217 | -0.028025 | -0.0051011 | -0.0037983 | [19 empty] | 46 | 240.0195 | 50.0518 | 1 | A | 99977.6646 | -0.022335 | 0.0082148 | 2.1324 | watts | -0.022937 | 0 |  | 1 |  | -0.010115 | -1.2313 | 0.82497 | 240.0513 | 49.9996 | 1 | A | 99987.7793 | -0.012221 | 0.0099577 | 2.1419 | watts | -0.011 | -0.011335 | -1.3799 | 240 | 50 | 1 | A | 99988.9132 | -0.011087 | 0.008602 | 2.4861
[22 empty] | 5 | -0.0072903 | -0.0012241 | 0.0017512 | 2007 | -0.0026737 | -0.00654 | 0.0048288 | 0.00023728 | -0.0084893 | -0.0038663 | -0.00025863 |  | 5 | -0.0072903 | -0.0012241 | 0.0017512 | 2007 | -0.0026737 | -0.00654 | 0.00023728 | -0.0084893 | -0.0045169 | -0.0033672 | [19 empty] | 47 | 240.0197 | 50.0499 | 0.5 Lag | A | 99989.9457 | -0.010054 | 0.00996 | 2.0679 | watts | -60.0145 | 0 | Lag | 1 |  | -0.01561 | -1.5673 | 0.94472 | 240.0531 | 49.9953 | 0.5 lag | A | 100005.5559 | 0.0055559 | 0.010543 | 2.1026 | watts | -0.01 | -5.4346e-05 | -0.0054565 | 240 | 50.01 | 0.5 lag | A | 99989.8659 | -0.010134 | 0.010077 | 2.1518
[22 empty] | 25 | -0.0037198 | -0.0037198 | 0.0039944 | 2009 | -0.0044385 | -0.001086 | 0.0082507 | -0.010992 | -0.019189 | -8.1633e-05 | -0.004402 |  | 50 | -0.0069729 | -0.0037198 | 0.0039944 | 2009 | -0.0044385 | -0.001086 | -0.010992 | -0.019189 | 0.0010226 | -0.014509 | [19 empty] | 48 | 240.0179 | 50.049 | 0.8 Lead | A | 99986.9555 | -0.013045 | 0.008499 | 2.1034 | watts | 36.8843 | 0 | Lead | 1 |  | -0.0019346 | -0.22763 | 0.91497 | 240.0482 | 49.9997 | 0.8 lead | A | 99988.8901 | -0.01111 | 0.0092889 | 2.1742 | watts | 0 | -0.013045 | -1.5348 | 240 | 50.01 | 0.8 lead | A | 99999.5788 | -0.00042117 | 0.0086005 | 2.4297
[22 empty] | 50 | -0.0069729 | -0.0037198 | 0.0039944 | 2011 | 0.0018446 | -0.0026011 | 0.0083457 | -0.025765 | -0.030112 | -0.0020099 | 0.0017275 |  | 100 | -0.010428 | 0.0010496 | 0.0028539 | [26 empty] | 57 | 109.9971 | 100.0189 | 1 | A B C | 100000.2554 | 0.0002554 | 0.0082008 | 2.134 | watts | 0.078524 | Problems with excel spreadsheet so lost |  | 1 |  | -0.0025225 | -0.30759 | 0.90577 | 110.0057 | 99.9761 | 1 | A B C | 100002.7779 | 0.0027779 | 0.0090539 | 2.2186 | watts | 0.002 | -0.0017446 | -0.21274 | 110 | 99.97 | 1 | A B C | 100002.083 | 0.002083 | 0.0086032 | 2.4445
[22 empty] | 100 | -0.010428 | 0.0010496 | 0.0028539 | 2013 | 0.0039944 | -0.0037198 | 0.0051976 | -0.0061014 | -0.016115 | -0.0072903 | -0.0052896 | [31 empty] | 58 | 109.9973 | 100.0273 | 0.5 Lag | A B C | 99988.5451 | -0.011455 | 0.0099345 | 2.0689 | watts | -59.9236 | some data with bad readings at 0° & - 36° | Lag | 1 |  | -0.013677 | -1.3767 | 0.92474 | 110.0103 | 99.966 | 0.5 lag | A B C | 100002.2223 | 0.0022223 | 0.010743 | 2.0906 | watts | -0.011 | -0.00045487 | -0.045786 | 110 | 99.98 | 0.5 lag | A B C | 99988.6913 | -0.011309 | 0.010027 | 2.1526
[26 empty] | 2015 | 0.0069411 | 0.0039288 | 0.004763 |  |  | -0.0042056 | -0.00064106 | [31 empty] | 59 | 109.9956 | 100.0202 | 0.8 Lead | A B C | 100004.0539 | 0.0040539 | 0.008511 | 2.1028 | watts | 36.9781 | on 100A & 240V. Needed to change the | Lead | 1 |  | -0.0031688 | -0.37232 | 0.88591 | 110.0054 | 99.9741 | 0.8 lead | A B C | 100007.2227 | 0.0072227 | 0.0096071 | 2.1368 | watts | 0 | 0.0040539 | 0.47631 | 110 | 99.98 | 0.8 lead | A B C | 100000.241 | 0.00024105 | 0.0085947 | 2.4151
[65 empty] | 60 | 240.0184 | 100.0194 | 1 | A B C | 99998.4416 | -0.0015584 | 0.0081659 | 2.1385 | watts | 0.070749 | counter trigger level from 9.2V to 8.2V. |  | 1 |  | -0.0032251 | -0.39495 | 0.8742 | 240.07 | 99.9846 | 1 | A B C | 100001.6667 | 0.0016667 | 0.009341 | 2.172 | watts | -0.006 | 0.0044416 | 0.54392 | 240 | 99.97 | 1 | A B C | 99993.873 | -0.006127 | 0.0086045 | 2.5537
[65 empty] | 61 | 240.0219 | 100.0268 | 0.5 Lag | A B C | 99987.301 | -0.012699 | 0.0099314 | 2.0691 | watts | -59.9299 | 0 | Lag | 1 |  | -0.010477 | -1.0549 | 0.92399 | 240.0687 | 99.9711 | 0.5 lag | A B C | 99997.7778 | -0.0022222 | 0.010748 | 2.09 | watts | -0.019 | 0.006301 | 0.63445 | 240 | 99.99 | 0.5 lag | A B C | 99980.9658 | -0.019034 | 0.0099103 | 2.1652
[65 empty] | 62 | 240.0195 | 100.0262 | 0.8 Lead | A B C | 99996.0084 | -0.0039916 | 0.0084962 | 2.1039 | watts | 36.9735 | 0 | Lead | 1 |  | -0.010659 | -1.2545 | 0.82068 | 240.0631 | 99.9871 | 0.8 lead | A B C | 100006.6671 | 0.0066671 | 0.010353 | 2.0805 | watts | -0.005 | 0.0010084 | 0.11869 | 240 | 99.98 | 0.8 lead | A B C | 99995.2962 | -0.0047038 | 0.0086194 | 2.4039
Vars | Radian RM-11 | [12 empty] | Vars | Radian RM-11 | [50 empty] | Vars | Radian RM-11
[80 empty] | Change (%) | Change/ | New EU/ | S21493
Voltage U (V) | Current I (A) | Phase Difference φ (°) | Reactive Power Factor | Energised Current Input(s) | Measured Impulses (Imp/(var.h)) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k | Ratio of uncertainties | En vs 2013 | En vs 2011 | En vs 2009 | En vs 2007 | Voltage U (V) | Current I (A) | Reactive Power Factor | Energised Current Input(s) | Measured Impulses (Imp/var.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k |  | En vs 2011 | Ratio of uncertainties | En vs 2009 | En vs 2007 | Voltage U (V) | Current I (A) | Reactive Power Factor | Energised Current Input(s) | Measured Impulses (Imp/Var.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k |  | EN vs 2009 | Ratio of uncertainties | En vs 2007 | Voltage U (V) | Current I (A) | Reactive Power Factor | Energised Current Input(s) | Measured Impulses (Imp/Var.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k | Voltage U (V) | Current I (A) | Reactive Power Factor | Measured Impulses (Imp/Var.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k |  |  | Vars measurements | [6 empty] | Change % | Change / unc |  | Voltage V (V) | Current I (A) | Var Factor | Energised Phase(s) | Measured Impulses (Imp/Var.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k |  | Power Bridge phase angle | Comments |  | Power factor rounding | Fudge the power factor to positive |  | Exp Uncert | Old EU | Voltage V (V) | Current I (A) | Var Factor | Energised Phase(s) | Measured Impulses (Imp/var.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k |  | 23-2-2001 S21411 | Change | Change / unc | Voltage V (V) | Current I (A) | Power Factor l | Energised Phases | Measured Impulses (Imp/var.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k
229.9968 | 0.99996 | 90 | 1 | A | 100002.4671 | 0.0024671 | 0.013795 | 2.083 | 0.95412 | 0.60659 | -7248750.8543 | 0.016525 | -16676.3289 | 229.9184 | 0.99965 | 1 | A | 99994.0994 | -0.0059006 | 0.014458 | 2.1604 |  | -0.56299 | 0.99601 | -0.57776 | -0.71715 | 230.0255 | 0.99974 | 1 | A | 100002.2391 | 0.0022391 | 0.014516 | 2.1542 |  | -0.014717 | 1.1197 | -0.15355 | 230.0469 | 0.99981 | 1 | A | 100002.4527 | 0.0024527 | 0.012964 | 2.2144 | 230.0196 | 0.99974 | 1 | 100004.468 | 0.004468 | 0.0078082 | 2.1469 |  |  |  | 90 |  | 30 |  | 127 |  | 0.0127 | 1.6265 | 24 | 109.9972 | 0.99998 | 1 | A | 99995.4765 | -0.0045235 | 0.015892 | 2.138 | vars | -89.9096 | 0 |  | 1 | -1 | -0.023972 | -1.5084 | 1.1404 | 110.0131 | 0.99978 | 1 | A | 100019.4482 | 0.019448 | 0.013936 | 2.1648 | vars | -0.012 | 0.0074765 | 0.47045 | 110 | 1 | 1 | A | 99987.6483 | -0.012352 | 0.013375 | 2.2429 | 0.99974 | 1 |  | -0.0059006 | 0.0022391 | 0.0024527 | 0.004468 | -0.0082323 | 0.016669 | -0.016509
229.9968 | 1 | 30.0004 | 0.5 Lag | A | 100000.5493 | 0.00054935 | 0.014776 | 2.0849 | 0.9195 | 0.33503 | -6767286.5486 | 0.010841 | -15568.5599 | 229.9133 | 0.99967 | 0.5 Lag | A | 99995.5989 | -0.0044011 | 0.01607 | 2.1056 |  | -0.29809 | 0.98635 | -0.086629 | -0.45357 | 230.0045 | 0.99976 | 0.5 Lag | A | 100000.3892 | 0.00038916 | 0.016292 | 2.0961 |  | 0.20857 | 1.1158 | -0.15336 | 230.0466 | 0.99985 | 0.5 Lag | A | 99996.991 | -0.003009 | 0.014601 | 2.1075 | 230.0215 | 0.99982 | 0.5 Lag | 100002.8878 | 0.0028878 | 0.0097525 | 2.0694 |  |  | 0.99996 | 0.0024671 | 0.013795 | 0.00054935 | 0.014776 | 0.00034864 | 0.014006 | 0.014271 | 1.4633 | 25 | 109.9982 | 0.99996 | 0.5 Lag | A | 99994.53 | -0.00547 | 0.016757 | 2.0981 | vars | -29.9133 | 0 | Lag | 1 | -1 | -0.024362 | -1.4539 | 1.1149 | 110.0137 | 0.99986 | 0.5 lag | A | 100018.8925 | 0.018892 | 0.015031 | 2.1019 | vars | -0.008 | 0.00253 | 0.15098 | 110 | 1 | 0.5 lag | A | 99991.7034 | -0.0082966 | 0.014726 | 2.1224 | 0.99976 | 0.5 Lag |  | -0.0044011 | 0.00038916 | -0.003009 | 0.0028878 | -0.011383 | 0.011112 | -0.012141
229.9967 | 0.99996 | 126.8725 | 0.8 Lead | A | 100000.3486 | 0.00034864 | 0.014006 | 2.0832 | 0.94848 | 0.55605 |  |  |  | 229.9111 | 0.99965 | 0.8 Lead | A | 99992.5605 | -0.0074395 | 0.014767 | 2.145 | [34 empty] | 1.9999 | 0.0014318 | 0.013811 | -0.0027283 | 0.014872 | -0.0014387 | 0.014047
229.9969 | 1.9999 | 89.9998 | 1 | A | 100001.4318 | 0.0014318 | 0.013811 | 2.0819 | 0.95402 | 0.59392 |  |  |  | 229.9143 | 2.0013 | 1 | A | 99993.2293 | -0.0067707 | 0.014477 | 2.1588 | [34 empty] | 4.9999 | 0.00035219 | 0.013797 | -0.00064106 | 0.014758 | -0.00020332 | 0.014003
229.9968 | 2.0001 | 30 | 0.5 Lag | A | 99997.2717 | -0.0027283 | 0.014872 | 2.0807 | 0.92385 | 0.49779 |  |  |  | 229.9126 | 2.0014 | 0.5 Lag | A | 99989.8686 | -0.010131 | 0.016098 | 2.1042 | [34 empty] | 5.9999 | 0.0021203 | 0.013799 | 0.00022777 | 0.014768 | 0.00059537 | 0.014017
229.9969 | 2 | 126.8733 | 0.8 Lead | A | 99998.5614 | -0.0014387 | 0.014047 | 2.081 | 0.94929 | 0.54381 |  |  |  | 229.9109 | 2.0013 | 0.8 Lead | A | 99990.9222 | -0.0090778 | 0.014798 | 2.1429
229.9969 | 4.9999 | 89.9998 | 1 | A | 100000.3522 | 0.00035219 | 0.013797 | 2.0829 | 0.95447 | 0.51437 | -7247257.1375 | -0.21807 | -16673.2103 | 229.9171 | 4.9998 | 1 | A | 99993.2553 | -0.0067447 | 0.014456 | 2.1603 |  | -0.69909 | 0.9999 | -0.41898 | -0.56487 | 229.9988 | 5.0001 | 1 | A | 100003.3609 | 0.0033609 | 0.014457 | 2.1605 |  | 0.28008 | 1.1145 | 0.13421 | 230.0472 | 5.0007 | 1 | A | 99999.3118 | -0.00068818 | 0.012971 | 2.2156 | 230.0217 | 5.0004 | 1 | 100001.4207 | 0.0014207 | 0.007396 | 2.2092 | [9 empty] | 0.010737 | 1.4517 | 26 | 240.0222 | 0.99998 | 1 | A | 99991.7677 | -0.0082323 | 0.015912 | 2.1381 | vars | -89.9157 | 0 |  | 1 | -1 | -0.024902 | -1.565 | 1.1402 | 240.0636 | 0.99978 | 1 | A | 100016.6694 | 0.016669 | 0.013955 | 2.1662 | vars | -0.017 | 0.0087677 | 0.55103 | 240.1 | 1 | 1 | A | 99983.4913 | -0.016509 | 0.013404 | 2.2434 | 5.0001 | 1 |  | -0.0067447 | 0.0033609 | -0.00068818 | 0.0014207 | -0.0093162 | 0.013891 | -0.017693
229.9968 | 5.0002 | 29.9999 | 0.5 Lag | A | 99999.3589 | -0.00064106 | 0.014758 | 2.0858 | 0.91926 | 0.31499 | -6775645.4304 | -0.1605 | -15588.0596 | 229.9119 | 4.9998 | 0.5 Lag | A | 99994.7104 | -0.0052896 | 0.016054 | 2.1064 |  | -0.43709 | 0.99996 | -0.055287 | -0.31338 | 230.0222 | 4.9998 | 0.5 Lag | A | 100001.7275 | 0.0017275 | 0.016055 | 2.1064 |  | 0.38179 | 1.1044 | 0.12371 | 230.0473 | 5.0007 | 0.5 Lag | A | 99995.598 | -0.004402 | 0.014538 | 2.1094 | 230.0207 | 5.0008 | 0.5 Lag | 99999.7414 | -0.00025863 | 0.0092876 | 2.092 | [9 empty] | 0.0061556 | 0.66278 | 27 | 240.0231 | 0.99996 | 0.5 Lag | A | 99988.6169 | -0.011383 | 0.016722 | 2.0983 | vars | -29.9183 | 0 | Lag | 1 | -1 | -0.022495 | -1.3452 | 1.1247 | 240.0605 | 0.99986 | 0.5 lag | A | 100011.1123 | 0.011112 | 0.014868 | 2.1082 | vars | -0.012 | 0.00061686 | 0.036889 | 240 | 1 | 0.5 lag | A | 99987.8588 | -0.012141 | 0.014358 | 2.1408 | 4.9998 | 0.5 Lag |  | -0.0052896 | 0.0017275 | -0.004402 | -0.00025863 | -0.0064142 | 0.011112 | -0.012837
229.9966 | 4.9999 | 126.8724 | 0.8 Lead | A | 99999.7967 | -0.00020332 | 0.014003 | 2.0834 | 0.94818 | 0.44749 |  |  |  | 229.9074 | 4.9996 | 0.8 Lead | A | 99993.5303 | -0.0064697 | 0.014769 | 2.145
229.9968 | 5.9999 | 90.0004 | 1 | A | 100002.1203 | 0.0021203 | 0.013799 | 2.0825 | 0.95253 | 0.48442 |  |  |  | 229.9101 | 5.998 | 1 | A | 99995.4358 | -0.0045642 | 0.014487 | 2.1583
229.9968 | 6.0003 | 30.0002 | 0.5 Lag | A | 100000.2278 | 0.00022777 | 0.014768 | 2.0854 | 0.91714 | 0.39028 |  |  |  | 229.9103 | 5.9981 | 0.5 Lag | A | 99994.4642 | -0.0055358 | 0.016102 | 2.1048
229.9969 | 5.9999 | 126.8727 | 0.8 Lead | A | 100000.5954 | 0.00059537 | 0.014017 | 2.0829 | 0.94831 | 0.43693 |  |  |  | 229.9076 | 5.9979 | 0.8 Lead | A | 99994.4711 | -0.0055289 | 0.014781 | 2.1441
[27 empty] | 229.9968 | 24.9961 | 1 | A | 100002.1932 | 0.0021932 | 0.01452 | 2.1552 |  | 0.23234 | 1.117 | 0.26102 | 230.0461 | 24.9623 | 1 | A | 99998.8196 | -0.0011804 | 0.012999 | 2.2118 | 230.0233 | 24.9749 | 1 | 99998.4032 | -0.0015968 | 0.0076821 | 2.1615 | [9 empty] | 0.0082348 | 1.072 | 28 | 109.9982 | 5.0015 | 1 | A | 99992.386 | -0.007614 | 0.015906 | 2.1383 | vars | -89.9603 | 0 |  | 1 | -1 | -0.024283 | -1.5267 | 1.1389 | 110.0175 | 5.0004 | 1 | A | 100016.6694 | 0.016669 | 0.013965 | 2.1645 | vars | -0.014 | 0.006386 | 0.40149 | 110 | 5 | 1 | A | 99986.0448 | -0.013955 | 0.013387 | 2.2445 | 24.9961 | 1 |  |  | 0.0021932 | -0.0011804 | -0.0015968 | -0.0098317 | 0.019448 | -0.020026
[27 empty] | 229.99 | 24.9968 | 0.5 Lag | A | 99995.0897 | -0.0049103 | 0.016243 | 2.0967 |  | 0.25672 | 1.096 | 0.040989 | 230.0459 | 24.9609 | 0.5 Lag | A | 99990.9198 | -0.0090802 | 0.01482 | 2.0963 | 230.0193 | 24.9764 | 0.5 Lag | 99994.4239 | -0.0055761 | 0.0098621 | 2.069 | [9 empty] | 0.0025731 | 0.26091 | 29 | 109.9975 | 5.0013 | 0.5 Lag | A | 99995.6672 | -0.0043328 | 0.016733 | 2.0983 | vars | -29.9626 | 0 | Lag | 1 | -1 | -0.021002 | -1.2551 | 1.1252 | 110.0154 | 5.0007 | 0.5 lag | A | 100016.6694 | 0.016669 | 0.014871 | 2.1093 | vars | -0.01 | 0.0056672 | 0.33868 | 110 | 5 | 0.5 lag | A | 99989.8182 | -0.010182 | 0.014271 | 2.1504 | 24.9968 | 0.5 Lag |  |  | -0.0049103 | -0.0090802 | -0.0055761 | -0.0081492 | 0.011112 | -0.025751
[27 empty] | 229.9913 | 49.9952 | 1 | A | 99994.3501 | -0.0056499 | 0.014491 | 2.156 |  | 0.2057 | 0.99074 | 0.47492 | 230.0486 | 50.0018 | 1 | A | 99991.3694 | -0.0086306 | 0.014626 | 2.1382 | 230.0295 | 50.0444 | 1 | 99987.4682 | -0.012532 | 0.0078381 | 2.1427 | [9 empty] | 0.0095502 | 1.2184 | 30 | 240.0202 | 5.0015 | 1 | A | 99990.6838 | -0.0093162 | 0.015914 | 2.1376 | vars | -89.9655 | 0 |  | 1 | -1 | -0.023207 | -1.4582 | 1.1412 | 240.0601 | 5.0004 | 1 | A | 100013.8908 | 0.013891 | 0.013945 | 2.1664 | vars | -0.018 | 0.0086838 | 0.54565 | 240.1 | 5 | 1 | A | 99982.3069 | -0.017693 | 0.013388 | 2.2437 | 49.9952 | 1 |  |  | -0.0056499 | -0.0086306 | -0.012532 | -0.022082 | -0.0055552 | -0.036378
[27 empty] | 229.9635 | 49.9962 | 0.5 Lag | A | 99995.486 | -0.004514 | 0.016089 | 2.1043 |  | 0.62124 | 0.87733 | -0.071279 | 230.037 | 50.0053 | 0.5 Lag | A | 99985.4911 | -0.014509 | 0.018338 | 2.1031 | 230.025 | 50.0402 | 0.5 Lag | 99996.6328 | -0.0033672 | 0.0099691 | 2.0576 | [9 empty] | 0.00043108 | 0.043242 | 31 | 240.0201 | 5.0013 | 0.5 Lag | A | 99993.5858 | -0.0064142 | 0.016736 | 2.0985 | vars | -29.9675 | 0 | Lag | 1 | -1 | -0.017527 | -1.0473 | 1.1254 | 240.0563 | 5.0006 | 0.5 lag | A | 100011.1123 | 0.011112 | 0.014871 | 2.1087 | vars | -0.013 | 0.0065858 | 0.39352 | 240.1 | 5 | 0.5 lag | A | 99987.1626 | -0.012837 | 0.014301 | 2.1489 | 49.9962 | 0.5 Lag |  |  | -0.004514 | -0.014509 | -0.0033672 | -0.0037983 | 0.010001 | -0.026048
#REF! | #REF! | #REF! | #REF! |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 229.8659 | 100.0341 | 1 | A B C | 99993.8986 | -0.0061014 | 0.014509 | 2.1559 |  | 1.3552 | 0.93096 | 0.33711 | -0.43688 | 229.9943 | 99.9814 | 1 | A B C | 99974.2353 | -0.025765 | 0.015585 | 2.1131 |  | -0.94785 | 1.0678 | -1.6684 | 230.0546 | 99.9803 | 1 | A B C | 99989.0075 | -0.010992 | 0.014595 | 2.1424 | 230.0257 | 100.0089 | 1 | 100000.2373 | 0.00023728 | 0.0078069 | 2.1473 | [9 empty] | 0.021938 | 2.81 | 39 | 109.9981 | 24.9996 | 1 | A | 99996.8877 | -0.0031123 | 0.016075 | 2.1272 | vars | -89.8756 | 0 |  | 1 | -1 | -0.017559 | -1.0923 | 1.1427 | 110.0148 | 24.9657 | 1 | A | 100014.4465 | 0.014447 | 0.014067 | 2.1545 | vars | -0.026 | 0.022888 | 1.4238 | 110 | 25 | 1 | A | 99974.1925 | -0.025807 | 0.015361 | 2.1482 | 99.9814 | 1 |  |  | -0.025765 | -0.010992 | 0.00023728 | -0.0217 | 0.021116 | -0.0362
#REF! | #REF! | #REF! | #REF! |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 229.8619 | 100.0331 | 0.5 Lag | A B C | 99983.8847 | -0.016115 | 0.016171 | 2.1007 |  | 0.86554 | 0.79318 | 0.19006 | -0.47159 | 229.9977 | 100.0496 | 0.5 Lag | A B C | 99969.8882 | -0.030112 | 0.020387 | 2.1055 |  | -0.53577 | 1.1192 | -1.0606 | 230.0366 | 99.9732 | 0.5 Lag | A B C | 99980.8112 | -0.019189 | 0.018216 | 2.1084 | 230.0224 | 100.0128 | 0.5 Lag | 99991.5107 | -0.0084893 | 0.0097613 | 2.0678 | [9 empty] | 0.019535 | 2.0013 | 40 | 109.9971 | 24.99 | 0.5 Lag | A | 99997.4335 | -0.0025665 | 0.017066 | 2.0861 | vars | -30.1175 | 0 | Lag | 1 | -1 | -0.018124 | -1.062 | 1.1317 | 110.0144 | 24.968 | 0.5 lag | A | 100015.558 | 0.015558 | 0.01508 | 2.0987 | vars | -0.02 | 0.017433 | 1.0215 | 110 | 25 | 0.5 lag | A | 99979.97 | -0.02003 | 0.014563 | 2.1315 | 100.0496 | 0.5 Lag |  |  | -0.030112 | -0.019189 | -0.0084893 | -0.028025 | 0.011112 | -0.040269
#REF! | #REF! | #REF! | #REF! |  | #REF! | #REF! | #REF! | #REF! | [5 empty] | 229.8623 | 100.0293 | 0.8 Lead | A B C | 99983.7834 | -0.016217 | 0.014792 | 2.1428 | [43 empty] | 41 | 240.0202 | 24.9903 | 1 | A | 99990.1683 | -0.0098317 | 0.015943 | 2.1352 | vars | -89.8785 | 0 |  | 1 | -1 | -0.02928 | -1.8366 | 1.1436 | 240.0706 | 24.9657 | 1 | A | 100019.4482 | 0.019448 | 0.013941 | 2.1661 | vars | -0.02 | 0.010168 | 0.63781 | 240 | 25 | 1 | A | 99979.9742 | -0.020026 | 0.015948 | 2.1464
[65 empty] | 42 | 240.02 | 24.9897 | 0.5 Lag | A | 99991.8508 | -0.0081492 | 0.017169 | 2.0831 | vars | -30.1239 | 0 | Lag | 1 | -1 | -0.019262 | -1.1219 | 1.156 | 240.0678 | 24.9689 | 0.5 lag | A | 100011.1123 | 0.011112 | 0.014852 | 2.1089 | vars | -0.026 | 0.017851 | 1.0397 | 240 | 25 | 0.5 lag | A | 99974.2487 | -0.025751 | 0.014334 | 2.1452
[65 empty] | 55 | 109.9956 | 50.0529 | 1 | A | 99992.3118 | -0.0076882 | 0.015903 | 2.1371 | vars | -89.9842 | 0 |  | 1 | -1 | -0.023246 | -1.4617 | 1.0959 | 110.0143 | 49.9948 | 1 | A | 100015.558 | 0.015558 | 0.014511 | 2.1294 | vars | -0.026 | 0.018312 | 1.1515 | 110 | 50 | 1 | A | 99973.9911 | -0.026009 | 0.013397 | 2.2427
[65 empty] | 56 | 109.9945 | 50.0515 | 0.5 Lag | A | 99994.9222 | -0.0050778 | 0.016948 | 2.0916 | vars | -30.0208 | 0 | Lag | 1 | -1 | -0.019524 | -1.152 | 1.0747 | 110.0109 | 49.9985 | 0.5 lag | A | 100014.4465 | 0.014447 | 0.01577 | 2.0813 | vars | -0.023 | 0.017922 | 1.0575 | 110 | 50 | 0.5 lag | A | 99977.4116 | -0.022588 | 0.014309 | 2.1478
[65 empty] | 53 | 240.0195 | 50.053 | 1 | A | 99977.918 | -0.022082 | 0.015884 | 2.1372 | vars | -89.9795 | 0 |  | 1 | -1 | -0.016527 | -1.0405 | 1.1389 | 240.0627 | 49.9944 | 1 | A | 99994.4448 | -0.0055552 | 0.013947 | 2.1662 | vars | -0.036 | 0.013918 | 0.87621 | 240 | 50 | 1 | A | 99963.6221 | -0.036378 | 0.013475 | 2.2286
[65 empty] | 54 | 240.0165 | 50.0524 | 0.5 Lag | A | 99996.2017 | -0.0037983 | 0.016856 | 2.0944 | vars | -30.0235 | 0 | Lag | 1 | -1 | -0.013799 | -0.81866 | 1.0931 | 240.0619 | 49.9986 | 0.5 lag | A | 100010.001 | 0.010001 | 0.01542 | 2.0875 | vars | -0.026 | 0.022202 | 1.3171 | 240 | 50 | 0.5 lag | A | 99973.9522 | -0.026048 | 0.01424 | 2.1517
[65 empty] | 63 | 109.9973 | 100.0261 | 1 | A B C | 99980.978 | -0.019022 | 0.015889 | 2.1382 | vars | -89.9207 | 0 |  | 1 | -1 | -0.035136 | -2.2114 | 1.1301 | 110.0119 | 99.9634 | 1 | A B C | 100016.1137 | 0.016114 | 0.01406 | 2.1543 | vars | -0.035 | 0.015978 | 1.0056 | 110 | 100 | 1 | A B C | 99965.3407 | -0.034659 | 0.013694 | 2.1989
[65 empty] | 64 | 109.9963 | 100.0244 | 0.5 Lag | A B C | 99978.912 | -0.021088 | 0.016789 | 2.0964 | vars | -29.9235 | 0 | Lag | 1 | -1 | -0.032756 | -1.951 | 1.0943 | 110.0091 | 99.9751 | 0.5 lag | A B C | 100011.668 | 0.011668 | 0.015342 | 2.0891 | vars | -0.033 | 0.011912 | 0.7095 | 110 | 100 | 0.5 lag | A B C | 99967.2821 | -0.032718 | 0.014453 | 2.1362
[65 empty] | 65 | 240.0189 | 100.0221 | 1 | A B C | 99978.2996 | -0.0217 | 0.015931 | 2.1366 | vars | -89.9259 | 0 |  | 1 | -1 | -0.042816 | -2.6876 | 1.1253 | 240.0762 | 99.9743 | 1 | A B C | 100021.1156 | 0.021116 | 0.014158 | 2.149 | vars | -0.036 | 0.0143 | 0.89759 | 240 | 100 | 1 | A B C | 99963.7999 | -0.0362 | 0.013457 | 2.2337
[65 empty] | 66 | 240.02 | 100.0218 | 0.5 Lag | A B C | 99971.9754 | -0.028025 | 0.016722 | 2.0981 | vars | -29.928 | 0 | Lag | 1 | -1 | -0.039137 | -2.3405 | 1.1006 | 240.061 | 99.9765 | 0.5 lag | A B C | 100011.1123 | 0.011112 | 0.015193 | 2.0925 | vars | -0.04 | 0.011975 | 0.71616 | 240 | 100 | 0.5 lag | A B C | 99959.7309 | -0.040269 | 0.01426 | 2.1477
[49 empty] | 230/240v | unity pf
[48 empty] | 2015-02-02 | 2013-01-14 | 2011-02-07 | 2009-02-25 | 2007-02-23 | 2005-02-23 | 2003-02-21 | 2001-02-23
[47 empty] | 1 | 0.0025232 | -0.00087428 | -0.00091797 | -0.0020296 | -0.0012436 | -0.0010001 | 0.0027779 | -0.007
[47 empty] | 5 | 0.0029853 | -0.0012241 | -0.00034762 | -0.0012861 | -0.0025649 | -0.0016932 | 0.0027779 | -0.003
[47 empty] | 25 |  |  | 0.00022076 | -0.001091 | -0.0011661 | -0.0029945 | 0.0033334 | -0.002
[47 empty] | 50 | 0.0039288 | -0.0037198 | -0.0026011 | -0.001086 | -0.00654 | -0.022335 | -0.012221 | -0.011
[47 empty] | 100 | 0.0050558 | 0.0010496 | 0.0028779 | -0.0020435 | 4.0752e-05 | -0.0015584 | 0.0016667 | -0.006
[49 empty] | 230/240v | unity var f
[48 empty] | 2015-02-02 | 2013-01-14 | 2011-02-07 | 2009-02-25 | 2007-02-23 | 2005-02-23 | 2003-02-21 | 2001-02-23
[47 empty] | 1 | 0.0024671 | -0.0059006 | 0.0022391 | 0.004468 | 0.004468 | -0.0082323 | 0.016669 | -0.017
[47 empty] | 5 | 0.00035219 | -0.0067447 | 0.0033609 | 0.0014207 | 0.0014207 | -0.0093162 | 0.013891 | -0.018
[47 empty] | 25 |  |  | 0.0021932 | -0.0015968 | -0.0015968 | -0.0098317 | 0.019448 | -0.02
[47 empty] | 50 |  |  | -0.0049103 | -0.012532 | -0.012532 | -0.022082 | -0.0055552 | -0.036
[47 empty] | 100 |  |  | -0.004514 | 0.00023728 | 0.00023728 | -0.0217 | 0.021116 | -0.036

## TABLE RD31 (hidden)
|  | Watts | Radian RD31
| 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Key | row select from "Data" | Voltage U (V) | Current I (A) | Phase Difference φ (°) | #N/A | #N/A | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k | watts/vars | Channel | IEC phase angle | Comments | Ch1 Type A/ Type B | Ch3 Type A/ Type B | PF label | Power factor rounding | Uncert/CMC
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
11105.0090watts | [6 empty] | 0 | 0.16771 | 2.1
|  | Voltage U (V) | Current I (A) | Phase Difference φ (°) | #N/A | #N/A | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
#N/A |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A
31100.5090watts | [6 empty] | 0 | 0.16771 | 2.1
Note on O5: t.stewart: Manual checks may be required to ensure the assumption that Type B uncertainty dominates Type A. (So where Type B is assumed to be positively correlated between CH1 and CH3, adding their uncertainties will be correct.)
Note on Q6: l.christian: First part picks out the 0, 90,180, 270 points and the second part the quadrant- 0 to 90 is EVEN, 90 to 180 is ODD, 180 to 270 is EVEN and 270 to 0 is ODD. Sign of PF handled in col F

## D3000 Star Table (hidden)
D3000 Star Power Results Summary
xx Amp Watts
row select from "Star Data" | Energised Phases | Voltage U (V) | Current I (A) | Phase Difference φ (°) | #N/A | #N/A | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k | watts/vars | PF Phase (IEC) | Power factor rounding | PF Label | Uncert/CMC
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
xx Amp Vars
row select from "Star Data" | Energised Phases | Voltage U (V) | Current I (A) | Phase Difference φ (°) | #N/A | #N/A | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k | watts/vars | PF Phase | Power factor rounding | PF Label | Uncert/CMC
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
Note on O5: l.christian: First part picks out the 0, 90,180, 270 points and the second part the quadrant- 0 to 90 is EVEN, 90 to 180 is ODD, 180 to 270 is EVEN and 270 to 0 is ODD. Sign of PF handled in col F

## D3000 Delta Table (hidden)
D3000 Delta Power Results Summary
xx Amp Watts
row select from "2W Delta Data" | Energised Phases | Voltage U12, U32 (V) | Current I1, I3 (A) | #N/A | Phase Angle U12 - I1 (°) | Phase Angle U32 - I3 (°) | #N/A | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k | watts/vars | Nominal Phase | IEC phase angle Ch1 | IEC phase angle Ch3 | PF Phase | Power factor rounding | PF Label | Uncert/CMC
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 2 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 2 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A | #N/A
xx Amp Vars
row select from "2W Delta Data" | Energised Phases | Voltage U12, U32 (V) | Current I1, I3 (A) | #N/A | Phase Angle U12 - I1 (°) | Phase Angle U32 - I3 (°) | #N/A | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k | watts/vars | Nominal Phase | IEC phase angle Ch1 | IEC phase angle Ch3 | PF Phase | Power factor rounding | PF Label | Uncert/CMC
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 2 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 2 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A | #N/A
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 3 | #N/A | #N/A
Note on S6: l.christian: First part picks out the 0, 90,180, 270 points and the second part the quadrant- 0 to 90 is EVEN, 90 to 180 is ODD, 180 to 270 is EVEN and 270 to 0 is ODD. Sign of PF handled in col F

## TABLES TVH4.322 (hidden)
REPORT TABLES
Basic Load curve
TVH4
| Energised Phase(s) | Current Range (A) | Voltage Range (V) | Voltage U (V) | Current I (A) | Power Factor | Measured Impulses (Imp/W.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k |  |  |  | Power Bridge phase angle (Non-IEC) | Rounding for pf
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 |  |  |  | Previous reports
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 |  |  |  | 5(10) 160-320 |  | 5A 230V
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 |  |  |  | All 3 phases
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 |  |  |  | pf |  | S21452 | S21361 | S21262
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 |  |  |  |  | 2004-04-30 | 2002-04-29 | 2000-05-05 | 1998-04-15
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 |  |  |  | 1 | #N/A | 0.047 | 0.04 | 0.05
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 |  |  |  | 0.5 lag | #N/A | 0.013 | 0.04 | 0.02
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 |  |  |  | 0.5 lead | #N/A | 0.097 | 0.07 | 0.08
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 |  |  |  | 5(10) 160-320 |  | 0.25A 230V
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 |  |  |  | All 3 phases
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 |  |  |  |  | 2004-04-30 | 2002-04-29 | 2000-05-05 | 1998-04-15
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 |  |  |  | 1 | #N/A | 0.078 | 0.08 | 0.04
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 |  |  |  | 0.5 lag | #N/A | 0.143 | 0.1 | 0.54
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
Basic Load curve continued
TVH4
| Energised Phase(s) | Current Range (A) | Voltage Range (V) | Voltage U (V) | Current I (A) | Power Factor | Measured Impulses (Imp/W.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k | [16 empty] | Top of range
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 |  | 0pf | 63.5V | Top of range |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | [7 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | [7 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 |  | 0pf | 63.5V | Mid range |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 0 | 0 | 0 | 0 | 0 | 0 | [5 empty] | Mid range
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 |  | 0pf | 63.5V | Bottem of range |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 0 | 0 | 0 | 0 | 0 | 0 | [5 empty] | Bottem of range
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
[16 empty] | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
[16 empty] | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0
Effect of Voltage
TVH4
| Energised Phase(s) | Current Range (A) | Voltage Range (V) | Voltage U (V) | Current I (A) | Power Factor | Measured Impulses (Imp/W.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
Current Ranges
TVH4
| Energised Phase(s) | Current Range (A) | Voltage Range (V) | Voltage U (V) | Current I (A) | Power Factor | Measured Impulses (Imp/W.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
VA Reactive
TVH4
| Energised Phase(s) | Current Range (A) | Voltage Range (V) | Voltage U (V) | Current I (A) | Var Factor | Measured Impulses (Imp/Var.h) | Error E (%) | Expanded Uncertainty (%) | Coverage Factor k
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
| #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1
[23 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0
Note on R7: Laurie Christian: Looks at WattBridge phase so independent of source phase definition

## Star Data (hidden)
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80 | 81 | 82 | 83 | 84 | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 117 | 118 | 119 | 120 | 121 | 122 | 123 | 124 | 125 | 126 | 127 | 128 | 129 | 130 | 131 | 132 | 133 | 134 | 135 | 136 | 137 | 138 | 139 | 140 | 141 | 142 | 143 | 144 | 145 | 146 | 147 | 148 | 149 | 150 | 151 | 152 | 153 | 154 | xx | xx | 157 | 158 | 159 | xx | xx | 162 | 163 | 164 | 165 | 166 | 167 | 168 | 169 | 170 | 171 | 172 | 173 | 174 | 175 | 176 | 177 | 178 | 179 | 180 | 181 | 182 | 183 | 184
Row no. | manually enter "watts" or "vars" | RD31 Ch1 or Ch123 |  | Nom. U1 | Nom. U2 | Nom. U3 | Nom I1 | Nom. I2 | Nom. I3 | Nom. Phase (IEC) | Nom. f | V Range | I Range | Counts per Wh | Type | Ref. U1 | Ref. U2 | Ref. U3 | Ref. I1 | Ref. I2 | Ref. I3 | Ref. phi1 (corrected to IEC for RD31) | Ref. phi2 (corrected to IEC for RD31) | Ref. phi3 (corrected to IEC for RD31) | Ref. f | Ref. Watts1 | Ref. Watts2 | Ref. Watts3 | Ref. Total Watts | Ref. Vars1 | Ref. Vars2 | Ref. Vars3 | Ref. Total Vars | Ref. VA1 | Ref. VA2 | Ref. VA3 | Ref. Timestamp | Mtr. U1 | Mtr. U2 | Mtr. U3 | Mtr. I1 | Mtr. I2 | Mtr. I3 | Mtr. phi1 | Mtr. phi2 | Mtr. phi3 | Mtr. f | Mtr. Watts1 | Mtr. Watts2 | Mtr. Watts3 | Mtr. Total Watts | Mtr. Vars1 | Mtr. Vars2 | Mtr. Vars3 | Mtr. Total Vars | Mtr. VA1 | Mtr. VA2 | Mtr. VA3 | Mtr. Timestamp | Error U1 | Error U2 | Error U3 | Error I1 | Error I2 | Error I3 | Error phi1 | Error phi2 | Error phi3 | Error f | Error Watts1 | Error Watts2 | Error Watts3 | Error Total Watts | Error Vars1 | Error Vars2 | Error Vars3 | Error Total Vars | Error VA1 | Error VA2 | Error VA3 | V.1 | V.2 | V.3 | V.Net | A.1 | A.2 | A.3 | A.Net | W.1 | W.2 | W.3 | W.Net | VA.1 | VA.2 | VA.3 | VA.Net | VAR.1 | VAR.2 | VAR.3 | VAR.Net | Hz.1 | Hz.2 | Hz.3 | Hz.Net | oP.1 | oP.2 | oP.3 | oP.Net | PF.1 | PF.2 | PF.3 | PF.Net | AnlgSens.1 | AnlgSens.2 | AnlgSens.3 | AnlgSens.Net | doP.1 | doP.2 | doP.3 | doP.Net | V Delta.1 | V Delta.2 | V Delta.3 | V Delta.Net | W Delta.1 | W Delta.2 | W Delta.3 | W Delta.Net | VA Delta.1 | VA Delta.2 | VA Delta.3 | VA Delta.Net | VAR Delta.1 | VAR Delta.2 | VAR Delta.3 | VAR Delta.Net | VAR X WYE.1 | VAR X WYE.2 | VAR X WYE.3 | VAR X WYE.Net | VAR X Delta.1 | VAR X Delta.2 | VAR X Delta.3 | VAR X Delta.Net | RD31 Power | Point key | RD31 Err | Crctd Total Ref Power | Repeated point count | Last reading of a set? | RD31 exp unc | RD31 cover Factor | RD31 std unc | Method std unc (enter to cell above) | D3000 type B (enter to cell above) | stdev readings | Total pwr std unc | cmb std unc | method dfm (enter to cell above) | D3000 dfm (enter to cell above) | Type A dfm | Ch1 dfm | cmb dfm | cmb cover | exp unc | Removed readings | Meter % err | average | Imp/(Wh orVarh) | Average V
4 | watts | #DIV/0! | [142 empty] | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Note on E4: t.stewart: Paste csv file data (from controller.py output) to cell E5 onwards. Copy formula down for columns: A:D and EP:FO Manually enter for watts or vars in column B.

## 2W Delta Data (hidden)
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80 | 81 | 82 | 83 | 84 | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 117 | 118 | 119 | 120 | 121 | 122 | 123 | 124 | 125 | 126 | 127 | 128 | 129 | 130 | 131 | 132 | 133 | 134 | 135 | 136 | 137 | 138 | 139 | 140 | 141 | 142 | 143 | 144 | 145 | 146 | 147 | 148 | 149 | 150 | 151 | 152 | 153 | 154 | 155 | 156 | 157 | 158 | 159 | 160 | 161 | 162 | 163 | 164 | 165 | 166 | 167 | 168 | 169 | 170 | 171 | 172 | 173 | 174 | 175 | 176 | 177 | 178 | 179 | 180 | 181 | 182 | 183 | 184
[162 empty] | xxx | xxx | [5 empty] | xxx | xxx
Row no. | manually enter "watts" or "vars" | Ch1 Energised | Ch3 Energised | Nom. U1 | Nom. U2 | Nom. U3 | Nom I1 | Nom. I2 | Nom. I3 | Nom. Phase (IEC) | Nom. f | V Range | I Range | Counts per Wh | Type | Ref. U1 | Ref. U2 | Ref. U3 | Ref. I1 | Ref. I2 | Ref. I3 | Ref. phi1 | Ref. phi2 | Ref. phi3 | Ref. f | Ref. Watts1 | Ref. Watts2 | Ref. Watts3 | Ref. Total Watts | Ref. Vars1 | Ref. Vars2 | Ref. Vars3 | Ref. Total Vars | Ref. VA1 | Ref. VA2 | Ref. VA3 | Ref. Timestamp | Mtr. U1 | Mtr. U2 | Mtr. U3 | Mtr. I1 | Mtr. I2 | Mtr. I3 | Mtr. phi1 | Mtr. phi2 | Mtr. phi3 | Mtr. f | Mtr. Watts1 | Mtr. Watts2 | Mtr. Watts3 | Mtr. Total Watts | Mtr. Vars1 | Mtr. Vars2 | Mtr. Vars3 | Mtr. Total Vars | Mtr. VA1 | Mtr. VA2 | Mtr. VA3 | Mtr. Timestamp | Error U1 | Error U2 | Error U3 | Error I1 | Error I2 | Error I3 | Error phi1 | Error phi2 | Error phi3 | Error f | Error Watts1 | Error Watts2 | Error Watts3 | Error Total Watts | Error Vars1 | Error Vars2 | Error Vars3 | Error Total Vars | Error VA1 | Error VA2 | Error VA3 | V.1 | V.2 | V.3 | V.Net | A.1 | A.2 | A.3 | A.Net | W.1 | W.2 | W.3 | W.Net | VA.1 | VA.2 | VA.3 | VA.Net | VAR.1 | VAR.2 | VAR.3 | VAR.Net | Hz.1 | Hz.2 | Hz.3 | Hz.Net | oP.1 | oP.2 | oP.3 | oP.Net | PF.1 | PF.2 | PF.3 | PF.Net | AnlgSens.1 | AnlgSens.2 | AnlgSens.3 | AnlgSens.Net | doP.1 | doP.2 | doP.3 | doP.Net | V Delta.1 | V Delta.2 | V Delta.3 | V Delta.Net | W Delta.1 | W Delta.2 | W Delta.3 | W Delta.Net | VA Delta.1 | VA Delta.2 | VA Delta.3 | VA Delta.Net | VAR Delta.1 | VAR Delta.2 | VAR Delta.3 | VAR Delta.Net | VAR X WYE.1 | VAR X WYE.2 | VAR X WYE.3 | VAR X WYE.Net | VAR X Delta.1 | VAR X Delta.2 | VAR X Delta.3 | VAR X Delta.Net | RD31 Ch1 | RD31 Ch3 | Ch1 key | Ch3 key | Ch1 err | Ch3 err | Crctd Ch1 | Crctd Ch3 | Crctd Total Ref Power | Repeated point count | Last reading of a set? | Ch1 exp unc | Ch3 exp unc | Ch1 cover Factor | Ch3 cover Factor | Ch1 std unc | Ch3 std unc | Method std unc (enter to cell above) | D3000 type B (enter to cell above) | stdev readings | Total pwr std unc | cmb std unc | Ch1 dfm | Ch3 dfm | method dfm (enter to cell above) | D3000 dfm (enter to cell above) | Type A dfm | cmb dfm | cmb cover | exp unc | Removed readings | Meter % err | average | Imp/(Wh orVarh) | Average V
4 | watts | FALSE | FALSE | [141 empty] | 0 | 0 | Channel Off | Channel Off | #N/A | #N/A | #N/A | #N/A | #N/A | 0 | TRUE | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | xxx | xxx | #N/A | #N/A | #NAME? | #N/A | #N/A | xxx | xxx | -1 | #NAME? | #NAME? | #NAME? |  | #N/A | #N/A | #N/A | #DIV/0!
Note on E4: t.stewart: Paste csv file data (from controller.py output) to cell E5 onwards. Copy formula down for columns: A:D and EP:FX Manually enter for watts or vars in column B.

## TVH4 Unc (hidden)
Uncertainty budget for TVH.4
Active energy |  |  |  | manufacturer's spec | [5 empty] | 1 sigma uncertainty
|  |  |  | 2 sigma (%) | [5 empty] | 1 pf | 0.8 pf | 0.5 pf | accuracy
|  |  |  | 1 pf | 0.8 pf | 0.5 pf |  |  |  | (%) | (%) | (%) | of unc % | deg frm
influence of temperature, per deg C |  |  |  | 0.01 |  | 0.02 |  | temp range, deg C | 1 | 0.005 | 0.008 | 0.01 | 30 | 5.5556
magnetic field, 500 uT |  |  |  | 0.5 |  | 0.5 |  | mag field, uT | 5 | 0.0025 | 0.0025 | 0.0025 | 50 | 2
voltage dependency |  |  |  | 0.3 |  | 0.4 |  | voltage departure from nom | 0.001 | 0.00015 | 0.00018 | 0.0002 | 30 | 5.5556
frequency, +/- 5% variation |  |  |  | 0.1 |  | 0.1 |  | freq departure from nom | 1.4e-05 | 0.00014 | 0.00014 | 0.00014 | 20 | 12.5
| power factor, cos phi |  |  |  | 1 | 0.8 | 0.5
| overall uncertainty, % |  |  |  | #NAME? | #NAME? | #NAME?
| degrees of freedom |  |  |  | #NAME? | #NAME? | #NAME?
Reactive energy |  |  |  | manufacturer's spec | [5 empty] | 1 sigma uncertainty
|  |  |  | 2 sigma (%) | [5 empty] | 1 var fctr | 0.8 var fctr | 0.5 var fctr | accuracy
|  |  |  | 1 var fctr | 0.8 var fctr | 0.5 var fctr |  |  |  | (%) | (%) | (%) | of unc % | deg frm
influence of temperature, per deg C |  |  |  | 0.03 |  | 0.05 |  | temp range, deg C | 1 | 0.015 | 0.021 | 0.025 | 30 | 5.5556
magnetic field, 500 uT |  |  |  | 1 |  | 1 |  | mag field, uT | 5 | 0.005 | 0.005 | 0.005 | 50 | 2
voltage dependency |  |  |  | 0.3 |  | 0.4 |  | voltage departure from nom | 0.001 | 0.00015 | 0.00018 | 0.0002 | 30 | 5.5556
frequency, +/- 5% variation |  |  |  | 0.5 |  | 0.5 |  | freq departure from nom | 1.4e-05 | 0.0007 | 0.0007 | 0.0007 | 20 | 12.5
| var factor, sin phi |  |  |  | 1 | 0.8 | 0.5
| overall uncertainty, % |  |  |  | #NAME? | #NAME? | #NAME?
| degrees of freedom |  |  |  | #NAME? | #NAME? | #NAME?
Lookup Table of Uncertainties
|  | 0 | 0watts | #NAME? | #NAME?
|  | 36.87 | 36.87watts | #NAME? | #NAME?
|  | 60 | 60watts | #NAME? | #NAME?
|  | 90 | 90vars | #NAME? | #NAME?
|  | 53.13 | 53.13vars | #NAME? | #NAME?
|  | 30 | 30vars | #NAME? | #NAME?

## RM11 Unc
Type B uncertainties for RM11, to be used when CALIBRATING the RM11.
The only type B influence specified in the manual is the effect of temperature.
This is "0.001 % per degree C maximum".
For a lab temperature of 20 +/- 1 deg C, take the 2 sigma uncertainty as 0.0005%, with 6 degrees of freedom.
Our RM11 suffers from a quantization problem - this may not be the case for all RM11s
12 Feb 2007: Laurie's additions: RM10 specs (errors (% of readings)- note these are expected errors not uncertainties in those errors)
Watthour
Power Factor | Typical | Maximum
1 | 0.01 | 0.05
0.5 | 0.02 | 0.05
P<0.5 |  | 0.05/P | Assume uncertainty (one sigma) of 0.0005 + 0.01/2/PF as a % of reading for watts
VAR hour
Reactive power factor
1 | 0.025 | 0.1
0.5 | 0.035 | 0.1 | Assume uncertainty (one sigma) of 0.0005 + 0.02/2/QF as a % of reading for vars
23 March 2007: Laurie's or Bruce's additions: RM11 specs (errors (% of readings)- note these are expected errors not uncertainties in those errors) | [11 empty] | 1 | 0.0015 |  | angle |  | RSS with 0.000645
Watthour | [11 empty] | 0.5 | 0.0035 |  | 0 | 0.0015 | 0.0016328 | (0.0005+0.001/COS(RADIANS(ABS(CellRef))))
Power Factor | Typical | Maximum |  |  |  |  | From S21672 Watt variance at 50A unity pf = | [7 empty] | 30 | 0.0016547 | 0.001776 | (0.0005+0.001/COS(RADIANS(ABS(CellRef))))
1 | 0.003 | 0.025 |  | 0.004 |  |  | 0.0013672 | See "This RM11" worksheet (I88)
0.5 | 0.007 | 0.025 |  |  |  |  | The spec shows a change of 0.002% from unity to 0.5pf, | [7 empty] | 90 | 0.002 | 0.0021014 | (-0.0005+(0.0025/SIN(RADIANS(ABS(N35)))))
P<0.5 |  | 0.025/P |  |  |  |  | at 1sigma this would be 0.001% | [7 empty] | 60 | 0.0023868 | 0.0024724 | (-0.0005+(0.0025/SIN(RADIANS(ABS(N35)))))
[7 empty] | 0.0023672 | At 0.5pf
RSS with 0.000645
[15 empty] | 0 | 0.0015 | 0.0016328 | (-0.008*COS(RADIANS(ABS(CellRef))) + 0.011)/2
[15 empty] | 60 | 0.0035 | 0.0035589 | (-0.008*COS(RADIANS(ABS(CellRef))) + 0.011)/2
VAR hour
Reactive power factor | [6 empty] | From S21672 Var variance at 50A unity Qf = | [7 empty] | 90 | 0.00375 | 0.0038051 | (-0.01*SIN(RADIANS(ABS(CellRef)))+ 0.025)/4
1 | 0.015 | 0.05 |  | 0.005 |  |  | 0.0021479 | See "This RM11" worksheet (J112) |  |  |  | 1 | 0.00375 |  | 30 | 0.005 | 0.0050414 | (-0.01*SIN(RADIANS(ABS(CellRef)))+ 0.025)/4
0.5 | 0.02 | 0.05 |  |  |  |  | The spec shows a change of 0.005% from unity to 0.5pf, |  |  |  |  | 0.5 | 0.005
at 1sigma this would be 0.0025%
[7 empty] | 0.0046479 | At 0.5Qf
Frequency Spec on Vars
Practical formula for row 12
50 Hz | 2% / hz |  | from this spreadsheet our frequency variation using the HEG is:
IF(M12="watts",(0.0005+0.001/COS(RADIANS(ABS(E12)))),(0.001+0.001/SIN(RADIANS(ABS(E12)))))
|  |  | Center Freq = | 50.0002 | hz | [5 empty] | or
|  |  | 1 sd = | 0.00020374 | hz | [6 empty] | IF(M12="watts",SQRT((0.0005+0.001/COS(RADIANS(ABS(E12))))^2+0.000645^2),SQRT((0.001+0.001/SIN(RADIANS(ABS(E12))))^2+0.000645^2))
|  |  | 1 sd uncert = | 0.00040749 | %
| RSS with the above temperature contribution | [10 empty] | IF(M12="watts",(0.0005+0.001/COS(RADIANS(ABS(E12)))),IF(M12="vars",(0.001+0.001/SIN(RADIANS(ABS(E12)))),"What are you doing?"))
|  |  | Type B Temperature + frequency = | 0.00064502 | % (for vars) | [5 empty] | or
[11 empty] | Bruce | IF(M12="watts",SQRT((0.0005+0.001/COS(RADIANS(ABS(E12))))^2+0.0005^2),IF(M12="vars",SQRT((0.001+0.001/SIN(RADIANS(ABS(E12))))^2+0.000645^2),"What are you doing?"))
[11 empty] | Laurie | IF(M12="watts",SQRT(((-0.008*COS(RADIANS(ABS(E12)))+0.011)/2)^2+0.0005^2),IF(M12="vars",SQRT(((-0.01*SIN(RADIANS(ABS(E12)))+ 0.025)/4)^2+0.000645^2),"What are you doing?"))
Copied from S21739_livened_newIVDconsts_corrctConst Laurie.xls
Copied from S21673_livened.xls
From report S21583a (a previous report for this Avo MA10-08)
Formula from:
From an eyeball examination of the std devs, take the overall uncertainty to be
|  | RM11Unc | 0.008 | % of applied VA |  |  |  |  | S21641 | E36040065
IF(K60="watts",0.0005,0.0015)
These uncertainties are based on a large number of measurements, but
| except at low currents the uncertainties are conservative. | [7 empty] | S21583a |  |  | Vircom EMS |  | Last calibration of this Avo
| Hence take the degrees of freedom to be | [7 empty] | IF(K64="watts",0.006*Q64*R64/BU64,0.008+0.004*Q64*R64/BV64)
| Above 0.1 A |  | 20
| 0.1 A and below |  | 5
[9 empty] | S21659 | E36040081 |  | Metrix
| MA-10-08 specifications | [7 empty] | No change from traditional .0005%, en's ok except for 50A. (the trigger level in the previous calibration was the problem)
Watthour
| Unity |  |  | 0.01% typical |  | 0.05% max |  |  | S21672 | E36040083 |  | Northpower
| 0.5 lag |  |  | 0.02% typical |  |  |  |  | IF(K71="watts",SQRT(((-0.008*COS(RADIANS(ABS(C71)))+0.011)/2)^2+0.0005^2),IF(K71="vars",SQRT(((-0.01*SIN(RADIANS(ABS(C71)))+ 0.025)/4)^2+0.000645^2),"What are you doing?"))
| Power factor P<0.5 (-60 to -90 degrees) |  |  | - |  | 0.05%/P max
VAR hour
| 0 Lag power factor (-90 degrees) |  |  | 0.025% typical |  | 0.1% max
| 0.866 lag (-30 degrees) |  |  | 0.035% typical |  | 0.1% max
Q hour- not available on the MA-10-08
| 0 Lag (0 degrees) |  |  | 0.035% typical |  | 0.1% max
| 0.5 Lag (-60 degrees) |  |  | 0.025% typical |  | 0.1% max
Vircom MA-10-08 is varying by more than 0.008% from one calibration to next
for the unity var factor points. Increase Type B for them to 0.012%. See
Type B column formula on Data page.
For same unit the stability of the watts results is better than 0.008% so
have decreased this uncertainty to 0.006%
This report: S21739 (2009). Use uncertainty formula in cell O46 as this is gives reasonable changes in Prime Points, is related to specs, and also give reasonable changes relative to 2007.
Not necessary to use S21583 values since these are not warranted by changes observed.

## TVE102_3 (hidden)
Type B uncertainties for TVE 102/3
Manufacturers Specifications | [10 empty] | Variation in Lab
ACTIVE ENERGY |  |  |  |  | REACTIVE ENERGY | [5 empty] | ACTIVE ENERGY | [5 empty] | REACTIVE ENERGY
Voltage dependency within the range 50 to 350 volts, |  |  |  |  | Voltage dependency within the range 50 to 350 volts, | [11 empty] | Voltage dependency within the range 50 to 350 volts,
re base value of 220 volts |  |  |  |  | re base value of 220 volts | [5 empty] | Voltage dependency, for V within 1 % nominal | [5 empty] | re base value of 220 volts
|  | cos phi | typ % | max % |  |  | sin phi | typ % | max % |  |  |  | cos phi | typ % | max % |  |  |  | sin phi | typ % | max %
|  | 1 | 0.1 | 0.15 |  |  | 1 | 0.2 | 0.3 |  |  |  | 1 | 3.3333e-06 | 5e-06 |  |  |  | 1 | 6.6667e-06 | 1e-05
|  | 0.5 | 0.15 | 0.25 |  |  | 0.5 | 0.3 | 0.5 |  |  |  | 0.5 | 5e-06 | 8.3333e-06 |  |  |  | 0.5 | 1e-05 | 1.6667e-05
frequency dependency within +/- 5 % of 50 Hz |  |  |  |  | frequency dependency within +/- 5 % of 50 Hz | [5 empty] | frequency dependency, for f within .0005% 50 Hz | [5 empty] | frequency dependency within +/- 5 % of 50 Hz
|  | cos phi | typ % | max % |  |  | sin phi | typ % | max % |  |  |  | cos phi | typ % | max % |  |  |  | sin phi | typ % | max %
|  | 1 | 0.02 | 0.03 |  |  | 1 | 0.04 | 0.06 |  |  |  | 1 | 5e-06 | 7.5e-06 |  |  |  | 1 | 1e-05 | 1.5e-05
|  | 0.5 | 0.03 | 0.05 |  |  | 0.5 | 0.06 | 0.1 |  |  |  | 0.5 | 7.5e-06 | 1.25e-05 |  |  |  | 0.5 | 1.5e-05 | 2.5e-05
temperature dependency per deg C |  |  |  |  | temperature dependency per deg C | [5 empty] | temperature dependency per deg C | [5 empty] | temperature dependency per deg C
| In % | cos phi | typ % | max % |  | In % | sin phi | typ % | max % |  |  | In % | cos phi | typ % | max % |  |  | In % | sin phi | typ % | max %
| 50 to 120 | 1 | 0.003 | 0.005 |  | 50 to 120 | 1 | 0.006 | 0.01 |  |  | 50 to 120 | 1 | 0.003 | 0.005 |  |  | 50 to 120 | 1 | 0.006 | 0.01
| 10 to 50 | 1 | 0.005 | 0.008 |  | 10 to 50 | 1 | 0.01 | 0.015 |  |  | 10 to 50 | 1 | 0.005 | 0.008 |  |  | 10 to 50 | 1 | 0.01 | 0.015
| 50 to 120 | 0.5 | 0.005 | 0.008 |  | 50 to 120 | 0.5 | 0.01 | 0.015 |  |  | 50 to 120 | 0.5 | 0.005 | 0.008 |  |  | 50 to 120 | 0.5 | 0.01 | 0.015
| 10 to 50 | 0.5 | 0.007 | 0.01 |  | 10 to 50 | 0.5 | 0.015 | 0.02 |  |  | 10 to 50 | 0.5 | 0.007 | 0.01 |  |  | 10 to 50 | 0.5 | 0.015 | 0.02
external fields, 500 uT |  |  |  |  | external fields, 500 uT | [5 empty] | external fields, 500 uT | [5 empty] | external fields, 500 uT
| unbal load | 1 | 0.1 | 0.2 |  | unbal load | 1 | 0.2 | 0.4 |  |  | unbal load | 1 | 0.001 | 0.002 |  |  | unbal load | 1 | 0.002 | 0.004
ASSUMPTIONS | [8 empty] | Overall Type B Uncertainty
[9 empty] | Active Energy | [7 empty] | combined |  | Unc vs cos phi
Take manufacturers "max" spec to be 2 sigma | [9 empty] | Amps |  | cos phi | voltage | freqy | temp | field | unc 1 sigma |  | x cos phi | const
Take lab temp to be within +/- 1 deg C | [8 empty] | .5 to 1.2 | 5 to 12 | 50 to 120 | 1 | 2.5e-06 | 3.75e-06 | 0.0025 | 0.001 | #NAME? |  | #NAME? | #NAME?
Take frequency of CH source to be within 0.005 % | [8 empty] | .1 to .5 | 1 to 5 | 10 to 50 | 1 | 2.5e-06 | 3.75e-06 | 0.004 | 0.001 | #NAME? |  | #NAME? | #NAME?
Take lab mag field to be less than 5 uT | [8 empty] | .5 to 1.2 | 5 to 12 | 50 to 120 | 0.5 | 4.1667e-06 | 6.25e-06 | 0.004 | 0.001 | #NAME?
Take test voltage to be within 1 % of nominal | [8 empty] | .1 to .5 | 1 to 5 | 10 to 50 | 0.5 | 4.1667e-06 | 6.25e-06 | 0.005 | 0.001 | #NAME?
Take number of degrees of freedom in final uncertainty to be 8
since there is a fairly high confidence in the numbers | [8 empty] | Active Energy | [7 empty] | combined |  | Unc vs sin phi
[12 empty] | cos phi | voltage | freqy | temp | field | unc 1 sigma |  | x sin phi | const
[9 empty] | .5 to 1.2 | 5 to 12 | 50 to 120 | 1 | 5e-06 | 7.5e-06 | 0.005 | 0.002 | #NAME? |  | #NAME? | #NAME?
[9 empty] | .1 to .5 | 1 to 5 | 10 to 50 | 1 | 5e-06 | 7.5e-06 | 0.0075 | 0.002 | #NAME? |  | #NAME? | #NAME?
[9 empty] | .5 to 1.2 | 5 to 12 | 50 to 120 | 0.5 | 8.3333e-06 | 1.25e-05 | 0.0075 | 0.002 | #NAME?
[9 empty] | .1 to .5 | 1 to 5 | 10 to 50 | 0.5 | 8.3333e-06 | 1.25e-05 | 0.01 | 0.002 | #NAME?

## D3000Unc (hidden)
D3000 Uncertainty Specifications
[7 empty] | assume | hence
Voltage dependency quoted as: | [6 empty] | 2 sigma | 1 sigma
48V to 160 V and 160 V to 320 V, +/- 0.05% |  |  |  | Hence % per volt, low range: |  |  | 0.00044643 | 0.00022321
|  |  |  | Hence % per volt, high range: |  |  | 0.0003125 | 0.00015625
Frequency dependency quoted as:
47 to 63 Hz, +/- 0.05% |  |  |  | Hence % per Hz: |  |  | 0.003125 | 0.0015625
Temperature dependency quoted as:
0 to 40 deg C, +/- 0.005% |  |  |  | Hence % per degree: |  |  | 0.005 | 0.0025
Low Voltage Range |  |  |  | High Voltage Range
|  |  | Uncerty, % |  |  |  | Uncerty, %
Parameter | Variation | Units |  | Parameter | Variation | Units |  |  | There is insignificant difference between the high and low voltage ranges.
Voltage | 0.5 | % | 0.00011161 | Voltage | 0.5 | % | 7.8125e-05 |  | The assumption that the specs quoted are 2 sigma is probably conservative,
Frequency | 0.002 | % | 3.125e-06 | Frequency | 0.002 | % | 3.125e-06 |  | so assume there is 15% confidence in these values. Hence dfm =
Temperature | 1 | Hz | 0.0025 | Temperature | 1 | Hz | 0.0025
|  | 1 sigma unc (%) | 0.0025025 |  |  | rss: | 0.0025012
|  | dfm: | 22

## TVH2.32 (hidden)
TVH2.32 Specifications
[7 empty] | assume | hence
Voltage dependency, 60 to 140% Un, quoted as: | [6 empty] | 2 sigma | 1 sigma
cos phi = 1: 0.015% per 10% delta U |  |  |  | Hence % per volt, unity pf |  |  | 0.0015 | 0.00075
cos phi = 0.5, 0.02% per 10% delta U |  |  |  | Hence % per volt, 0.5 pf |  |  | 0.002 | 0.001
Frequency dependency quoted as:
0.01% for 5% delta f |  |  |  | Hence % per Hz: |  |  | 0.004 | 0.002
Temperature dependency quoted as:
0.003% per degree C |  |  |  | Hence % per degree: |  |  | 0.003 | 0.0015
Active energy |  |  |  | Reactive energy
|  |  | Uncerty, % |  |  |  | Uncerty, %
Parameter | Variation | Units |  | Parameter | Variation | Units
Voltage | 0.5 | % | 0.000375 | Voltage | 0.5 | % | 0.0005
Frequency | 0.02 | Hz | 4e-05 | Frequency | 0.02 | Hz | 4e-05
Temperature | 1 | degree | 0.0015 | Temperature | 1 | degree | 0.0015
|  | 1 sigma unc (%) | 0.0015467 |  |  | rss: | 0.0015816
|  | dfm: | 22
Seeing there is not much difference between active and reactive energy,
take worst case for simplicity.
|  | TVH2_32Unc | 0.0015816
|  | TVH2_32DFM | 22

## Unit Tests (hidden)
Unit Tests | [5 empty] | Pass | Fail | Total
|  |  | Threshold: | 1e-06 | Status: | 0 | 1 | 1
Module: | Wattfuncs
| Function | Expected | Actual | OK? | Parameters (if any)
| Watts |  |  |  | Divider ratio K | Divider ratio K | Ref Volts V1(cor'd) | Shunt impedance R3 | Shunt impedance R3 | Watts dial alpha | Watts buffer gain | R2/R1 | R2/R1 | Vars dial beta | Vars buffer gain | g_1 conductance | Y1 admittance | R2 impedance | R2 impedance | Detector amp volts | Detector phase | Det Amp Gain G | Det Amp phase Gj | Det amp conductance g_d | Det amp admittance Yd
|  | 50 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -1.85e-06 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 35.3553 | #NAME? | #NAME? | 70.7107 | 70.7107 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -1.85e-06 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 200 | #NAME? | #NAME? | 100 | 0 | 2 | 1 | 0 | -0.5 | 1 | 1 | 0 | -3.7e-06 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 250 | #NAME? | #NAME? | 100 | 0 | 1 | 0.2 | 0 | -0.5 | 1 | 1 | 0 | -1.85e-06 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 35.3553 | #NAME? | #NAME? | 100 | 0 | 1 | 0.70711 | 0.70711 | -0.5 | 1 | 1 | 0 | -1.85e-06 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | -50 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | 0.5 | 1 | 1 | 0 | -1.85e-06 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 100 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 2 | 1 | 0 | -3.7e-06 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 35.3553 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 0.70711 | 0.70711 | -1.85e-06 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 50 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 50 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -1.85e-06 | 0.5 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 100 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -0.5 | 1 | 0.0001 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 50 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -1.85e-06 | 1 | 0 | 1000000 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 50.2 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -1.85e-06 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0 | 1000 | 0 | 0 | 0
|  | 50.0707 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -1.85e-06 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0.7854 | 1000 | 0 | 0 | 0
|  | 52 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -1.85e-06 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0 | 100 | 0 | 0 | 0
|  | 50.0707 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -1.85e-06 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0 | 1000 | -0.7854 | 0 | 0
|  | 50.3 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -1.85e-06 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0 | 1000 | 0 | 0.0001 | 0
|  | 50.2 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -1.85e-06 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | -1.5708 | 1000 | 0 | 0 | 0.0001
| Vars |  |  |  | Divider ratio K | Divider ratio K | Ref Volts V1(cor'd) | Shunt impedance R3 | Shunt impedance R3 | Watts dial alpha | Watts buffer gain | R2/R1 | R2/R1 | Vars dial beta | Vars buffer gain | g_1 conductance | Y1 admittance | R2 impedance | R2 impedance | Detector amp volts | Detector phase | Det Amp Gain G | Det Amp phase Gj | Det amp conductance g_d | Det amp admittance Yd
|  | 50 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 35.3553 | #NAME? | #NAME? | 70.7107 | 70.7107 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 200 | #NAME? | #NAME? | 100 | 0 | 2 | 1 | 0 | -2.78e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 250 | #NAME? | #NAME? | 100 | 0 | 1 | 0.2 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 35.3553 | #NAME? | #NAME? | 100 | 0 | 1 | 0.70711 | 0.70711 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | -50 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | 0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 100 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -2.72e-06 | 1 | 1 | 0 | -0.5 | 2 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 50 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 0.70711 | 0.70711 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 50 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 50 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 0.5 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 50 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -0.5 | 1 | 10000000 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 100 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0002 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 100.2 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 20000 | 0 | 1 | 0 | 1000 | 0 | 0 | 0
|  | 50.2121 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0.7854 | 1000 | 0 | 0 | 0
|  | 51 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0 | 100 | 0 | 0 | 0
|  | 50.2121 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0 | 1000 | -0.7854 | 0 | 0
|  | 50.1 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0 | 1000 | 0 | 0.0001 | 0
|  | 49.8 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | -1.5708 | 1000 | 0 | 0 | 0.0001
| Amps |  |  |  | Divider ratio K | Divider ratio K | Ref Volts V1(cor'd) | Shunt impedance R3 | Shunt impedance R3 | Watts dial alpha | Watts buffer gain | R2/R1 | R2/R1 | Vars dial beta | Vars buffer gain | g_1 conductance | Y1 admittance | R2 impedance | R2 impedance | Detector amp volts | Detector phase | Det Amp Gain G | Det Amp phase Gj | Det amp conductance g_d | Det amp admittance Yd
|  | 0.5 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 0.5 | #NAME? | #NAME? | 70.7107 | 70.7107 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 1 | #NAME? | #NAME? | 100 | 0 | 2 | 1 | 0 | -2.78e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 2.5 | #NAME? | #NAME? | 100 | 0 | 1 | 0.2 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 0.5 | #NAME? | #NAME? | 100 | 0 | 1 | 0.70711 | 0.70711 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 0.5 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | 0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 1 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -2.72e-06 | 1 | 1 | 0 | -0.5 | 2 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 0.5 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 0.70711 | 0.70711 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 0.70711 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 0.5 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 0.5 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 1.118 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -0.5 | 1 | 0.0001 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 1 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0002 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 1.002 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 20000 | 0 | 1 | 0 | 1000 | 0 | 0 | 0
|  | 0.50212 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0.7854 | 1000 | 0 | 0 | 0
|  | 0.51039 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0 | 100 | 0 | 0 | 0
|  | 0.50212 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0 | 1000 | -0.7854 | 0 | 0
|  | 0.50101 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0 | 1000 | 0 | 0.0001 | 0
|  | 0.498 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | -1.5708 | 1000 | 0 | 0 | 0.0001
| Volts | 60.903 | #NAME? | #NAME?
|  | 202.9813 | #NAME? | #NAME?
|  | 111.6469 | #NAME? | #NAME?
| Phase |  |  |  | Divider ratio K | Divider ratio K | Ref Volts V1(cor'd) | Shunt impedance R3 | Shunt impedance R3 | Watts dial alpha | Watts buffer gain | R2/R1 | R2/R1 | Vars dial beta | Vars buffer gain | g_1 conductance | Y1 admittance | R2 impedance | R2 impedance | Detector amp volts | Detector phase | Det Amp Gain G | Det Amp phase Gj | Det amp conductance g_d | Det amp admittance Yd
|  | 90 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 45 | #NAME? | #NAME? | 70.7107 | 70.7107 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 90 | #NAME? | #NAME? | 100 | 0 | 2 | 1 | 0 | -2.78e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 90 | #NAME? | #NAME? | 100 | 0 | 1 | 0.2 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 45 | #NAME? | #NAME? | 100 | 0 | 1 | 0.70711 | 0.70711 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | -90 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | 0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 90 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -2.72e-06 | 1 | 1 | 0 | -0.5 | 2 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 45 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 0.70711 | 0.70711 | -1.85e-06 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 45 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -0.5 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 90 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 0.5 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 45 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0.0001 | 0.0001 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 90 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0002 | 10000 | 0 | 0 | 0 | 1000 | 0 | 0 | 0
|  | 89.8856 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 20000 | 0 | 1 | 0 | 1000 | 0 | 0 | 0
|  | 89.9193 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0.7854 | 1000 | 0 | 0 | 0
|  | 87.7543 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0 | 100 | 0 | 0 | 0
|  | 89.9193 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0 | 1000 | -0.7854 | 0 | 0
|  | 89.6569 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | 0 | 1000 | 0 | 0.0001 | 0
|  | 89.7699 | #NAME? | #NAME? | 100 | 0 | 1 | 1 | 0 | -1.39e-06 | 1 | 1 | 0 | -0.5 | 1 | 0 | 0.0001 | 10000 | 0 | 1 | -1.5708 | 1000 | 0 | 0 | 0.0001
AlphaCount
BetaCount
CSUW
CSUV
DegFmW
DegFmV
| rss | 1.4142 | #NAME? | #NAME? | 1 | 1
|  | 1.4142 | #NAME? | #NAME? | -1 | -1
|  | 1 | #NAME? | #NAME? | 1
|  | 5.4772 | #NAME? | #NAME? | 1 | 2 | 3 | 4
| DoF_A | false | #NAME? | FAIL | 1 | 2 | 3 | 1 | 2 |  | Check that unequal lengths are flagged properly
|  | 1 | #NAME? | #NAME? | 1 | 1
|  | 2 | #NAME? | #NAME? | 1 | 1 | 1 | 1
|  | 4 | #NAME? | #NAME? | 1 | 1 | 2 | 2
|  | 4 | #NAME? | #NAME? | 1 | 2 | 3 | 2 | 2 | 2
| CoverFactor | 12.71 | #NAME? | #NAME? | The test values are taken from the table at http://www.physics.nist.gov/Pubs/guidelines/table.b.1.html
|  | 3.18 | #NAME? | #NAME? | Note the necessity of significantly lowering the threshold since this is a rough interpolation (the coarsest coefficient used is 0.5%).
|  | 2.57 | #NAME? | #NAME?
|  | 2.23 | #NAME? | #NAME? | Local threshold: |  | 0.02
|  | 2.09 | #NAME? | #NAME?
|  | 1.984 | #NAME? | #NAME?
Module: | Uncertainties
| Function | Expected | Actual | OK? | Parameters (if any)

## ChangeLog
Change Log
Date | Sheet | Changes
2007-02-08 | Data | Added conditional formatting for the parameter input cells, red cells indicate a failure of some basic sanity checks. |  |  | Specifically, the source, voltage range, frequency range and watts/vars selector are checked.
2007-02-08 | Unit Tests | Added a unit tests sheet. Currently Incomplete (only atan2 and rss).
2007-02-09 | VB | Deleted all unused functions from the Wattfuncs module: WorV, WorV2, LagLead, AlphaCountFine, BetaCountFine
2007-02-09 | Unit Tests | Added unit tests for Watts, Vars, Volts, DoF_A and CoverFactor. Removed test for atan2 since it was testing the excel version.
2007-02-12 | Unit Tests | Corrections to alpha and beta values in the Watts and Vars checks to cancel out the dial offset adjustments and give the nice round answers we desire (rather than calculating the correct answers).
2007-02-13 | Unit Tests | Added a unit test for Amps.
2007-02-14 | Unit Tests | Added a unit test for Phase.
2007-02-16 | Sheet 10/TVH2.32 | Renamed Sheet 10 to TVH2.32.
2007-02-19 | Data | Added Data Validation to the frequency, range, shunt selection, phase selection and source selection cells of the parameters block.
2007-02-27 | PowerCalc | Added a negative sign to each of the angle arguments in cells FG10 and FG12 to give us the correct PF or VF for all four quadrants. i.e. we convert our American phase angle to a European one.
2007-03-22 | TABLE RM11 | Updated the RM-11 template. Table headings: use U instead of V; italicise U, I and k; change var factor to "Reactive Power Factor".
2007-03-29 | RM11Unc | Updated the RM11 uncertainties to accomodate behavior observed in S21672.
2007-03-29 | Data | Put "#N/A" into Cell N2. Type B uncertainty column.
2008-03-05 | Data | Corrected wrong spreadsheet version number on Data page
2008-11-12 | constants | Version n: new values of constants
2008-11-12 | Data | Added some suggested Prime Points
2009-03-06 | constants | Version o: new values of constants
2009-03-19 | TABLE RM11 | Updated the RM-11 template. Table headings: use Imp/W.h and Imp/var.h instead of Imp/Watt.h and Imp/Var.h
2009-05-14 | Data | Corrected and included further RD-31 data headings to the far right of Data page.
2009-05-14 | TABLE RM11 | Included check of expanded uncertainty against current CMC's in column P. Also changed table font to Times New Roman to match Report font.
2009-05-14 | Uncertainties | The standard uncertainty values for the Watt-bridge have been reverted back to the values from Pb_Auto_V1g (Feb 2005), as these are the last approved values consistent with our current CMC's.
2009-05-26 | constants | In version "o" the constants were incorrectly copied across, causing the Vlookups from other worksheets to incorrectly reference the old constants. This has been corrected.
2009-06-26 | WattFuncs Macro | Updated the hard coded constants for IVD offsets. An error in the way in which they are applied has also been corrected.
2010-03-02 | TABLE RD31 | Added worksheet for RD31 results. Also has a "point key" column added required for looking up corrections for full star and delta 3 phase calibrations.
Power factor label now calculates for full four quadrant. The power factor sign is calculated in the table.
2010-03-02 | D3000 Star Table | Added worksheet for D3000 Star results. Table rows are selected from "Star Data" worksheet. Power factor is full four quadrant calculation
2010-03-02 | Star Data | Added worksheet for analysis of csv file output from controller.py. Data is pasted from E5 and other formula copied down. Calculates from a group of readings the average error, std unc, exp unc, dfm, cover factor, whilst looking up and applying correction for use of RD31 as calibration transfer.
Worksheet currently does not support single phase analysis.
2010-03-02 | D3000 Delta Table | Added worksheet for D3000 Delta results. Table rows are selected from "2W Delta Data" worksheet. Four quandrant Power factor calculation plus each channel phase angle calculation.
2010-03-02 | 2W Delta Data | Added worksheet for analysis of csv file output from controller.py. Data is pasted from E5 and other formula copied down. Calculates from a group of readings the average error, std unc, exp unc, dfm, cover factor, whilst looking up and applying correction based on each RD31 channel calibration.
2010-11-17 | WattFuncs Macro | Updated the hard coded constants for IVD offsets, measured as part of the WB calibration.
2010-11-22 | constants | New values of constants updated from Nov 2010 WB calibration.
2011-01-14 | constants | New values of constants updated from Dec 2010 WB calibration. Nov calibration values not trusted due to issue with Watt-bridge internal cabling.
2011-02-28 | constants | Updated shunt and WG2 PRT values
2011-04-12 | D3000 Star Table | Corrected Voltage and Current table headings. | Fixed mistake in RD31 error, unc & cover factor lookup on D3000 Star Data and 2W Delta Data pages.
2011-04-18 | constants | Updated value for MSL 0.2 Ohm shunt. This is based more on an average value over the period the shunt was calibrated (BM & TJS).
2011-04-29 | Uncertainties | Changed cells K10, K12 and O10,O12 so that the condition for dial settings below 0.4, an uncertainty of 15 ppm applies. This matches the comments for when the last CMC's were approved.
The standard uncertainty of 15 ppm for dial settings below 0.4 should be good down to a watts:vars dial combination of 0.02:0 or a watts:vars of 0.05:0, based on the variability of the offset (see "IVD offset SR830 readings Mar11.xls").
2011-04-29 | WattFuncs Macro | Updated the hard coded constants for the IVD offsets. The calculation for these have been simplified based on the use of a constant frequency of 53 Hz.
The numbers now represent a figure that is effectively the IVD voltage offset at the bridge centre.
2011-05-02 | constants | Updated values for Buffer amplifier gain (cells D32-D35), this includes ~9 ppm change in gain of the inverter amplifier. These values apply after the op42 op-amps were changed on the A6 board on 30/3/11.
2011-05-02 | constants | Values for ESI(1) and ESI(10) updated.
2011-05-05 | Data | changed cell E1 to read Pb_Auto_V1u instead of t.
2011-05-06 | Uncertainties | Added comments to the Dial uncertainty values.
2011-05-18 | Uncertainties | Changed the values for each dial linearity to 2 ppm. Added two more entries for each of Watts and Vars dial offset uncertainty
2011-05-18 | UncCalc | Modified the formula in cells K10,K12 and O10,O12 to be an uncertainty of dial setting (gain plus linearity) plus offset dial uncertainty.
2011-06-14 | TABLE RM11 | Updated Lead/Lag calculation to use IEC phase angle and to match RD31 & D3000 worksheets. Changed Power/Reactive Power Factor heading to depend on watts/vars label of first table line.
2011-08-22 | constants | Placeholder - need to implement correction for counter crystal error (HP53131A (S# 316A04115) currently measured at -1.2 ppm (19.9 degrees))
2012-06-26 |  | Updated all phase column headings to either IEC phase or Non-IEC.
2012-06-26 | constants | "CountCorr" named cell added for counter correction. Measured at -1.58 ppm for HP53132A (S# KR91200497) on 25/6/12
2012-06-26 | PowerCalc | Cells DP-DT and FO-FS updated with calculation to correct for counter error - using "CountCorr" on constants page.
2012-06-26 | constants | Updated constants from Apr/May 2012 calibration of Wattbridge
2012-06-26 | 2W Delta Data | Corrected columns ET, EU and FA-FD with correct lookup column of TABLE RD31.
2012-06-26 | TABLE RD31 | Change key to have Mround to 0.5 and "0.0" for voltage text. Matching changes made to 2W Delta Data
2012-06-26 | WattFuncs Macro | Updated the hard coded constants for IVD offsets, measured as part of the WB calibration.
2013-01-10 | RD31 Data | Created worksheet putting RD31 readings (RD31 data headings removed from "Data" worksheet)
2013-01-10 | K2006 Data | Created worksheet for K2006 readings
2013-01-10 | TABLE RD31 | Correctly updated lead/lag calculation to other than the first row.
2013-01-10 | TABLE RM11 | Correctly updated lead/lag calculation to use IEC phase angle and to match RD31 & D3000 worksheets.
2013-01-14 | constants | Updated value of C1 from recent calibration of capacitor
2013-01-18 | K2006 Data | Added column headings for K2006 Data.
2013-06-13 | Data | Updated column B to include FLUKE (6105A & 6106A) and FLUHIGH (52120A) as possible sources. Column E title (E9-10) updated with phase convention of Fluke sources for correct lead/lag phase setting.
2013-06-17 | Data | Input phase angles (column E) restricted to being between -360 and 360 degrees. (pb_auto_v1xx will convert phase to between -180 to 180 degrees for the Fluke 6105A system.)
2013-07-30 | UncCalc and Uncertainties | Changed formula in columns K and O to include phase angle of IVD dials. The values are included in opposite dial and the numerical values are numbers 24 and 25 on the Uncertainties page.
2013-10-31 | UncCalc | Corrected C1 phase angle (loss angle) calculation. Changed cell Q3 from 23 to 24. This will result in the loss angle uncertainty being calculated as 15 urad instead of previously being 15 ppm of g_1.
2013-11-26 | WattFuncs Macro | Corrected IVD offset calculation to be proportional to V1. (Included multiplication of IVD offset by RefVolts).
2013-12-05 | Uncertainties | Updated uncertainties and Dof's to values evaluated for the APMP K5.1 power comparison (report S21885)
2013-12-05 | constants | Updated value of g_1 based on value determined for APMP k5.1 comparison |  | changes for ab version
2013-12-05 | PowerCalc | Updated FZ column calculation of C1 based its temperature when calibrated (23.9 degC)
2013-12-05 | Table worksheets | Added column for phase angle in results tables
2014-05-22 | constants | Added the calibration temperature of C1 as a new constant. This is to try to better ensure it gets updated when C1 (and WB) is calibrated.
2014-05-22 | PowerCalc | Updated FZ column calculation of C1 using the new C_1_CalTemp constant.
2014-05-28 | constants | Updated constants from Feb-May 2014 calibration of Wattbridge
2014-06-04 | WattFuncs Macro | Updated the hard coded constants for IVD offsets, measured as part of the WB calibration.
2014-07-16 | Star Data | Corrected Ref meter phase label as IEC for RD31. The sign is changed before being written to the csv file.
2014-12-02 | Table worksheets | For impulses per watt hour or var hour the units now have brackets around watt.h and var.h. So now reads (Imp/(W.h)) or (Imp/(var.h)). This ensures the 'h' is technically correct as part of the denominator.
2015-01-27 | constants | Update constants for WG2 and 3478A. Also change counter correction to zero for the change of counter to the Agilent 53230A. |  |  |  | changes for ad version